## Service report
TECHNICIAN REPORT
NO REPORT | TANGGAL | NAMA CUSTOMER | NAMA ALAT |  | SN | JENIS PEKERJAAN | URAIAN PEKERJAAN | STATUS | KETERANGAN | PELAKSANA | MASA BERAKHIR GARANSI | SERVICE KONTRAK
19000612 | 31 April 2026 | RSUPN Cipto Mangunkusumo | Mesin Anasthesia | Draeger/Atlan A300 XL | ASMM0060 | Survey | Pengecekan unit demo. Baterai failure, ganti baterai. unit berfungsi dengan baik. Dilakukan pendampingan pasien | Ready to use |  | Nanda
08000838 | 2026-04-29 | RSUD Kajen Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHA0233 | Survey | Sensor check fail. Hit software, check 2/2 distributing valve. Port 2/2 distributing valve rusak | Pending |  | Amna
08000838 | 2026-04-29 | RSUD Kajen Pekalongan | Mesin Anasthesia | Draeger/Primus | ASLD0023 | Survey | Sensor check fail. Hit software, check 2/2 distributing valve. Port 2/2 distributing valve rusak | Pending |  | Amna
08000838 | 2026-04-29 | RSUD Kajen Pekalongan | Mesin Anasthesia | Draeger/Primus | ASMA0191 | Survey | Sensor check fail. Hit software, check 2/2 distributing valve. Port 2/2 distributing valve rusak | Pending |  | Amna
08000838 | 2026-04-29 | RSUD Kajen Pekalongan | Mesin Anasthesia | Draeger/Primus | ASLD0024 | Survey | Sensor check fail. Hit software, check 2/2 distributing valve. Port 2/2 distributing valve rusak | Pending |  | Amna
19000610 | 2026-04-28 | RSUP dr. Kariadi | Dental Unit | Kavo/1058 Life | 1002719 | Survey | Pengecekan handpiece bor (SN: 1018501) bearing set dan cap perlu diganti | Pending |  | Nanda
08000840 | 2026-04-28 | RSUD Kayen Pati | Mesin Anasthesia | Draeger/Primus | ASNH0252 | Service | Alarm co2 apnea. System check, Cosy check. Perlu penggantian sample line dan waterlock | Ready to use |  | Amna
08000840 | 2026-04-28 | RSUD Kayen Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASSH0043 | Service | Alarm co2 apnea. System check, Cosy check. Perlu penggantian sample line dan waterlock | Ready to use |  | Amna
19000611 | 2026-04-28 | RSUP dr. Kariadi | Baby Incubator | Draeger/Isolette 8000 | MC09035 | Survey | Air probe failed. Cek unit, exchange probe dgn unit lain, OK; Sensor module isolette bermasalah | Pending |  | Nanda
07000985 | 2026-04-24 | RSUP dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Atlan A3500 XL | ASMN0009 | Survey | Pengecekan kondisi alat dan kelengkapan aksesoris unit demo | Ready to use |  | Wahyu
9000856 | 2026-04-23 | RSUD Dr. Chasan Boesoerie | Ventilator | Draeger/Evita V300 | ASND0029 | Survey | FIO2 error, battery drop, air supply fail, nebulizer fail. Tarik error. Perlu penggantian pressure regulator | Pending |  | Ruslan
08000839 | 2026-04-23 | RS Mitra Bangsa Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0041 | Service | O2 sensor fail. System check. Perlu oenggantian O2 capsule | Pending |  | Wahyu
08000839 | 2026-04-23 | RS Mitra Bangsa Pati | Ventilator | Draeger/Babylog VN600 | ASTB0024 | Service | O2 sensor fail. System check. Perlu oenggantian O2 capsule | Pending |  | Amna
19000608 | 2026-04-22 | RSUD Temanggung | Infant Warmer | Draeger/Resuscitaire | JY10905 | Survey | Error 21. Cek unit, service mode, setting 5V preference. Running OK | Ready to use |  | Nanda
19000609 | 2026-04-22 | RSUD Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Survey | Pengecekan unit. Unit berfungsi normal | Ready to use |  | Nanda
07000984 | 2026-04-22 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draefer/Babylog VN500 | ASNL0107 | Maintenance | Preventive Maintenance total care | Ready to use |  | Wahyu
07000983 | 2026-04-22 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draefer/Babylog VN500 | ASKL0092 | Maintenance | Preventive Maintenance total care | Ready to use |  | Wahyu
9000853 | 2026-04-22 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | CPD41220H22 | Survey | Training cara burning hasil expose burning to CD | Ready to use |  | Ruslan
22000025 | 2026-04-21 | Siloam Hospital Manado | Hemodialisa | Bbraun/Analog | 264993 | Service | Pressure vena bocor. Ganti part monometer | Ready to use |  | Ari
19000607 | 2026-04-21 | RSUD dr. Adhyatma | Mesin Anasthesia | Draeger/Primus | ASHJ0179 | Survey | Unit sering mati saat unit prepaer dinyalakan. Cek baterai, cek VGC, Restrictor mepet dengan VGC. | Pending |  | Nanda
19000606 | 2026-04-21 | RSUD dr. Adhyatma | Mesin Anasthesia | Draeger/Primus | ASHH0122 | Survey | Pressure sensor fail, device check, membran piston sudah menipis. Disarankan penggantian service set 6 tahun dan restrictor | Pending |  | Nanda
07000982 | 2026-04-21 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draefer/Babylog VN500 | ASNL0106 | Maintenance | Preventive Maintenance total care | Ready to use |  | Wahyu
19000605 | 2026-04-21 | RSUD Tidar Magelang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0496 | Survey | Pengecekan unit. Pergerakan baik, hasil expose bagus. Unit berfungsi dengan baik | Ready to use |  | Nanda
9000852 | 2026-04-22 | RSUD Dr. Chasan Boesoerie | MOT | SYAS | - | Maintenance | Pencucian HVQC, penggantian mcb saklar c-arm, perbaikan kebocoran pendant. Unit berfungsi normal kembali | Ready to use |  | Ruslan
6001028 | 2026-04-20 | RSUD Cilacap | Baby Incubator | Draeger/Isolette C2000 | CF32804 | Service | Contoller failure; Bersihkan kipas blower, heater plate. Unit OK | Ready to use |  | Dwi
6001028 | 2026-04-20 | RSUD Cilacap | Baby Incubator | Draeger/Isolette C2000 | CF32807 |  | Contoller failure; Bersihkan kipas blower, heater plate. Unit OK | Ready to use |  | Dwi
6001028 | 2026-04-20 | RSUD Cilacap | Bubble CPAP | Draeger/Oxymixer | QM20230175 |  | Contoller failure; Bersihkan kipas blower, heater plate. Unit OK | Ready to use |  | Dwi
18000719 | 2026-04-20 | RSUD ODSK Provinsi Sulawesi Utara | Bubble CPAP | Draeger/Oxymixer | QMHF2022208431 | Service | Alarm mixer. Tekanan O2 dan tabung tidan sinkron dengan tekanan air. Setting O2 bar sesuai dengan tekanan air | Ready to use |  | Irfan
08000836 | 2026-04-20 | RSUD dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASSE0030 | Survey | System check. O2 capsule expired | pending |  | Amna
08000836 | 2026-04-20 | RSUD dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0109 | Survey | System check. O2 capsule expired | pending |  | Amna
08000836 | 2026-04-20 | RSUD dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Atlan A350 | ASUK0231 | Survey | System check. O2 capsule expired | pending |  | Amna
17000238 | 2026-04-20 | RS Immanuel | Infus Pump | Bbraun/Infusomat Compact Plus | 001222 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Bahar
6001027 | 2026-04-18 | RSUD KRT Setjonegoro | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0302 | Survey | Pengecekan unit | pending |  | Dwi
6001026 | 2026-04-18 | RSUD KRT Setjonegoro | Ventilator | Draeger/Savina 300 | ASLL0400 | Survey | Internal battery is missing, FIO2 measurement failed. Tarik error. Disarankan penggantian service set 2 year, O2 sensor, VE Trace 2pcs | pending |  | Dwi
9000855 | 2026-04-17 | RSUD Weda | Dental Unit | Kavo/1058 Life | 10205459 | Survey | Pergerakan kursi tidak bisa, scaller sering macet. Pengecekan board PCBA chair dan selang scaller; Disarankan penggantian selang scaller dan board pcba | pending |  | Ruslan
12000427 | 2026-04-17 | RSUD Aloei Saboe | Ventilator | Draeger/Evita V300 | ASME0153 | Survey | Pengecekan unit, o2 supply dan air tidak terbaca | pending |  | Marshall
12000426 | 2026-04-16 | RSUD Ainun Habibie | Infant Warmer | Draeger/Resuscitaire | DY03434 | Survey | Blender O2 bocor. Overhaul kit rusak | pending |  | Marshall
12000425 | 2026-04-16 | RSUD Ainun Habibie | Infant Warmer | Draeger/Resuscitaire | DY03438 | Survey | error 05, blender o2 bocor. Heater disarankan diganti | pending |  | Marshall
12000424 | 2026-04-16 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJH0030 | Survey | error 05, blender o2 bocor. Heater disarankan diganti | pending |  | Marshall
12000424 | 2026-04-16 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJE0060 | Survey | Tidak ada kelembaban. Conv. kit heater rusak | pending |  | Marshall
12000424 | 2026-04-16 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJE0059 | Survey | Tidak ada kelembaban. Conv. kit heater rusak | pending |  | Marshall
12000424 | 2026-04-16 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJE0061 | Survey | Tidak ada kelembaban. Conv. kit heater rusak | pending |  | Marshall
12000424 | 2026-04-16 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJE0053 | Survey | Tidak ada kelembaban. Conv. kit heater rusak | pending |  | Marshall
12000424 | 2026-04-16 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJH0031 | Survey | Tidak ada kelembaban. Conv. kit heater rusak | pending |  | Marshall
18000723 | 2026-04-16 | RSUD Maria Walanda Maramis | Ventilator | Draeger/Babylog VN800 | ASPF0044 | Survey | Baterai drop, tidak mau mengisi. Unit masih bisa digunakan tanpa baterai. | pending |  | Irfan
18000723 | 2026-04-16 | RSUD Maria Walanda Maramis | Ventilator | Draeger/Babylog VN800 | ASPF0047 | Survey | Baterai drop, tidak mau mengisi. Unit masih bisa digunakan tanpa baterai. | pending |  | Irfan
12000423 | 2026-04-15 | RSUD Dr. M. M. Dunda | C-arm | Apelem/Evo R+ | 551-18-031-636 | Survey | Pengecekan ulang kerusakan unit; Travo rusak, PSU 24V rusak, power board rusak, sigma board rusak, software bluescreen, inverter rusak | pending |  | Marshall
18000720 | 2026-04-15 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | B.Braun/Infusomat | 1166 | Service | LCD error (cairan masuk ke unit). Bongkar alat dan dibersihkan serta dikeringkan. Unit sudah berfungsi kembali | Ready to use |  | Irfan
19000604 | 2026-04-15 | RS JIH Solo | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0125 | Maintenance | Maintenance Preventive Care | Ready to use |  | Nanda
19000603 | 2026-04-15 | RS JIH Solo | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Maintenance | Maintenance Preventive Care | Ready to use |  | Nanda
08000835 | 2026-04-15 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASUM0001 | Survey | System check, warmer fail. Penggantian Tuelle, Gasket , Spout T, Clampring | Ready to use |  | Amna
19000602 | 2026-04-14 | RSUP Surakarta | USG | Mindray/ DP20T | LK7B36001327 | Service | Connecting printer ke unit. Unit sudah bisa printing dengan tombol pada keyboard | Ready to use |  | Nanda
08000834 | 2026-04-14 | RSI Arafah Rembang | Mesin Anasthesia | Draeger/Fabius Plus | ASUA0091 | Maintenance | System check. Unit OK | Ready to use |  | Amna
16000316 | 2026-04-14 | M&P Clinic | Dental Unit | Planmeca/Compact i | UC16191927 | Maintenance | Preventive Maintenance | Ready to use |  | Farhat, Ipan
19000614 | 2026-04-13 | RSUD dr. Moewardi | USG | Mindray/Resona I9 | JA8-3A00114 | Survey | Installasi unit demo | Ready to use |  | Nanda
19000614 | 2026-04-13 | RSUD dr. Moewardi | Mesin Anasthesia | Draeger/Atlan A100 | ASMM0059 | Survey | Installasi unit demo | Ready to use |  | Nanda
08000833 | 2026-04-13 | RSI PKU Muhammadiyah Pekajangan | Mesin Anasthesia | Draeger/Atlan A300 | ASUM0315 | Maintenance | System check. Unit OK | Ready to use |  | Amna
08000833 | 2026-04-13 | RSI PKU Muhammadiyah Pekajangan | Mesin Anasthesia | Draeger/Atlan A300 | ASUE0013 | Maintenance | System check. Unit OK | Ready to use |  | Amna
16000315 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000834 | Survey | Suction tidak menyedot. Cek jalur suction, ada patah pada dudukannya | Pending |  | Farhat, Ipan
16000315 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000825 | Survey | Suction tidak menyedot. Cek jalur suction, ada patah pada dudukannya | Pending |  | Farhat, Ipan
16000315 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000833 | Survey | Suction tidak menyedot. Cek jalur suction, ada patah pada dudukannya | Pending |  | Farhat, Ipan
16000315 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000853 | Survey | Suction tidak menyedot. Cek jalur suction, ada patah pada dudukannya | Pending |  | Farhat, Ipan
16000313 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000823 | Survey | Suction tidak menyedot. Cek jalur suction, ada patah pada dudukannya | Pending |  | Farhat, Ipan
16000313 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000843 | Survey | bocor angin pada meja element. Cek meja elemenm cek selangnya, bocor pada switch. Dental simulator masih bisa digunakan hanya ada suara bocor angin | Pending |  | Farhat, Ipan
16000313 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000838 | Survey | bocor angin pada meja element. Cek meja elemenm cek selangnya, bocor pada switch. Dental simulator masih bisa digunakan hanya ada suara bocor angin | Ready to use |  | Farhat, Ipan
16000313 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000842 | Survey | bocor angin pada meja element. Cek meja elemenm cek selangnya, bocor pada switch. Dental simulator masih bisa digunakan hanya ada suara bocor angin | Ready to use |  | Farhat, Ipan
16000313 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000820 | Survey | bocor angin pada meja element. Cek meja elemenm cek selangnya, bocor pada switch. Dental simulator masih bisa digunakan hanya ada suara bocor angin | Ready to use |  | Farhat, Ipan
16000313 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000811 | Survey | bocor angin pada meja element. Cek meja elemenm cek selangnya, bocor pada switch. Dental simulator masih bisa digunakan hanya ada suara bocor angin | Ready to use |  | Farhat, Ipan
16000313 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000855 | Survey | bocor angin pada meja element. Cek meja elemenm cek selangnya, bocor pada switch. Dental simulator masih bisa digunakan hanya ada suara bocor angin | Ready to use |  | Farhat, Ipan
16000313 | 2026-04-13 | FKG Moestopo | Dental Simulator | Kavo/DSE Expert | 0000834 | Survey | bocor angin pada meja element. Cek meja elemenm cek selangnya, bocor pada switch. Dental simulator masih bisa digunakan hanya ada suara bocor angin | Ready to use |  | Farhat, Ipan
9000851 | 2026-04-13 | RSUD Dr. Chasan Boesoerie | Dirty Linen Trolley | SYAS | 260269601000400001 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000851 | 2026-04-13 | RSUD Dr. Chasan Boesoerie | Dirty Linen Trolley | SYAS | 260269601000400002 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000851 | 2026-04-13 | RSUD Dr. Chasan Boesoerie | Dirty Linen Trolley | SYAS | 260269601000400003 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000851 | 2026-04-13 | RSUD Dr. Chasan Boesoerie | Dirty Linen Trolley | SYAS | 260269601000400004 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
19000613 | 2026-04-11 | RSJ Prof Dr Soerojo Magelang | USG | Mindray/Resona I9 | - | Survey | Packing unit demo | Ready to use |  | Nanda
12000422 | 2026-04-10 | RSUP Ratatotok Buyat | Lampu Operasi | Draeger/Polaris 600 | ASND0066 | Survey | Lampu arm 2 mati, cek unit engsel arm 2 lepas. Pengencangan engsel arm 2, unit berfungsi OK | Ready to use |  | Marshall
17000237 | 2026-04-10 | RSUD dr. Adjidharmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Survey | Motor grid rusak, UPS PU BOX fault tidak charging. Cek pergerakan alat, software. Disarankan penggantian battery dan motor grid | Ready to use | Alat masih bisa digunakan dengan keterbatasan | Bahar
08000832 | 2026-04-10 | RS Aisyah Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0048 | Survey | O2 sensor fail dan flow sensor fail. Check O2 sensor, flow sensor. O2 sensor dan flow sensor expired | Pending |  | Amna
12000420 | 2026-04-09 | RSUD Kotamobagu | C-arm | Apelem/Evo R+ | 551-18-031-654 | Survey | Koneksi to sigma board failed. Sigma board rusak | Pending | Sudah diberikan penawaran | Marshall
19000600 | 2026-04-09 | RSUD Dr. Gunawan Mangunkusomo Ambarawa | Incubator Transport | Draeger/TI500 | KY11399 | Service | Baterry fault. Soket baterai ada yang bengkok. Sudah diperbaiki | Ready to use |  | Nanda
19000600 | 2026-04-09 | RSUD Dr. Gunawan Mangunkusomo Ambarawa | Incubator Transport | Draeger/TI500 | KY11400 | Service | Baterry fault. Soket baterai ada yang bengkok. Sudah diperbaiki | Ready to use |  | Nanda
17000236 | 2026-04-09 | RSUD Berkah Pandeglang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0497 | Survey | Resistor break untuk meja turun rusak. Cek pergerakan platinum, test expose, cek software. Alat masih bisa digunakan | Ready to use |  | Bahar
18000724 | 2026-04-09 | RSUP Jayapura | Digital Radiografi Fluoroskopi | Apelem/Platinum | H248L0157 | Maintenance | Preventive Maintenance (PM) | Ready to use |  | Irfan
19000599 | 2026-04-08 | RSUP dr. Kariadi | Ventilator | Draeger/babylog VN500 | ASLN0005 | Survey | Unit restart sendiri. Reload software ulang. Running on going | Pending |  | Nanda
9000790 | 2026-04-08 | RSD Tidore | Digital Radiografi | Apelem/Camargue DR | CPD41220H22 | Service | Software PC workstation error; Reload software pc workstation, penggantian SSD 1TB, penggantian RAM DDR4. Unit sudah berfungsi dengan baik | Ready to use |  | Ruslan
6001019 | 2026-04-08 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASHA0095 | Survey | Gas supply sensor fail, O2 supply fail, air supply fail. Tarik error. Perlu penggantian flow sensor | Pending |  | Dwi
3000939 | 2026-04-08 | RSD Tidore | Digital Radiografi | Apelem/Camargue DR | CPD41220H22 | Service | Software PC workstation error; Reload software pc workstation, penggantian SSD 1TB, penggantian RAM DDR4. Unit sudah berfungsi dengan baik | Ready to use |  | Imam
17000235 | 2026-04-08 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0488 | Survey | Grid limit switch, kv low limit dan UPS rusak. Test gerak alat, test expose, cek software. Disarankan mengganti inverter board, HV module, Fan, UPS console, UPS Workstation | Pending |  | Bahar
12000417 | 2026-04-07 | RSUD Kotamobagu | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0270 | Survey | Pengecekan unit. Unit berfungsi normal (Workstation, console, pergerakan meja, generator xray, ups pu box dan ups console) | Ready to use |  | Marshall
18000721 | 2026-04-07 | RSUD ODSK Provinsi Sulawesi Utara | Ventilator | Draeger/babylog VN800 | ASTH0036 | Service | Kalibrasi internal gagal. Cek flow sensor dan kabel harnes. Flow sensor rusak (putus); Penggantian flow sensor dengan unit cadangan, unit sudah berfungsi | Ready to use |  | Irfan
16000312 | 2026-04-07 | Signature Dental care | Dental Unit | Kavo/1058 Life | 10207191 | Service | Mainboard rusak; Penggantian mainboard | Ready to use |  | Farhat
17000234 | 2026-04-07 | RSUD dr. Drajat Prawiranegara | X-ray Digital | Apelem/Camargue DR | 2510220562 | Survey | Cek pergerakan alat, cek software, test expose. Unit masih berfungsi dengan baik | Ready to use |  | Bahar
17000233 | 2026-04-07 | RSUD dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Survey | Cek keseluruhan. Test expose. Unit berfungsi dengan baik | Ready to use |  | Bahar
08000830 | 2026-04-07 | RSUD dr. R. Soedjati Soemodiarjo | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | H24L0156 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
3000938 | 2026-04-07 | RSUD Kotamobagu | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0270 | Survey | Pengecekan seluruh fungsi unit. Workstation OK, console OK, pergerakan meja OK, generator xray OK, ceiling OK, Ups console OK, ups pubox OK | Ready to use |  | Imam
18000725 | 2026-04-07 | RSUD Dok II Jayapura | Instrument | BBraun/Craniotomoy Set | - | Instalasi | Instrument lengkap | Ready to use |  | Irfan
12000414 | 2026-04-06 | RSUD Kotamobagu | C-arm | Apelem/Evo R+ | 551-18-031-654 | Service | Penggantian board B5 dan reload software. Unit OK | Ready to use |  | Marshall
19000596 | 2026-04-06 | RSUD dr. Moewardi | USG | Mindray/M9 | CCI-04002265 | Service | Track ball tidak bisa naik turun. Cek roller back. Trackball perlu diganti | Pending |  | Nanda
8000827 | 2026-04-06 | RSUD Dr. Loekmono Hadi Kudus | Mesin Anasthesia | Draeger | ASPH0121 | Service | APL leak. Check breathing system, check flow sensor, check water trap, check sample line. Running OK | Ready to use |  | Amna
3000937 | 2026-04-06 | RSUD Kotamobagu | C-arm | Apelem/Evo R+ | 55118031654 | Service | Power supply error. Pemasangan board B5 flu/rad, dan reload software | Pending | Unit belum berufngsi karna ada kerusakan lainnya yaitu adanya komponen yang terbakar pada sigma board | Imam, Marshall
17000232 | 2026-04-06 | RSUP Fatmawati | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0493 | Survey | Software lemot, cek data software, test expose. Unit berfungsi dengan baik | Ready to use |  | Bahar
12000413 | 2026-04-02 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | J166F0184 | Survey | Unit terendam air, pengecekan unit. Tarik error. Unit perlu penggantian Pu Box | Pending |  | Marshall, Ari
8000826 | 2026-04-02 | RSUD Batang | Mesin Anasthesia | Draeger/Atlan A300 | ASTH0152 | Survey | System check ventilator fail. System check, penggantian kit ventilator assembly. Running OK | Ready to use |  | Amna
3000936 | 2026-04-02 | RSUD Pagelaran | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/136/C2 | Maintenance | Preventive Maintenance Warranty. Unit OK, console OK, workstation OK | Ready to use |  | Imam
07000980 | 2026-04-02 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASKE0070 | Survey | Pengecekan unit. Blower failure, internal baterai failure, eksternal baterai failure, air filter expired, particle filter expired | Ready to use | Unit masih bisa digunakan dengan keterbatasan | Wahyu
07000979 | 2026-04-02 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASKL0094 | Survey | Pengecekan kondisi alat. Blower unit failure, internal baterai failure, eksternal baterai failure, air filter expired, particle filter expired | Ready to use | Unit masih bisa digunakan dengan keterbatasan | Wahyu
07000978 | 2026-04-02 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASKE0069 | Survey | Pengecekan unit. Touchscreen slow respon, blower unit failure, internal baterai failure, eksternal baterai failure, air filter expired, particle filter expired | Ready to use | Unit masih bisa digunakan dengan keterbatasan | Wahyu
07000976 | 2026-04-02 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASHN0040 | Survey | Pengecekan kondisi alat. Baterai pack, kit air filter dan particle filter perlu diganti | Ready to use | Unit masih bisa digunakan dengan keterbatasan | Wahyu
07000975 | 2026-04-02 | RSUD Panembahan Senopati | Ventilator | Draeger/Babylog VN500 | ASHJ0032 | Survey | Pengecekan unit. PCBA Control unit M1.6, baterai, dan kit air filter serta particle filter perlu diganti | Ready to use | Unit masih bisa digunakan dengan keterbatasan | Wahyu
07000974 | 2026-04-02 | RSUP dr. Soeradji Tirtonegoro | Infant Warmer | Draeger/Babyleo TN500 | ASTL0117 | Survey | Sensor box faulty, peripheral skin sensor faulty. Cek error log, communication sensor box. Reconnect sensor box. Unit OK | Ready to use |  | Wahyu
18000728 | 2026-04-01 | RSUD Anuntaloko | Bor tulang | Aesculap/ELAN4 | 2513100 | Training | Training user | Ready to use |  | Irfan R
07000965 | 2026-03-31 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0091 | Survey | Internal gas calibration failed, gas supply unit failed. Cek error log,reload software. Gas supply tidak dapat terhubung dengan ventilator, blower unit perlu diganti | Pending |  | Wahyu
9000789 | 2026-03-31 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASHJ0178 | Survey | Internal battery discharge; Masuk service mode, cek kondisi baterai dan charge baterai selama 8 jam. Baterai normal | Ready to use |  | Ruslan
6001024 | 2026-03-31 | RSUD Cilacap | USG | Mindray/M9 | CC1-51006326 | Survey | Printing error. Setting ulang koneksi printing. Unit sudah OK | Ready to use |  | Dwi
60001023 | 2026-03-30 | RSUD Ajibarang | Ventilator | Draeger/Babylog VN500 | ASMK0054 | Survey | Baterai failure. Baterai perlu diganti | Pending |  | Dwi
8000825 | 2026-03-30 | RSUD Kajen Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHA0233 | Service | valve/ventilator fail. System check, ganti MV2. Belum normal | Pending | Unit perlu diganti MV2 dan service set 3 tahun | Amna
19000595 | 2026-03-30 | RSUD Simo Boyolali | Jaudince Meter | Draeger/JM-105 | H3505875 | Survey | Muncul notif time due cal. Unit dibawa untuk dikalibrasi | Pending |  | Nanda
17000231 | 2026-03-30 | RS Al Islam | Syringe Pump | Bbraun/Perfusor Compact Plus | 102716 | Maintenance | Unit berufngsi dengan baik | Ready to use |  | Bahar, Farhat
17000230 | 2026-03-30 | RS Al Islam | Infus Pump | Bbraun/Infusomat Compact Plus | 11335 | Maintenance | Unit berufngsi dengan baik | Ready to use |  | Bahar, Farhat
8000824 | 2026-03-27 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Babylog VN600 | ASSK0028 | Maintenance | General check | Ready to use |  | Amna
8000824 | 2026-03-27 | RS PKU Muhammadiyah Mayong | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ0081 | Maintenance | General check | Ready to use |  | Amna
3000935 | 2026-03-27 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Maintenance | Semua unit berfungsi dengan baik dan siap digunakan | Ready to use |  | Imam
8000828 | 2026-03-26 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Atlan | ASTJ0118 | Survey | System check fail. Refresh training | Ready to use |  | Amna
6001017 | 2026-03-26 | RST Wijayakusuma | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0176 | Survey | O2 sensor fail, abterai fail, leaking system. Tarik error log. Perlu diganti Cosy, o2 sensor dan service set 3 tahun | Pending |  | Dwi
9000788 | 2026-03-19 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | CPD41220H22 | Survey | DROC system error. Restart pc dan cek software. | Pending | Disarankan reload software | Ruslan
22000020 | 2026-03-17 | RSUD ODSK Provinsi Sulawesi Utara | Hemodialisa | Bbraun/Dialog Evo | 216630 | Service | Error bicarbonat mixing, Tukar bodypump ke BICP | Ready to use |  | Ari
22000019 | 2026-03-17 | RSUD ODSK Provinsi Sulawesi Utara | Hemodialisa | Bbraun/Dialog Evo | 518858 | Service | Test DF, bersihkan body pump | Ready to use |  | Ari
3000934 | 2026-03-17 | RSUD Dr. R. Soedjati Soemodiarjo | X-ray Digital | Apelem/Camargue DR | 251025003 | Maintenance | General check, fucntion check, table check, generator check. Berfungsi dengan baik semua unit | Ready to use |  | Imam, Amna
07000959 | 2026-03-17 | RSUP dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0109 | Service | Pembacaan O2 error. Penggantian housing asm O2 sensor, O2 sensor, kalibrasi O2 sensor dan post test. Unit OK | Ready to use |  | Wahyu
9000793 | 2026-03-17 | RSUD Dr. Chasan Boesoerie | MOT | SYAS | - | Survey | Rel pintu swing lepas/rusak; Perbaikan rel motor pintu swing | Ready to use |  | Ruslan
19000594 | 2026-03-16 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog | ASLN0005 | Survey | Unit restart, kayar touchscreen tidak sensitif. Cal unit dan baterry check. Unit sedang running | Pending |  | Nanda
3000933 | 2026-03-16 | RSUD Dr. R. Soedjati Soemodiarjo | Digital Radiografi Fluoroskopi | Apelem/Platinum | H248L0156 | Maintenance | General check, fucntion check, table check, generator check. Berfungsi dengan baik semua unit | Ready to use |  | Imam, Amna
6001015 | 2026-03-16 | RSUD Ajibarang | USG | Mindray/DC30 | 9P-18009329 | Service | Penggantian probe convex 3C5P | Ready to use |  | Dwi
19000593 | 2026-03-13 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASLN0005 | SUrvey | Unit restart terus. Device check dan tarik error log. Error log akan dianalisa dahulu | Pending |  | Nanda
8000831 | 2026-03-13 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASHD0034 | SUrvey | Sensor check fail. Hit software, VGC power on self test, calibration gas controller. Running OK | Ready to use |  | Amna
6001014 | 2026-03-13 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASUM0001 | SUrvey | Warmer breathing susyem failure, expiratory flow sensor failure. Tarik error log; Tuelle_T perlu diganti | Pending | Akan dilakukan claim warranty | Dwi
19000592 | 2026-03-12 | RSUD dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | Scanner tidak bisa konek. Setting IP Address. Unit OK | Ready to use |  | Nanda
17000229 | 2026-03-12 | RSUP Dr. Cipto Mangunkusumo | Patient Monitor | Draeger/Vista 120S | S2SNM1733 | Survey | Installasi unit demo | Ready to use |  | Bahar, Nanda, Farhat
21000034 | 2026-03-12 | RSUD Ulin | Laser Surgical | Jena Surgical/Multipulse Hoplus | J191150176 | Survey | Sambungan selang patah; Sambungan selang dibawa ke HO | Pending |  | Habibi
18000727 | 2026-03-12 | RSUD Anuntaloko | Bor tulang | Aesculap/ELAN4 | 2513100 | Instalasi | Unit berufngsi dengan baik | Ready to use |  | Irfan R
12000419 | 2026-03-12 | RSUD Totokabila | Ventilator | Draeger/Evita 300 | ASND0131 | Survey | Tombol rotari knob error. Bersihkan socket dan konektor ke board base. Coba tukar dengan tombol unit lain namun masih error; Indikasi rusak pada board base | Pending |  | Marshall
7000964 | 2026-03-12 | RSUD Merah Putih | Ventilator | Draeger/babylog VN800 | ASTK0008 | Service | Touchscreen tidak responsif. Cek error log, replace LCD Display, reload sfotware, touchscreen OK. (claim warranty) | Ready to use |  | Wahyu
12000412 | 2026-03-11 | RSUD Bolaang Mongondow Utara | X-ray Digital | Apelem/Camargue DR | IMA703918 | Survey | Baterai detector sudah dicharge namun sering drop. Pengecekan tegangan charger dan baterai; Baterai perlu diganti | Pending |  | Marshall
9000792 | 2026-03-11 | RSUD Dr. Chasan Boesoerie | MOT | SYAS | - | Survey | Pintu passbox tidak bisa terbuka. Bersihkan sensor pintu | Ready to use |  | Ruslan
18000727 | 2026-03-11 | RSUD Anuntaloko | Elan4 | Aesculap/ELAN4 | 2513100 | Instalasi | Unit berfungsi dengna baik | Ready to use |  | Irfan R
19000591 | 2026-03-10 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Atlan A300 | ASPH0141 | Survey | Gas measurement fail, etco2 fail, ilca3 fail | Pending | Perlu penggantian ILCA3 dan baterai internal | Nanda
12000418 | 2026-03-10 | RSUD Ainun Habibie | MOT | SYAS | - | Survey | Pengecekan unit indoor dan outdoor HVAC OK5 dan OK6. Prefilter, medium filter dan kondensator sudah kotor | Ready to use |  | Marshall
09000786 | 2026-03-10 | RSUD Dr. Chasan Boesoerie | MOT | SYAS | - | Survey | Unit pendant dan saklar dinding mati; Pengecekan MCB pada trafo isolasi. Unit sudah normal kembali | Ready to use |  | Ruslan
09000785 | 2026-03-10 | RSUD Dr. Chasan Boesoerie | MOT | SYAS | - | Service | Scrubsink bermasalah sehingga air mengalir terus menerus. Pengecekan pada foot pedal valve | Ready to use |  | Ruslan
20000470 | 2026-03-10 | RSUD Sayang Rakyat | Digital Radiologi Fluoroscopy | Apelem/Platinum | M196F0403 | Survey | Hasil gambar artefak. Setting AEC dan cek detektor; Unit perlu diklaibrasi detektor | Ready to use | Unit bisa digunakan dengan keterbatasan | Sujana, Nasrul
20000471 | 2026-03-11 | RS Universitas Hassanudin | Mesin Anasthesia | Draeger/Primus | ASRD0045 | Survey | Inspiration flow fail & breathing leak system; Perlu penggantian service set 3 tahun | Pending |  | Sujana
20000471 | 2026-03-11 | RS Universitas Hassanudin | Mesin Anasthesia | Draeger/Primus | ASRH0039 | Survey | Inspiration flow fail & breathing leak system; Perlu penggantian service set 3 tahun | Pending |  | Sujana
7000966 | 2026-03-09 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASJN0023 | Maintenance | Maintenance Total Care. Running well | Ready to use |  | Wahyu
22000018 | 2026-03-09 | Siloam Hospital Manado | Hemodialisa | Bbraun/Dialog Evo | 214224 | Service | Self test VD failed. Cek valve VD, buka board dan bersihkan | Ready to use |  | Ari
07000966 | 2026-03-09 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN6500 | ASJN0023 | Maintenance | Maintenance Total care. Self test, running OK | Ready to use |  | Wahyu
19000590 | 2026-03-09 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 101625 | Survey | Unit tidak keluar shock. HVG rusak dan EMH expired | Pending |  | Nanda
6001012 | 2026-03-08 | RSUD Hj. Anna Lesmanah | Mesin Anasthesia | Draeger/Primus | ASLD0042 | Maintenance | Maintenance Total Care | Ready to use |  | Dwi
6001013 | 2026-03-08 | RSUD Hj. Anna Lesmanah | Mesin Anasthesia | Draeger/Primus | ASKJ0050 | Survey | Preventive Maintenance Warranty | Ready to use |  | Dwi
6001013 | 2026-03-08 | RSUD Hj. Anna Lesmanah | Mesin Anasthesia | Draeger/Primus | ASJF0043 | Survey | Preventive Maintenance Warranty | Ready to use |  | Dwi
8000823 | 2026-03-06 | RSUD dr. R. Soedjati Soemodiarjo | USG | Mindray/Z50T | HH3E58000152 | Service | 3D/4D software tidak berfungsi. Penggantian module 4D kit | Ready to use |  | Amna
19000588 | 2026-03-06 | RSUP dr. Kariadi | Dental Unit | Kavo/1058 Life | 1002719 | Survey | Tekanan kompresor dibawah 2bar. Ada kebocoran pada dryer. Bypass dari piston ke tanki, kompresor hanya 1.5 bar | Pending | Perlu penggantian kompressor | Nanda
19000587 | 2026-03-06 | RSUP dr. Kariadi | Dental Unit | Kavo/1058 Life | 1002719 | Survey | Handpiece tidak bisa muter. Cek unit. Handpiece dibawa untuk di cek lebih lanjut | Pending |  | Nanda
12000405 | 2026-03-06 | RSUD Totokabila | MOT | SYAS | - | Service | Pencucian indoor dan outdoor unit HVAC | Ready to use |  | Marshall
07000963 | 2026-03-06 | FKG UGM | Dental Simulator | Kavo/DSE Expert | 10001195 | Survey | Pasien simulator tidak bergerak. Cek error log, service mode, setting ulang. Unit OK | Ready to use |  | Wahyu
9000783 | 2026-03-06 | RSD Tidore | Ventilator | Draeger/Evita VN800 | ASTN0001 | Survey | FIO2 error ketika setting O2 70-80% | Pending |  | Ruslan
7000963 | 2026-03-06 | FKG UGM | Dental Simulator | Kavo/DSE Expert | 10001195 | Survey | Patient simulator tidak bisa bergerak. Cek error log, service mode, learn end position. Unit OK | Ready to use |  | Wahyu
18000726 | 2026-03-05 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista120 | S2SRF0243 | Service | NIBP tekanan darah error. Pengecekan manset dan selang tube. Konektor manset longgar. Sudah diperbaiki dan unit berfungsi dengan normal | Ready to use |  | Irfan R
6001011 | 2026-03-05 | RSUD dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASJV0034 | Survey | display mati total. Kabel fleksible dan konektor diperbaiki, unit sudah berfungsi normal | Ready to use |  | Dwi
09000782 | 2026-03-05 | RSUD Dr. Chasan Boesoerie | MOT | SYAS | - | Survey | Scrubsink bermasalah sehingga air mengalir terus menerus. Pengecekan pada foot pedal valve | Ready to use |  | Ruslan
8000822 | 2026-03-05 | RSUD dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASUC0046 | Maintenance | Cleaning filter, check CO2y, flow sensor, sample line and water trap. All OK | Ready to use |  | Amna
12000401 | 2026-03-04 | RSUD Totokabila | MOT | SYAS | - | Survey | Unit outdor dan indoor kotor. Pengecekan filter dan kondensator unit. Perlu dilakukan pencucian dan penggantian filter | Ready to use |  | Marshall
19000585 | 2026-03-04 | Puskesmas Grabag II | USG | Mindray/DP20T | LK7B36001565 | Service | Tidak bisa loading OS. Download OS terbaru, install driver printer. Unit sudah berfungsi normal | Ready to use |  | Nanda
8000821 | 2026-03-04 | RSUD dr. R. Soedjati Soemodiarjo | Digital Radiologi Fluoroscopy | Apelem/Platinum Neo | H248L0156 | Service | Error 245; Kalibrasi table dan setting limit switch. Table OK | Ready to use |  | Amna
6001010 | 2026-03-03 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASHA0095 | Survey | Pressure measurement impaired. O2 air failed. Dilakukan pengecekan unit dengan 3 pcs Flow sensor | Pending |  | Dwi
09000781 | 2026-03-03 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0011 | Survey | Gas supply unit fail. Tarik error; Board E-sat PCB with firmware rusak | Pending |  | Ruslan
07000960 | 2026-03-03 | RSUP dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASRK0016 | Survey | Monitoring pasien error. Cek error log, coba ganti board pvt, ganti valve calibration, dan baterai internal, post test. Baterai internal perlu diganti | Ready to use | Unit bisa dipakai dengan keterbatasan | Wahyu
6001009 | 2026-03-02 | RSUD Brebes | Ventilator | Draeger/Evita V800 | ASPH0003 | Survey | Software error M48. Reload software versi 1.5, masih error. PBA M48 perlu diganti | Pending |  | Dwi
7000958 | 2026-03-02 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog | ASKL0091 | Maintenance | Maintenance Total Care. Penggantian board service cockpit dan reload software cockpit | Ready to use |  | Wahyu
22000017 | 2026-03-02 | RSUD ODSK Provinsi Sulawesi Utara | Hemodialisa | Bbraun/Dialog Evo | 5188847 | Service | Error solution intake disinfectan; Pembersihan board VD dan connectornya | Ready to use |  | Ari
19000586 | 2026-03-02 | RSJ Prof Dr Soerojo Magelang | USG | Mindray/Nuewa 19 Exp | KV3-52000073 | Service | Install unit demo | Ready to use |  | Nanda
21000033 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 389411 | Survey | Pengecekan alat. Motor pump, pengunci spuit dan baterai rusak | Pending |  | Habibi
21000032 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 345160 | Survey | Pengecekan alat. Pengunci spuit sudah retak | Pending |  | Habibi
21000031 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 389393 | Survey | Pengecekan alat. Board pump bermasalah | Pending |  | Habibi
21000031 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 432825 | Survey | Pengecekan alat. Board pump bermasalah | Pending |  | Habibi
21000030 | 2026-03-02 | RSUD Ratu Zalecha | Infus Pump | Bbraun/Infusomat Space P | 409638 | Survey | Pengecekan alat. Baterai bermasalah | Pending |  | Habibi
21000030 | 2026-03-02 | RSUD Ratu Zalecha | Infus Pump | Bbraun/Infusomat Space P | 505897 | Survey | Pengecekan alat. Baterai bermasalah | Pending |  | Habibi
21000030 | 2026-03-02 | RSUD Ratu Zalecha | Infus Pump | Bbraun/Infusomat Space P | 411339 | Survey | Pengecekan alat. Baterai bermasalah | Pending |  | Habibi
21000029 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 345155 | Survey | Pengecekan alat. Baterai bermasalah | Pending |  | Habibi
21000028 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 389512 | Survey | Pengecekan alat. Baterai bermasalah | Pending |  | Habibi
21000025 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 432800 | Survey | Pengecekan alat. Keypad bermasalah | Pending |  | Habibi
21000024 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 432839 | Survey | Pengecekan alat. Perlu penggantian spuit dan display | Pending |  | Habibi
21000023 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 398095 | Survey | Pengecekan alat. Pengunci spuit rusak, display dan baterai juga rusak | Pending |  | Habibi
21000022 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 345189 | Survey | Pengecekan alat. Display rusak | Pending |  | Habibi
21000021 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 432826 | Survey | Pengecekan alat, baterai rusak | Pending |  | Habibi
21000020 | 2026-03-02 | RSUD Ratu Zalecha | Syringe Pump | Bbraun/Perfusor Space | 479336 | Survey | Alat mati total, pengecekan akat; Motor pumo rusak, syringe holder patah, baterai rusak | Pending |  | Habibi
07000958 | 2026-03-02 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger | ASKL0091 | Maintenance | Maintenance Total Care. Reload software cockpit, self test, running OK | Ready to use |  | Wahyu
9000776 | 2026-02-27 | RSD Tidore | Ventilator | Draeger/Babylog VN800 | ASTN0001 | Survey | FIO2 low. Cek kemurnian O2 . Disarankan menggunakan O2 dengan purity 99% | Ready to use |  | Ruslan
9000775 | 2026-02-27 | RSD Tidore | Ventilator | Draeger/Babylog VN500 | ASJL0001 | Survey | Nebulizer fail, battery drop. Device check; Pressure regulator perlu diganti | Pending |  | Ruslan
9000774 | 2026-02-27 | RSD Tidore | Ventilator | Draeger/Babylog VN500 | ASJK0042 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Ruslan
11000439 | 2026-02-27 | RSUD Tenriawaru Bone | Hemodialisa | Bbraun | 242701 | Service | Alarm disinfektan valve open; Pengecekan board disinfektan valve, koneksi kabel board. Unit OK | Ready to use |  | Nasrul
12000409 | 2026-02-26 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Service | Printer error 632 (Heating disconnection). Reset thermal heater detector. Rest unit; Unit sudah tidak error | Ready to use |  | Marshall
17000228 | 2026-02-26 | RSAL Dr. Mintohardjo | Laser Surgical | Jena Surgical/Multipulse Hoplus | J191150378 | Survey | Suara kasar saat mesin beroperasi. Menambahkan aquades 1L. Suara kasar dari radiator aktif | Ready to use |  | Bahar
19000584 | 2026-02-26 | Puskesmas Grabag II | USG | Mindray/DP20T | LK7B376001565 | Service | Unit tidak bisa print. Install ulang OS, fail. Install ulang recovery system, fail. Unit belum bisa digunakan, akan di follow up kembali ke Mindray | Pending |  | Nanda
9000773 | 2026-02-26 | RSUD Weda | Mesin Anasthesia | Draeger/Primus | ASHK0159 | Survey | Ventilator piston fail, nebulizer fail; Perlu penggantian motor piston, flow poer, baterai dan iLCA 2 | Pending |  | Nasrul
9000770 | 2026-02-26 | RSUD Weda | Baby bed | SYAS | 260126211000400001 | Instalasi | Unit berfungsi dengan baik. Installasi 5 unit | Ready to use |  | Ruslan
8000820 | 2026-02-26 | RS Assuyuthiyyah | Mesin Anasthesia | Draeger/Fabius Plus | ASUA0003 | Maintenance | Maintenance unit | Ready to use |  | Amna
9000769 | 2026-02-26 | RSUD Weda | Examination Table | SYAS | 26016410100040001 | Instalasi | Installasi 10 unit. Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000771 | 2026-02-26 | RSUD Weda | Foot step | SYAS | 2601613000400001 | Instalasi | Installasi 10 unit. Unit berfungsi dengan baik | Ready to use |  | Ruslan
70000953 | 2026-02-26 | RSUD Panembahan Senopati | ESWL | Medispec/E3000 | 0719139 | Survey | air meluap dari mesin. Pengecekan pressure sensor, pressure sensor OK | Ready to use |  | Wahyu
70000954 | 2026-02-26 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASJH0121 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Wahyu
18000736 | 2026-02-26 | RS GMIM Pancaran Kasih | Ventilator | Draeger/Savina 300 | ASTJ0019 | Survey | System check tidak berhasil. Breathing circuit bocor; Penggantian breathing circuit dan kalibrasi alat, system check OK | Ready to use |  | Irfan
14000408 | 2026-02-25 | RSUP Surabaya | C-arm | Apelem/Evo R+ | 551240231085 | Survey | Hasil gambar saat mode operasi gelap, ada missing pulse dan tidak bisa expose. Cek sambungan monitor c-arm dan board serta kamera, baik. Unit masih belum bisa menunggu remote dari apelem | Pending |  | Fahim
22000016 | 2026-02-25 | RS GMIM Bethesda | Hemodialisa | Bbraun | 504499 | Service | DF pressure 12550 membran. Penggantian membran | Ready to use |  | Ari
6001008 | 2026-02-25 | RSUD Prembun | Bubble CPAP | Draeger/Oxymixer | QM20230153 | Service | Penggantian sparepart overhaul dan low flow meter | Ready to use |  | Dwi
6001008 | 2026-02-25 | RSUD Prembun | Bubble CPAP | Draeger/Oxymixer | QM20230154 | Service | Penggantian sparepart overhaul dan low flow meter | Ready to use |  | Dwi
7000952 | 2026-02-25 | RSUD Merah Putih | Jaudince Meter | Draeger/JM-105 | H3510375 | Service | Error 05; Penggantian internal battery | Ready to use |  | Wahyu
18000739 | 2026-02-25 | RSUD Pratama Bitung | Mobile X-ray | Apelem/Rafale | 8112416337 | Survey | Pengecekan unit karna belum diinstall sejak dibeli. Belum bisa test dengan listrik karna listrik rumah sakit tidak stabil | Pending |  | Irfan
7000951 | 2026-02-24 | RSUD Sleman | Ventilator | Draeger/Evita V800 | ASUL0018 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Wahyu
14000409 | 2026-02-23 | RSUP Surabaya | Meja Operasi | SYAS/OP850 | OP850XVZY00003 | Service | Meja bagian belakang turun. Cek pergerakan meja, cek motor hidrolic, valve dan setting valve. Unit masih belum bisa, akan diajukan claim warranty | Pending |  | Fahim
19000577 | 2026-02-23 | RS JIH Solo | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Survey | O2 sensor fail. Cek unit, cleaning, setting; Unit bisa digunakan namun disarankan mengganti service set 6 tahun | Ready to use |  | Nanda
19000580 | 2026-02-23 | RS JIH Solo | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0125 | Maintenance | Maintenance unit | Ready to use |  | Nanda
6001007 | 2026-02-23 | RS PKU Muhammadiyah Siti Aminah | Ventilator | Draeger/Fabius Plus | ASND0060 | Survey | Leaking system dan O2 failure. Pressure Cozy breathing expired | Pending |  | Dwi
22000015 | 2026-02-23 | Siloam Hospital Manado | Hemodialisa | Bbraun | 219216 | Service | Unit mati total. Penggantian SMPS (power supply) | Ready to use |  | Ari
22000014 | 2026-02-23 | Siloam Hospital Manado | Hemodialisa | Bbraun | 313796 | Service | Temperature test not OK. Solder ulang sensor temperature | Ready to use |  | Ari
9000778 | 2026-02-23 | RSUD Dr. H. Chasan Boesoirie | Laparascopy | Bbraun | 2510700 | Service | Packing unit untuk di kirim ke Bbraun | Ready to use |  | Ruslan
8000819 | 2026-02-23 | RS Sebening Kasih | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ0082 | Survey | O2 sensor expired | Pending |  | Amna
21000027 | 2026-02-23 | RSUD H. Boejasin | Infus Pump | Bbraun/Infusomat Space P | 443957 | Survey | Pengecekan alat. Baterai rusak | Pending |  | Habibi
21000026 | 2026-02-23 | RSUD H. Boejasin | Infus Pump | Bbraun/Infusomat Space P | 305163 | Survey | Unit mati lama tidak digunakan | Ready to use |  | Habibi
210000019 | 2026-02-23 | RSUD H. Boejasin | Syringe Pump | Bbraun | 376792 | Survey | layar mati sendiri dan keypad respon lama. Perlu penggantian display dan keypad | Pending |  | Habibi
21000005 | 2026-02-23 | RSUD H. Boejasin | MOT | SYAS | - | Survey | Pengecekan HVAC karna ruangan panas. HVAC rusak | Pending |  | Habibi
21000004 | 2026-02-23 | RSUD H. Boejasin | ESU | Bbraun | 003475 | Survey | Pengecekan alat karna sudah lama tidak digunakan. Handpiece perlu diperbaiki | Pending |  | Habibi
7000950 | 2026-02-23 | RSUD Prambanan | Ventilator | Draeger/Babylog VN800 | ASPF0050 | Survey | Humidifier kondensasi. Mechanical test, electrical case, internal battery test. Baterai internal rusak | Pending |  | Wahyu
18000738 | 2026-02-23 | RS Mata Manado | Mesin Anasthesia | Draeger/Perseus A500 | ASJM0178 | Survey | Baterai tidak mengisi. Cek tegangan baterai, normal. Test charging 1 hari | Pending |  | Irfan
12000407 | 2026-02-20 | RS Mata Manado | Mesin Anasthesia | Draeger/Perseus A500 | ASJM0178 | Survey | Power supply error. Tarik error unit; Error log dikirim untuk dibaca principle | Pending |  | Marshall
19000576 | 2026-02-20 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0125 | Survey | M540 not connected. Cleaning sensor, running, not OK; M540 rusak | Pending |  | Nanda
8000818 | 2026-02-20 | RS Budi Agung Juwana | Mesin Anasthesia | Draeger/Fabius Plus | ASTK0003 | Survey | O2 sensor expired | Pending |  | Amna
6001006 | 2026-02-20 | RSUD Brebes | Ventilator | Draeger/Evita V800 | ASPH0003 | Service | Unit mati total; Pemasangan electrical energy unit, baterai pack M7 | Pending | Unit masih error | Dwi
20000466 | 2026-02-19 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVZU000068 | Service | LCD kontrol mati; Penggantian lcd kontrol (warranty) | Ready to use |  | Sujana
8000817 | 2026-02-19 | RSJD Amino Gondohutomo | Mesin Anasthesia | Draeger/Atlan A300 | ASTF0110 | Survey | Ventilator fail. Penggantian kit ventilator assembly TSB 13. Unit masih belum bisa karna pressure sensor fail | Pending |  | Amna
6001000 | 2026-02-19 | RSUD Kardinah | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0082 | Survey | Gas sensor failure, measuring failure; Flow power dan mini modul perlu diganti | Pending |  | Dwi
18000732 | 2026-02-19 | RSUD ODSK Provinsi Sulawesi Utara | Incubator Transport | Draeger/Babyleo TN500 | ASTH0041 | Survey | Bed tidak bisa naik turun, pengecekan keypad dan motor penggerak bed; Motor south rusak | Pending |  | Irfan
20000464 | 2026-02-18 | RS Regional dr. Hasri Ainun Habibie | Bubble CPAP | Draeger/Oxymixer | QM20246035 | Service | Error mixer. Penggantian part. Unit OK | Ready to use |  | Sujana
11000435 | 2026-02-18 | RSUD Lakipadada | Ventilator | Draeger/Babylog VN800 | ASTE0042 | Survey | FIO2 measurement error. System test, hit software; Pressure regulator perlu diganti | Pending |  | Nasrul
6001005 | 2026-02-18 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Fabius Plus | ASTH008 | Survey | Refresh training | Ready to use |  | Dwi
11000436 | 2026-02-17 | RSUD Tenriawaru Bone | Mesin Anasthesia | Draeger/Primus | ASJM0022 | Survey | O2 sensor fail, gas supply fail | Pending |  | Nasrul
11000436 | 2026-02-17 | RSUD Tenriawaru Bone | Mesin Anasthesia | Draeger/Primus | ASNH0255 | Survey | Ventilator piston fail, nebulizer fail; Perlu penggantian motor piston, flow poer, baterai dan iLCA 2 | Pending |  | Nasrul
5000410 | 2026-02-16 | L&M Dental Care | Dental Unit | Planmeca/Compact i | UTTF273272 | Survey | Survey fisik dan fungsi alat. Gagang lampu patah, touchsreen dents rusak | Pending |  | Farhat
5000409 | 2026-02-16 | L&M Dental Care | Dental Unit | Planmeca/Compact i | UCIV172353 | Survey | Survey fisik dan fungsi alat.Ada bocor pada dentist element | Pending |  | Farhat
5000408 | 2026-02-16 | L&M Dental Care | Dental Unit | Planmeca/Compact i touch V2 | UTTF273273 | Survey | Survey fisik dan fungsi alat. Jok stool sobek, air bowl bocor netes, lampu turun akrna kendor, saliva patah ujung, selang threeway koyak | Pending |  | Farhat
9000768 | 2026-02-16 | RSUD Sanana | Bed Patient | SYAS/Premium 3 crank | - | Service | Pemindahan paket SYAS premium ke gedung baru | Ready to use |  | Ruslan
18000744 | 2026-02-11 | RSUD ODSK Provinsi Sulawesi Utara | Mesin Anasthesia | Draeger/Atlan XL0115 | ASRH0115 | Service | Flow sensor failure; Penggantian flow sensor | Ready to use |  | Irfan
9000766 | 2026-02-15 | RSUD Sanana | Mesin Anasthesia | Draeger/Atlan A300 | ASSD0233 | Survey | Pengecekan unit. Unit normal | Ready to use |  | Ruslan
9000765 | 2026-02-15 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD0300 | Survey | SPO2 tidak terbaca dan tekanan NIBP tidak stabil. Pengecekan sensor SPO2 dan NIBP Cuff; Cuff bocor dan sensor rusak | Pending |  | Ruslan
3000930 | 2026-02-15 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Meja tilting tidak bisa bergerak. Perbaikan meja, perbaikan pergerakan transversal, perbaikan monitor grid. Unit masih error | Pending |  | Imam, bahar
8000816 | 2026-02-14 | RS Bhina Bakti Husada | Mesin Anasthesia | Draeger/Atlan A300 XL | ASMM0069 | Service | Packing demo unit | Ready to use |  | Amna
11000438 | 2026-02-13 | RSUD Tenriawaru Bone | Hemodialisa | Bbraun | 264404 | Service | Alarm disinfektan valve open; Penggantian board disinfektan valve | Ready to use |  | Nasrul
9000767 | 2026-02-13 | RSUD Sanana | X-ray Digital | Apelem/Camargue DR | 251023012 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Ruslan
| 2026-02-13 | Signature Dental care | Dental Unit | Planmeca/ Compact i5 | UPCU574532 | Service | Pemasangan separator backup | Ready to use |  | Ipan
16000310 | 2026-02-13 | FKG Dr. Moestopo | Dental Simulator | Kavo/Compact | 00000850 | Survey | Bocor angin pada botol air bersih; Pengecekan oring pada air bersih, oring rusak/getas | Pending |  | Ipan
16000310 | 2026-02-13 | FKG Dr. Moestopo | Dental Simulator | Kavo/Compact | 00000809 | Survey | Bocor angin pada botol air bersih; Pengecekan oring pada air bersih, oring rusak/getas | Pending |  | Ipan
16000310 | 2026-02-13 | FKG Dr. Moestopo | Dental Simulator | Kavo/Compact | 00000823 | Survey | Bocor angin pada botol air bersih; Pengecekan oring pada air bersih, oring rusak/getas | Pending |  | Ipan
16000310 | 2026-02-13 | FKG Dr. Moestopo | Dental Simulator | Kavo/Compact | 00000823 | Survey | Bocor angin pada botol air bersih; Pengecekan oring pada air bersih, oring rusak/getas | Pending |  | Ipan
16000310 | 2026-02-13 | FKG Dr. Moestopo | Dental Simulator | Kavo/Compact | 00000841 | Survey | Bocor angin pada botol air bersih; Pengecekan oring pada air bersih, oring rusak/getas | Pending |  | Ipan
16000310 | 2026-02-13 | FKG Dr. Moestopo | Dental Simulator | Kavo/Compact | 00000845 | Survey | Bocor angin pada botol air bersih; Pengecekan oring pada air bersih, oring rusak/getas | Pending |  | Ipan
16000309 | 2026-02-13 | FKG Dr. Moestopo | Dental Simulator | Kavo/Compact | 00000865 | Survey | Bocor angin pada dental simulator. Pengecekan jalur kebocoran, bocor pada selang 2mm dan 6mm | Pending |  | Ipan
16000309 | 2026-02-13 | FKG Dr. Moestopo | Dental Simulator | Kavo/Compact | 00000863 | Survey | Bocor angin pada dental simulator. Pengecekan jalur kebocoran, bocor pada selang 2mm dan 6mm | Pending |  | Ipan
16000309 | 2026-02-13 | FKG Dr. Moestopo | Dental Simulator | Kavo/Compact | 00000859 | Survey | Bocor angin pada dental simulator. Pengecekan jalur kebocoran, bocor pada selang 2mm dan 6mm | Pending |  | Ipan
5000407 | 2026-02-13 | L&M Dental Care | Dental Unit | Planmeca/Compact i | VTIF273272 | Survey | Threeway selang rusak, touch pada layar dentist bermasalah | Pending |  | Farhat
8000815 | 2026-02-13 | RSUD RAA Soewondo Pati | Infant Warmer | Draeger/Resuscitaire | KC13741 | Survey | Error 05 dan 5way gas fail. Heater oke namun perlu penggantian 5way | Pending |  | Amna
8000815 | 2026-02-13 | RSUD RAA Soewondo Pati | Infant Warmer | Draeger/Resuscitaire | KC13738 | Survey | Error 05 dan 5way gas fail. 5 way dan heater rusak | Pending |  | Amna
8000815 | 2026-02-13 | RSUD RAA Soewondo Pati | Infant Warmer | Draeger/Resuscitaire | AC13489 | Survey | Error 05 dan 5way gas fail. Heater oke namun perlu penggantian 5way | Pending |  | Amna
7000957 | 2026-02-13 | RSUD Merah Putih | Ventilator | Draeger/Babylog VN800 | ASTK0008 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Wahyu
21000017 | 2026-02-12 | RSUD H. Andi Abdurrahman Noor | Infus Pump | Bbraun/Infusomat Space P | 458851 | Survey | Drive motor ISP, frame ISP dan baterai perlu diganti | Pending |  | Habibi
21000017 | 2026-02-12 | RSUD H. Andi Abdurrahman Noor | Infus Pump | Bbraun/Infusomat Space P | 456710 | Survey | Baterai rusak | Pending |  | Habibi
21000015 | 2026-02-12 | RSUD H. Andi Abdurrahman Noor | Syringe Pump | Bbraun/Perfusor Space | 389514 | Survey | Baterai rusak, pengunci spuit patah | Pending |  | Habibi
21000014 | 2026-02-12 | RSUD H. Andi Abdurrahman Noor | Syringe Pump | Bbraun/Perfusor Space | 389507 | Survey | Baterai rusak | Pending |  | Habibi
8000814 | 2026-02-12 | RSJD Amino Gondohutomo | Mesin Anasthesia | Draeger/Atlan A300 | ASTF0110 | Survey | Ventilator fail; Piston perlu diganti | Pending |  | Amna
18000743 | 2026-02-12 | RSUD Noongan | Mesin Anasthesia | Draeger/Perseus A500 | ASJF0093 | Survey | Device check fail, O2 supply failure, heater failure; Heater, battery pack dan VMGS part rusak | Pending |  | Irfan
18000742 | 2026-02-12 | RSUD Noongan | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0066 | Survey | Power supply fail, battery fail; Battery dan PSU M7 rusak | Pending |  | Irfan
18000741 | 2026-02-12 | RSUD Noongan | Mesin Anasthesia | Draeger/Perseus A500 | ASJL0006 | Survey | Unit mati total. Board M.1.6 rusak dan flow power rusak. | Pending |  | Irfan
19000575 | 2026-02-12 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Draeger/Fabius Plus | ASHD0007 | Survey | O2 flow power tidak bisa naik turun. Setting knop naik turun flow meter O2, cleaning air pressure yang kemasukan air; Unit bisa digunakan namun harus dikasih watertrap | Ready to use |  | Nanda
20000460 | 2026-02-12 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Survey | Gearbox pecah. Unit perlu penggantian gearbox dan water chamber | Pending |  | Sujana
19000573 | 2026-02-11 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0087 | Survey | Konsentrasi O2 terlalu tinggi. Cek unit, tarik error, ada indikasi PATO Fault, kalibrasi, OK. Unit sementara sedang di running | Pending |  | Nanda
21000018 | 2026-02-11 | RSUD H. Andi Abdurrahman Noor | Infus Pump | Bbraun/Infusomat Space P | 458840 | Survey | Baterai rusak | Pending |  | Habibi
21000013 | 2026-02-11 | RSUD H. Andi Abdurrahman Noor | Syringe Pump | Bbraun/Perfusor Space | 393016 | Survey | Baterai rusak | Pending |  | Habibi
21000012 | 2026-02-11 | RSUD H. Andi Abdurrahman Noor | Infus Pump | Bbraun/Infusomat Space P | 45884 | Survey | Layar gelap dan tombol tidak responsif. Display dan tombol rusak | Pending |  | Habibi
21000011 | 2026-02-11 | RSUD H. Andi Abdurrahman Noor | Infus Pump | Bbraun/Infusomat Space P | 592217 | Survey | Baterai rusak | Pending |  | Habibi
8000813 | 2026-02-11 | RS PKU Fastabiq Sehat Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASTH0060 | Service | Penggantian A-cone switch |  |  | Amna
14000404 | 2026-02-11 | RSUP Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum | G247L0146 | Survey | Terdapat bocor pada console. Cek settingan, restart unit dan cek pergerakan table. Kalibrasi table. Unit masih belum bisa | Pending |  | Fahim
20000455 | 2025-02-10 | RSUP Makassar | MOT | SYAS | - | Service | Penggantian mini pc | Ready to use |  | Sujana
21000010 | 2026-02-10 | RSUD H. Andi Abdurrahman Noor | Syringe Pump | Bbraun/Perfusor Space | 388505 | Survey | Alat mati; Driver head syringe, pengunci spuit dan baterai rusak | Pending |  | Habibi
3000929 | 2026-02-10 | RSUD Gabriel Manek | Digital Radiografi Fluoroskopi | Apelem/Platinum | E226F0723 | Service | Detector muncul artefak. Penggantian detector | Ready to use |  | Imam, bahar
21000009 | 2026-02-10 | RSUD H. Andi Abdurrahman Noor | Syringe Pump | Bbraun/Perfusor Space | 383985 | Survey | Unit mati. driver head rusak | Pending |  | Habibi
21000008 | 2026-02-10 | RSUD H. Andi Abdurrahman Noor | Infus Pump | Bbraun/Infusomat Space P | 508304 | Survey | Mati unit; Baterai rusak | Pending |  | Habibi
8000812 | 2026-02-10 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Atlan A300 XL | ASTJ0118 | Survey | Refresh training | Ready to use |  | Amna
9000779 | 2026-02-10 | RSUD Dr. H. Chasan Boesoirie | MOT | SYAS | - | Survey | Pipa freon unit outdoor HVAC membeku. Cek tekanan lowpressure. Sementara unit normal dan akan dilakukan pengisian freon jika sudah dikirim oleh tim KSS | Ready to use |  | Ruslan
6001002 | 2026-02-10 | RS Raffa Medika | Ventilator | Draeger/Babylog VN600 | ASTE0067 | Service | Penggantian kompressor | Ready to use |  | Dwi
19000572 | 2026-02-09 | RSUD dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Atlan A300 | ASPH0141 | Survey | Gas measurement fail. Cek unit, coba ganti motor pump, unit masih belum bisa. Tarik error log | Pending |  | Nanda
7000956 | 2026-02-09 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASLL0067 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Wahyu
21000007 | 2026-02-09 | RSUD H. Boejasin | Syringe Pump | Bbraun/Perfusor Space | 443953 | Survey | Pengunci spuit patah | Pending |  | Habibi
21000007 | 2026-02-09 | RSUD H. Boejasin | Infus Pump | Bbraun/Infusomat Space | 361428 | Survey | Baterai rusak | Pending |  | Habibi
5000406 | 2026-02-09 | MP Dental | Dental Unit | Planmeca/ Compact i | UCIG191427 | Survey | Bau amis pada pembuangan; Menutup celah dan cleaning suction pada jalur pembuangan | Ready to use |  | Farhat, Ipan
8000810 | 2026-02-09 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0115 | Service | Penggantian O2 sensor | Ready to use |  | Amna
8000810 | 2026-02-09 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0148 | Service | Penggantian O2 sensor | Ready to use |  | Amna
8000810 | 2026-02-09 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0049 | Service | Penggantian O2 sensor | Ready to use |  | Amna
8000810 | 2026-02-09 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASSC0025 | Service | Penggantian O2 sensor | Ready to use |  | Amna
22000013 | 2026-02-09 | Siloam Hospital Manado | Hemodialisa | Bbraun | 214217 | Service | Mesin bocor saat disinfektan; Perbaikan pelampung sensor NSVB | Ready to use |  | Ari
11000437 | 2026-02-09 | RSUP Dr. Wahidin Sudirohusodo | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/087/C1 | Service | Daily check fail. Cek suhu detektor, cek tegangan, cek prosedur daily check | Ready to use |  | Nasrul, Sujana
6001001 | 2026-02-09 | RSU An Ni'mah | Bubble CPAP | Draeger/Oxymixer | QM20255060 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Dwi
6001001 | 2026-02-09 | RSU An Ni'mah | Bubble CPAP | Draeger/Oxymixer | QM20255074 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Dwi
22000011 | 2026-02-07 | RSUD Dok II Jayapura | Hemodialisa | Bbraun | 208469 | Service | Temperature test not OK. Pembersihan dan kalibrasi sensor temperature | Ready to use |  | Ari
9000760 | 2026-02-07 | RSUD Dr. H. Chasan Boesoirie | Laparascopy | Bbraun | 2510700 | Survey | Pendampingan operasi | Ready to use |  | Ruslan
9000759 | 2026-02-07 | RSUD Dr. H. Chasan Boesoirie | MOT | SYAS | - | Survey | Alarm low pressure. Cek seluruh setting dan kebocoran pada unit outdoor. Cek tekanan lowpress dan highpress. Unit normal | Ready to use |  | Ruslan
22000010 | 2026-02-07 | RSUD Dok II Jayapura | Hemodialisa | Bbraun | 522717 | Service | Temperature test not OK. Pembersihan dan kalibrasi sensor temperature | Ready to use |  | Ari
18000731 | 2026-02-06 | RS GMIM Bethesda | Ventilator | Draeger/Savina 300 | ASTJ0025 | Survey | Kalibrasi valve gagal. Cek valve expirasi, membran pecah; Penggantian membran valve | Ready to use |  | Irfan
22000009 | 2026-02-06 | RSUD Dok II Jayapura | Hemodialisa | Bbraun | 522708 | Service | DF pressure 12350; Penggantian body pump UF, self test. OK | Ready to use |  | Ari
22000008 | 2026-02-06 | RSUD Dok II Jayapura | Hemodialisa | Bbraun | 522690 | Service | DF pressure intermitten; Penggantian body pump dab kalibrasi PDA. Self test, OK | Ready to use |  | Ari
22000007 | 2026-02-06 | RSUD Dok II Jayapura | Hemodialisa | Bbraun | 222978 | Service | DF pressure 12350; Penggantian body pump dan filter | Ready to use |  | Ari
5000405 | 2026-02-06 | Kusumah Dental Care | Dental Unit | Kavo/E30 | 10111157 | Survey | Pengecekan unit. Unit normal | Ready to use |  | Farhat, Ipan
3000928 | 2026-02-06 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | ET130. Perbaikan koneksi DRV lifter | Ready to use |  | Imam, bahar
7000941 | 2026-02-06 | RSUD Bagas Waras | Infant Warmer | Draeger/Caleo | ASJD0035 | Survey | Cover control display damage; Pengecekan kondisi cover | Pending |  | Wahyu
20000451 | 2026-02-05 | RSUP Makassar | MOT | SYAS | - | Service | HVAC panas; Penggantian inverter board | Ready to use |  | Sujana
18000745 | 2026-02-05 | RS Siloam Manado | Hemodialisa | Bbraun/HD | 227271 | Survey | Temperature not OK. Setelah di cek dan dibersihkan sensor TSDE_S temperature masih error. | Pending | Sensor TSDE_S perlu diganti | Marshall
7000955 | 2026-02-05 | RSUD Wonosari | Ventilator Transport | Draeger/Oxylog 2000 | ASLE0013 | Survey | VT tidak stabil; flow sensor rusak | Pending |  | Ari
7000955 | 2026-02-05 | RSUD Wonosari | Ventilator Transport | Draeger/Oxylog 2000 | ASLE0014 | Survey | VT tidak stabil; flow sensor rusak | Pending |  | Ari
22000012 | 2026-02-05 | RSUD Dok II Jayapura | Hemodialisa | Bbraun | 522714 | Service | Temperature too high. Cek PDA, penggantian part PDA dan kalibrasi | Ready to use |  | Ari
17000227 | 2026-02-05 | RS UNPAD | Infus Pump | Bbraun/Infusomat P | 001221 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Bahar, Imam
8000809 | 2026-02-05 | RSUD Dr. R. Soedjati Soemodiarjo | X-ray Digital | Apelem/Camargue DR | 251025003 | Service | Printer fujifilm error, tersangkut. Reposisi film | Ready to use |  | Amna
8000808 | 2026-02-05 | RSUD Dr. R. Soedjati Soemodiarjo | USG | Mindray/Z50T | HH3E58000152 | Survey | 4D tidak terdeteksi; Reinstall doppler recovery, 4D masih tidak bisa. | Pending | Pengajuan claim warranty 4D module | Amna
9000754 | 2026-02-05 | RSUD Jailolo | Dental Unit | Kavo/E30 | 10117281 | Service | Pemasangan board scaller dan scaller tube. Unit sudah berfungsi dengan normal | Ready to use |  | Ruslan
20000451 | 2026-02-05 | RSUP Makassar | MOT | SYAS | - | Service | Penggantian inverter board HVAC. Unit sudah berfungsi normal | Ready to use |  | Sujana
9000752 | 2026-02-05 | RSUD Jailolo | UPS | OAXAKA/AP3330 | 40045034F65393600001 | Instalasi | Unit sudah diinstall namun belum dikoneksikan dengan kWh meter. Namun uji fungsi dengan battery berfungsi dengan baik | Ready to use |  | Ruslan
18000730 | 2026-02-04 | RSUD ODSK Provinsi Sulawesi Utara | Ventilator | Draeger/Babylog VN500 | ASTH0036 | Service | Flow error, cek flow sensor. Penggantian flow sensor | Ready to use |  | Irfan
6000998 | 2026-02-04 | RSUD Ir. Soekarno | USG | Mindray/DC30 | JM3-48006338 | Service | Baterai tidak charging. Penggantian baterai pack DC-30 | Ready to use |  | Dwi
11000434 | 2026-02-04 | RS Siloam Makassar | Hemodialisa | Bbraun | 036392 | Service | Temperature high; Perbaikan sensor temperature | Ready to use |  | Nasrul, Sujana
11000433 | 2026-02-04 | RS Siloam Makassar | Hemodialisa | Bbraun | 229045 | Service | motor piston macet; perbaikan motor piston | Ready to use |  | Nasrul, Sujana
16000308 | 2026-02-04 | MP Dental | Dental Unit | Planmeca/Compact i | UCIG191427 | Service | Suction kurang kuat dan bunyi kencang; Pasang backup separator dan sedot air pada jalur pipa | Ready to use |  | Farhat, Ipan
200000452 | 2026-02-03 | RSUP Makassar | Pendant | SYAS | - | Service | Breaket pendant lepas. Pemasangan breaket pendant | Ready to use |  | Sujana
14000402 | 2026-02-03 | RSUP Surabaya | Meja Operasi | SYAS/OP850 | OP850XVZY00007 | Survey | Oli bocor. Cek fisik bagian atas meja oli bocor, bongkar atas meja namun kendala kunci untuk buka bagian atas tidak ada | Pending | Akan dijadwalkan pengecekan ulang | Fahim
20000453 | 2026-02-02 | RSUP Makassar | MOT | SYAS | - | Service | Passbox error tidak bisa terbuka. Pengecekan sensor dan setting penempatan passbox. Unit sudah berfungsi normal | Ready to use |  | Sujana
12000406 | 2026-02-03 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1545 | Survey | Baterai tidak mengisi. Cek konektor baterai, ada yang longgar; Bersihkan konektor dan rapatkan konektor | Ready to use |  | Marshall
18000746 | 2026-02-03 |  | Patient Monitor | Draeger/Vista 120 | SBSNF1473 | Service | Pemasangan baterai temperature | Ready to use |  | Irfan
18000746 | 2026-02-03 |  | Patient Monitor | Draeger/Vista 120 | SBSNF1500 | Service | Pemasangan baterai temperature | Ready to use |  | Irfan
18000746 | 2026-02-03 |  | Patient Monitor | Draeger/Vista 120 | SBSNF1479 | Service | Pemasangan baterai temperature | Ready to use |  | Irfan
18000746 | 2026-02-03 |  | Patient Monitor | Draeger/Vista 120 | SBSNF1470 | Service | Pemasangan baterai temperature | Ready to use |  | Irfan
9000700 | 2026-02-03 | RSUD Tobelo | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/088/C2 | Service | Koneksikan unit dengan PACS RS | Ready to use |  | Ruslan
9000699 | 2026-02-03 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0012 | Survey | Gas supply unit fail; Penggantian blower unit, G-set PCB dan firmware GS500 | Ready to use |  | Ruslan
6000997 | 2026-02-03 | RS Dadi Keluarga | Ventilator | Draeger/Babylog VN800 | ASPF0030 | Survey | Air supply kalibrasi gagal, baterai tidak bisa charge; Baterai internal rusak dan PIA2 sensor perlu diganti | Pending |  | Dwi
14000402 | 2026-02-03 | RSUP Surabaya | Meja Operasi | SYAS/OP850 | OP850XVZY00007 | Survey | Oli bocor. Cek fisik atas meja oli bocor. Tapi ada kendala kunci bagian hidrolic susah | Pending |  | Fahim
600099 | 2026-02-02 | RSUD Cilacap | Mesin Anasthesia | Draeger/Perseus A500 | ASUJ0001 | Service | Pressure control vail, error system test; Penggantian breathing circuit dan soda lime | Ready to use |  | Dwi
9000753 | 2026-02-02 | RSUD Dr. H. Chasan Boesoirie | Laparascopy | Bbraun | - | Service | Koneksikan semua unit ke trolly unit yang baru | Ready to use |  | Ruslan
22000005 | 2026-02-02 | Siloam Hospital manado | Hemodialisa | Bbraun | 246381 | Service | Bicarbonat dan maxing unit wrong direction. Cek valve dan tukar body pump bicp | Ready to use |  | Ari
19000571 | 2026-02-01 | RSUP dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0235 | Survey | O2 fail, gas measurement fail. Cek unit, tarik error, ada indikasi flow power rusak | Pending |  | Nanda
6000996 | 2026-01-31 | RSUD Ajibarang | Mesin Anasthesia | Draeger/Primus | ASPM0019 | Service | Leak breathing circuit, ETCO2 not detected. Tarik error. Penggantian breathing circuit, waterlock sample line | Ready to use |  | Dwi
19000570 | 2026-01-30 | RSUD Mijen | Baby Incubator | Draeger/Isolette 8000 | HC09150 | Survey | Sensor failure dan low temperature skin. Cek sensor module, slide in sensor module. Unit sudah OK | Ready to use |  | Nanda
8000807 | 2026-01-30 | RSUD Kab. Temanggung | Infant Warmer | Draeger/Resuscitaire | JY10905 | Survey | Error 21, test diagnostic 10, check ambient temperature, calibration temprerature. Unit OK | Ready to use |  | Amna
14000401 | 2026-01-30 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Service | HVG rusak. Penggantian HVG backup | Ready to use |  | Fahim
17000226 | 2026-01-30 | RSK Dharmais | Digital Radiografi Fluoroskopi | Apelem/Platinum | L216F0664 | Service | Kolimator tidak sesuai; Setting kolimator | Ready to use |  | Bahar
19000569 | 2026-01-29 | RSUD Pandan Arang | Incubator Transport | Draeger/TI500 | BH12345 | Survey | Baterai internal drop, baterai rusak. Harus menggantia baterai | Pending |  | Nanda
19000568 | 2026-01-29 | RSUD Pandan Arang | Infant Warmer | Draeger/Resuscitaire | CF10509 | Survey | Error 05. Cek alat, cek heater, cek board SSR. Sementara heater disambung namun tetap harus diganti heater | Ready to use |  | Nanda
21000006 | 2026-01-29 | RSUD Sultan Suriansyah | Dental Unit | Kavo/1058 Life | 10205321 | Survey | Lampu mati; Pengecekan alat, perbaikan. Lampu sudah berfungsi normal | Ready to use |  | Habibi
12000298 | 2026-01-29 | RS Advent Manado | Bor tulang | Bbraun/ELAN4 | 008036 | Service | Kabel power sobek. Perbaikan kabel power | Ready to use |  | Marshall
8000806 | 2026-01-29 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASHD0034 | Survey | Breathing system & leak fail, hit software, check diaphraphm piston dan MV2, check valve. Unit OK | Ready to use |  | Amna
9000764 | 2026-01-29 | RSD Tidore | Ventilator | Draeger/Babylog VN800 | ASTN0001 | Survey | Alarm MVe low. Cek kebocoran dan perbaiki breathing circuit dan sambungan ETT | Ready to use |  | Ruslan
9000763 | 2026-01-29 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASHJ0180 | Survey | Internal battery discharged. Baterai di cek mode service. Baterai normal dan berfungsi dengan baik | Ready to use |  | Ruslan
6000990 | 2026-01-29 | RSUD Kardinah | Baby Incubator | Draeger/Isolette C2000 | AF32541 | Survey | Engsel kanan patah, belahan kiri patah. | Pending |  | Dwi
7000943 | 2026-01-29 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASJH0127 | Service | Penggantian power supply M.7.3 dan internal baterai | Ready to use |  | Wahyu
20000450 | 2026-01-28 | RSUD Kota Baubau | Hemodialisa | Bbraun | 552070 | Service | Pemasangan tube acid dan penggantian maintenance kit | Ready to use |  | Sujana
20000449 | 2026-01-28 | RSUD Kota Baubau | Hemodialisa | Bbraun | 552071 | Service | Penggantian tube acid | Ready to use |  | Sujana
20000448 | 2026-01-28 | RS Siloam Buton | Hemodialisa | Bbraun | 241794 | Survey | Terdapat bunyi keras. Pengecekan dan pembersihan body pump, penggantian baterai internal dan baterai cmos | Ready to use |  | Sujana
20000447 | 2026-01-28 | RS Siloam Buton | Hemodialisa | Bbraun | 241793 | Survey | Terdapat bunyi keras. Pengecekan dan pembersihan body pump | Ready to use |  | Sujana
20000446 | 2026-01-28 | RS Siloam Buton | Hemodialisa | Bbraun | 218054 | Service | Bodypump bunyi keras. Pembersihan dan perbaikan stand bodypump | Ready to use |  | Sujana
14000400 | 2026-01-28 | RSUD RA Basoeni | Dental Unit | Kavo/1058 Life | 10205460 | Survey | Unit tidak bisa digunakan. Cek board control, cek board meja dokter error. | Pending |  | Fahim
18000750 | 2026-01-28 | RSUD Maria Walanda Maramis | Ventilator | Draeger/Evita V300 | ASPA0009 | Survey | Unit dilakukan tarik error dan akan dianalisis oleh tim draeger | Pending |  | Irfan R, Marshall
18000750 | 2026-01-28 | RSUD Maria Walanda Maramis | Ventilator | Draeger/ Savina 300 | ASNL0479 | Survey | Baterai rusak | Pending |  | Irfan R, Marshall
18000750 | 2026-01-28 | RSUD Maria Walanda Maramis | Ventilator | Draeger/ Savina 300 | ASNL0478 | Survey | Baterai rusak | Pending |  | Irfan R, Marshall
18000750 | 2026-01-28 | RSUD Maria Walanda Maramis | Ventilator | Draeger/ Savina 300 | ASHL0291 | Survey | Baterai rusak | Pending |  | Irfan R, Marshall
10000270 | 2026-01-28 | RSUD Beriman Balikpapan | Syringe Pump | Bbraun/Perfusor Compact Plus | 100010 | Survey | Baterai rusak | Pending |  | Adnan
10000269 | 2026-01-28 | RSUD Beriman Balikpapan | Syringe Pump | Bbraun/Perfusor Compact Plus | 27788 | Survey | Ketika alat dicabut dari sumber tegangan, alat alngsung mati. Baterai rusak | Pending |  | Adnan
6000995 | 2026-01-28 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASHM0015 | Service | Gas sensor bermasalah, O2 sensor tidak akurat membaca, sample line waterlock tidak terdetect; Pembersihan jalur PGM | Ready to use |  | Dwi
7000940 | 2026-01-28 | RSUD Bagas Waras | Infant Warmer | Draeger/Resuscitaire | JY10901 | Service | Kit corner bed damage; Penggantian kit corner blok | Ready to use |  | Wahyu
7000939 | 2026-01-28 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0144 | Service | Gas monitoring failure; Penggantian mini module pump | Ready to use |  | Wahyu
16000307 | 2026-01-27 | MP Dental | Kompressor | Durr/Tornado I | T529798084 | Instalasi | Unit sudah berfungsi | Ready to use |  | Farhat, Ipan
20000445 | 2026-01-27 | RS Siloam Buton | Hemodialisa | Bbraun | 108430 | Survey | Error mixing. Pengecekan valve VBicp dan pembersihan body pump. | Pending | Perlu penggantian valve VBicp | Sujana
20000444 | 2026-01-27 | RS Siloam Buton | Hemodialisa | Bbraun | 104603 | Survey | Error mixing. Pengecekan valve VBicp dan pembersihan body pump. | Pending | Perlu penggantian valve VBicp | Sujana
8000803 | 2026-01-27 | RS Bhina Bakti Husada | Mesin Anasthesia | Draeger/Atlan A300XL | ASMM0060 | Survey | Install demo alat | Pending |  | Amna
9000751 | 2026-01-27 | RSUD Weda | Mesin Anasthesia | Draeger/ Primus | ASNE0150 | Survey | LCD blank white screen. PBA mobi rusak | Pending |  | Ruslan
6000994 | 2026-01-27 | RSU An Ni'mah | Ventilator | Draeger/Babylog VN600 | ASSK0026 | Survey | Pengecekan unit. Baterai internal rusak, flow sensor expired | Pending |  | Dwi
19000567 | 2026-01-26 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0091 | Survey | Blower GS500 rusak. Tarik error. Blower rusak. Board M1.3 dan blower rusak | Pending |  | Nanda
20000443 | 2026-01-26 | RS Siloam Buton | Hemodialisa | Bbraun | 203006 | Survey | Error temperatur low. Pengecekan motor pump, sensor VPA dan body pump | Pending | Perlu penggantian motor pump, sensor VPA dan body pump | Sujana
20000442 | 2026-01-26 | RS Siloam Buton | Hemodialisa | Bbraun | 246972 | Survey | Darah naik ke bagian arteri. Pengecekan sensor arteri, sensor rusak | Pending |  | Sujana
20000441 | 2026-01-26 | RS Siloam Buton | Hemodialisa | Bbraun | 223615 | Service | Mesin bunyi keras dan tidak bisa mode hemodialisa. Pengecekan body pump, pembersihan body pump, penggantian baterai internal | Ready to use |  | Sujana
9000762 | 2026-01-26 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASHJ0179 | Survey | Unit mati total. Cek power supply. PSU rusak | Pending
8000804 | 2026-01-26 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0110 | Survey | O2 sensor fail, flow sensor fail. System check. O2 sensor capsule expired | Pending |  | Amna
8000804 | 2026-01-26 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0108 | Survey | O2 sensor fail, flow sensor fail. System check. O2 sensor capsule expired | Pending |  | Amna
8000804 | 2026-01-26 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0117 | Survey | O2 sensor fail, flow sensor fail. System check. O2 sensor capsule expired | Pending |  | Amna
6000989 | 2026-01-26 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASFL0120 | Survey | Display mati total. Disarankan oenggantian basis board | Pending |  | Dwi
6000989 | 2026-01-26 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASFL0121 | Survey | Display mati total. Disarankan oenggantian basis board | Pending |  | Dwi
7000938 | 2026-01-26 | RSA UGM | Jaundice Meter | Draeger/JM-105 | 3508916 | Service | Error 05. Penggantian baterai internal | Ready to use |  | Wahyu
12000297 | 2026-01-26 | RSUD Noongan | Mesin Anasthesia | Draeger/Perseus A500 | ASJF0093 | Survey | Software corrupt. Valve VMGS, sensor PDGMS, board M16 rusak | Pending |  | Marshall
12000295 | 2026-01-26 | RSUD Noongan | Mesin Anasthesia | Draeger/Perseus A500 | ASJL0006 | Survey | Software corrupt. Board M16 dan flow power rusak | Pending |  | Marshall
11000431 | 2026-01-26 | RS Primaya Inco Sarowoko | Hemodialisa | Bbraun | 549710 | Instalasi | Unit OK dan siap digunakan | Ready to use |  | Nasrul
11000431 | 2026-01-26 | RS Primaya Inco Sarowoko | Hemodialisa | Bbraun | 549713 | Instalasi | Unit OK dan siap digunakan | Ready to use |  | Nasrul
11000431 | 2026-01-26 | RS Primaya Inco Sarowoko | Hemodialisa | Bbraun | 549714 | Instalasi | Unit OK dan siap digunakan | Ready to use |  | Nasrul
11000431 | 2026-01-26 | RS Primaya Inco Sarowoko | Hemodialisa | Bbraun | 549711 | Instalasi | Unit OK dan siap digunakan | Ready to use |  | Nasrul
11000431 | 2026-01-26 | RS Primaya Inco Sarowoko | Hemodialisa | Bbraun | 549712 | Instalasi | Unit OK dan siap digunakan | Ready to use |  | Nasrul
9000756 | 2026-01-25 | RSUD Dr. H. Chasan Boesoirie | MOT | SYAS | - | Service | Pintu tidak tertutup rapat. Perbaikan rel motor. Pintu sudah berfungsi normal | Ready to use |  | Ruslan
20000454 | 2026-01-25 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum | G248L0145 | Service | Unit tidak bisa booting; Pengecekan power dan restart unit | Ready to use |  | Sujana
10000268 | 2026-01-24 | RS Hermina Balikpapan | Bor tulang | Bbraun/ELAN4 | 5437 | Survey | Outing pengancing mata bor pada handpiece sudah goyang; Handpiece dibawa ke bbraun untuk di cek | Pending |  | Adnan
11000430 | 2026-01-23 | RS PKU Muhammadiyah | Bubble CPAP | Draeger/Oxymixer | QM20215193 | Service | FIO2 tidak akurat; Penggantian maintenance kit oxymixer | Ready to use |  | Nasrul
7000937 | 2026-01-23 | RSUD Nyi Ageng Serang (KSO) | CT Scan | Siemens/Somatom | 120196 | Service | Printer fujifilm error EP01. Setting max value film | Ready to use |  | Wahyu
8000802 | 2026-01-23 | RS Islam Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0114 | Survey | N2O bocor. Cek central N20, line N2O. Running test, unit OK | Ready to use |  | Amna
16000306 | 2026-01-23 | RSKO Cibubur | Vaporizer | Draeger/Vapor 2000 | ASUM0496 | Service | Penggantian spare part | Ready to use |  | Ipan
16000306 | 2026-01-23 | RSKO Cibubur | Vaporizer | Draeger/Vapor 2000 | ASUM0173 | Service | Penggantian spare part | Ready to use |  | Ipan
16000306 | 2026-01-23 | RSKO Cibubur | Vaporizer | Draeger/Vapor 2000 | ASUN0427 | Service | Penggantian spare part | Ready to use |  | Ipan
12000294 | 2026-01-23 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1500 | Survey | PSU unit rusak | Pending |  | Marshall
22000050 | 2026-01-23 | Siloam Hospital manado | Hemodialisa | Bbraun | 214202 | Service | DF pressure 12550; Pengecekan membran dan penggantian membran | Ready to use |  | Ari
5000393 | 2026-01-22 | Smile Work Dental Care | Dental Unit | Kavo/E30 | 30212530 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Farhat, Ipan
5000394 | 2026-01-22 | Smile Work Dental Care | Dental Unit | Kavo/E30 | 30212529 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Farhat, Ipan
3000927 | 2026-01-22 | RSUD Dr. R. Soedjati Soemodiarjo | Digital Radiografi Fluoroskopi | Apelem/Platinum | H248L0156 | Survey | Training alat | Ready to use |  | Imam, Bahar
7000934 | 2026-01-21 | RSUD R.A.A Tjokronegoro | Bubble CPAP | Draeger/Oxymixer | QM20255065 | Survey | Uninstall unit demo | Ready to use |  | Wahyu
9000757 | 2026-01-20 | RSUD Dr. H. Chasan Boesoirie | MOT | SYAS | - | Service | Passbox tidak terbuka, sensor procimiti tidak bekerja. Buka passbox manual, dan buka kain yang terjepit di pintu. Unit sudah berfungsi normal | Ready to use |  | Ruslan
9000755 | 2026-01-20 | RSUD Jailolo | Dental Unit | Kavo/E30 | 10117281 | Survey | Scaller tidak berfungsi. Penggantian board scaller. Unit belum berfungsi karna harus sekalian dengan scaler tube | Pending |  | Ruslan
18000748 | 2026-01-20 | RSUD ODSK Provinsi Sulawesi Utara | Baby Incubator | Draeger/Babyleo TN500 | ASTH0042 | Survey | Suhu bayi tidak terbaca, sensor skin rusak. Penggantian sensor skin | Ready to use |  | Irfan R
19000561 | 2026-01-20 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Atlan A300 | ASPH0141 | Survey | Gas measurement fail, hit software, cek motor pump. Motor pump rusak | Pending |  | Nanda
6000993 | 2026-01-20 | RSUD Cilacap | USG | Mindray/Nuewa 19 | KV3-52000073 | Survey | Uninstall unit demo | Ready to use |  | Dwi
6000988 | 2026-01-20 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASUN0037 | Survey | O2 delivery failed. tarik error. Perlu penggantian PN 6873510 | Pending | Pengajuan claim | Dwi
7000933 | 2026-01-20 | RSUD Merah Putih | Jaundice Meter | Draeger/JM-105 | H3510375 | Survey | Error 05. Baterai JM105 rusak | Pending |  | Wahyu
220000049 | 2026-01-19 | Siloam Hospital manado | Hemodialisa | Bbraun | 214216 | Service | Self test failed. Pengecekan VP dan pembersihan detector | Ready to use |  | Ari
22000001 | 2026-01-19 | RS GMIM Bethesda | Hemodialisa | Bbraun | 504506 | Survey | Bicarbonat tidak konek; Pembersihan body pump dan pembersihan house top bottom | Ready to use |  | Ari
12000293 | 2026-01-19 | RSUD Noongan | X-ray Digital | Apelem/Camargue DR | IMA703798 | Survey | Lampu kolimator mati, detector tusak, pc lemot. Menunggu teknisi pusat untuk pengecekan lebih lanjut | Pending |  | Marshall
5000392 | 2026-01-19 | Signature Dental care | Handpiece | Kavo/E667L | 1020301 | Service | Handpiece tidak keluar air, jalur tersumbat; Pembersihan filter dan outlet air | Ready to use |  | Farhat, Ipan
6000987 | 2026-01-19 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASUM0001 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
7000942 | 2026-01-19 | RSUP dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Atlan A550 | ASNH0125 | Survey | Installasi unit demo | Ready to use |  | Wahyu
8000801 | 2026-01-15 | RSUD Dr. Loekmono Hadi Kudus | Infant Warmer | Draeger/Resuscitaire | AY11083 | Survey | Error 04. Test heater, test SSR, test voltage range | Pending |  | Amna
14000399 | 2026-01-15 | RSUD H. Slamet Martodirdjo | ESWL | Medispec/E3000 | 1017902 | Survey | Unit tidak bisa menembak. HVG rusak dan control box rusak | Pending |  | Fahim, Bahar
6000985 | 2026-01-15 | RSUI Harapan Anda Tegal | Baby Incubator | Draeger/Fabius Plus | ASEL0010 | Survey | Pengecekan unit. Disarankan penggantian O2 cell, flow sensor, service set 3 thn | Ready to use |  | Dwi
6000985 | 2026-01-15 | RSUI Harapan Anda Tegal | Baby Incubator | Draeger/Isolette C2000 | DSPC0022 | Survey | Suara bising pada blower fan; Perbaikan cleaning fan blower | Ready to use |  | Dwi
11000429 | 2026-01-15 | RSUP Tadjuddin Chalid | Digital Radiografi Fluoroskopi | Apelem/Platinum | J236F0882 | Service | Monitor console stuck (tidak bisa start exam). Cek CPU dan restart unit | Ready to use |  | Nasrul
20000461 | 2026-01-14 | RSUP Dr. Wahidin Sudirohusodo | Syringe Pump | Bbraun/Infusomat | 905582 | Service | Penggantian baterai | Ready to use |  | Sujana
20000461 | 2026-01-14 | RSUP Dr. Wahidin Sudirohusodo | Syringe Pump | Bbraun/Infusomat | 905388 | Service | Penggantian baterai | Ready to use |  | Sujana
20000461 | 2026-01-14 | RSUP Dr. Wahidin Sudirohusodo | Syringe Pump | Bbraun/Infusomat | 905400 | Service | Penggantian baterai | Ready to use |  | Sujana
7000932 | 2026-01-14 | RSUD Merah Putih | X-ray Digital | Apelem/Camargue DR | IMA703979 | Service | Vertical movement stuck; Penggantian tensator spring | Ready to use |  | Wahyu
18000749 | 2026-01-14 | RSUD Noongan | Baby Incubator | Draeger/Babyleo TN500 | ASJM0024 | Survey | Modul sensor humidity rusak | Pending |  | Irfan R, Marshall
18000749 | 2026-01-14 | RSUD Noongan | Ventilator | Draeger/babylog VN500 | ASMK035 | Survey | Kompressor extend rusak | Pending |  | Irfan R, Marshall
19000560 | 2026-01-14 | RSUD Kota Salatiga | Patient Monitor | Draeger/Infinity Gamma XL | 5517259967 | Survey | Adaptor hilang, butuh adaptor baru | Pending |  | Nanda
19000559 | 2026-01-14 | RSUD Kota Salatiga | Infant Warmer | Draeger/Resuscitaire | FM09495 | Survey | Error 05, heater putus | Pending |  | Nanda
8000805 | 2026-01-14 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0147 | Survey | O2 sensor fail, general check system check. O2 sensor capsule expired | Pending |  | Amna
8000805 | 2026-01-14 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASNB0020 | Survey | O2 sensor fail, general check system check. O2 sensor capsule expired | Pending |  | Amna
8000805 | 2026-01-14 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0043 | Survey | O2 sensor fail, general check system check. O2 sensor capsule expired | Pending |  | Amna
9000761 | 2026-01-14 | RSD TIdore | Baby Incubator | Draeger/Isolette 8000 | DH08425 | Survey | Hinge access rusak | Pending |  | Ruslan
6000984 | 2026-01-14 | RSUD Ir. Soekarno | USG | Mindray/DC30 | JM3-48006338 | Service | Penggantian baterai kit (Warranty) | Ready to use |  | Dwi
18000658 | 2026-01-13 | RS Advent Manado | Bor tulang | Aesculap/ELAN4 | 8036 | Maintenance | Preventive Maintenance | Ready to use |  | Irfan R
19000558 | 2026-01-13 | RSUD dr. Adhyatma | Baby Incubator | Draeger/Caleo | ASHK0031 | Survey | Unit dikerubungi semut, cek kabel dan dalam incubator. Unit OK | Ready to use |  | Nanda
19000557 | 2026-01-13 | RSUD dr. Adhyatma | Baby Incubator | Draeger/Caleo | ASHK0035 | Survey | Ada bau gosong, cek humidity, water heater berkerak karna user salah gunakan cairan infus DS. Pembersihan water heater | Pending | Unit sedang dilakukan pengecekan (running test) | Nanda
8000789 | 2026-01-13 | RSUD Bendan Pekalongan | Mesin Anasthesia | Draeger/Fabius Tiro | ASHK0143 | Survey | FIO2 fail. System check, leakage check, piston-valve check, CO2y check. O2 capsule expired | Pending | Perlu penggantian service set 3 tahun dan O2 sensir capsule serta CO2y valve | Amna
9000695 | 2026-01-13 | RSUD Jailolo | Dental Unit | Kavo/E30 | 10117281 | Survey | Handpiece sacler tidak berfungsi. Sudah diganti module namun kabel power masih kendala. | Pending | Menunggu kabel power datang | Ruslan
9000694 | 2026-01-13 | RSUD Jailolo | Ventilator | Draeger/Babylog VN800 | ASTC0001 | Survey | O2 supply fail. Device checkm HIT software; O2 supply harus 99% | Pending |  | Ruslan
10000267 | 2026-01-13 | RSUD Dayaku Aji | Pendant | SYAS | SY-PI2506005 | Service | Kebocoran arus pada lengan pendant, housebreak bocor; Rakit ulang housebreak dan pemasangan kabel stop kontak | Ready to use |  | Adnan
6000983 | 2026-01-13 | RSU Nirmala Purbalingga | Ventilator | Draeger/Savina 300 | ASRE0307 | Service | O2 kalibrasi gagal, baterai drop, expirasi valve gagal; Penggantian Oxytrave VE dan service set 2 tahun | Ready to use |  | Dwi
22000002 | 2026-01-13 | Siloam Hospital Manado | Hemodialisa | Bbraun | 264992 | Service | Temperature too high; Membersihan sensor temperator dan cek valve | Ready to use |  | Ari
10000266 | 2026-01-12 | RSUD Dayaku Aji | Pendant | SYAS | SY-PI2506007 | Service | Kebocoran housbreak dan stop kontak tidak ada arus yang masuk; Rakit ulang housebreak dan kabel stop kontak | Ready to use |  | Adnan
9000697 | 2026-01-12 | RSUD Jailolo | UPS | OAXAKA/AP3330 | 110045034F65393600001 | Service | Koneksikan kabel power UPS dan kabel output panel UPS ke panel distrubusi | Pending | Unit belum berfungsi karna belum ada sumber 3 phase | Ruslan
6000991 | 2026-01-12 | RSUD Cilacap | Mesin Anasthesia | Draeger/Perseus A500 | ASUJ0001 | Survey | Training alat | Ready to use |  | Dwi
8000786 | 2026-01-12 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0033 | Survey | Sensor check fail, system check. Check V-mix valve, unit sudah berfungsi | Ready to use |  | Amna
6000991 | 2026-01-12 | RSUD Cilacap | Mesin Anasthesia | Draeger/Perseus A500 | ASUJ0001 | Survey | Refresh Training | Ready to use |  | Dwi
7000935 | 2026-01-12 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASJH0125 | Service | VT tidak stabil, nebulizer failure. Penggantian board gas dosage M1.3, board PVT mi=onitor M4.1, cable harnes flow sensor, pressure regulator | Ready to use |  | Wahyu
11000428 | 2026-01-12 | RS Islam Faisal | ESWL | Medispec/EM1000 | 1214012 | Service | EMH tidak mengembang, bocor pada pembuangan. Motor actuator rusak | Pending |  | Nasrul
10000264 | 2026-01-11 | RSUD Dayaku Aji | Pendant | SYAS | SY-PI2506005 | Survey | Kebocoran housbreak dan stop kontak tidak ada arus yang masuk; Rakit ulang housebreak dan kabel stop kontak | Ready to use |  | Adnan
10000263 | 2026-01-11 | RSUD Dayaku Aji | Pendant | SYAS | SY-PI2506006 | Survey | Kebocoran housbreak dan stop kontak tidak ada arus yang masuk; Rakit ulang housebreak dan kabel stop kontak | Ready to use |  | Adnan
22000003 | 2026-01-10 | Siloam Hospital Manado | Hemodialisa | Bbraun | 313796 | Service | Blood leak test failure; Penggantian blood leak dan kalibrasi unit | Ready to use |  | Ari
9000693 | 2026-01-09 | RSUD Dr. H. Chasan Boesoirie | Bor tulang | Bbraun/Elan4 | 003988 | Service | Handpiece crathiotome dan handpiece L7 rusak. Unit sudah diperbaiki bbraun dan tarik unit backup | Ready to use |  | Ruslan
3000926 | 2026-01-09 | RSUD Langsa | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/057/C1 | Survey | Alat tidak menyala pasca banjir. Pengecekan semua bagian unit. Kerusakan pada PC Console, PC Workstation, Inverter, PCB 190, HSS Board, Tabung HTT (Generator), Trafo TF167, Trafo TF165, Trafo TF170, Trafo TF 155, Motor arm dan gearbox, Solid State Relay (SSR) | Pending |  | Imam, Bahar
10000262 | 2026-01-09 | RSUD Beriman Balikpapan | Infus Pump | Bbraun/Infusomat P | 115813 | Survey | LCD ngeblank, jika unit dinyalakan cahaya di LCD hanya setengah dan baterai sudah tidak menyimpan tegangan | Pending |  | Adnan
10000261 | 2026-01-09 | RSUD Beriman Balikpapan | Infus Pump | Bbraun/Infusomat P | 103204 | Survey | Alat mati total. Baterai sudah tidak menyimpan tegangan; Mainboard dan battery perlu diganti | Pending |  | Adnan
10000260 | 2026-01-09 | RSUD Beriman Balikpapan | Syringe Pump | Bbraun/Perfusor Compact Plus | 037636 | Survey | Battery rusak | Pending |  | Adnan
10000259 | 2026-01-09 | RSUD Beriman Balikpapan | Syringe Pump | Bbraun/Perfusor Compact Plus | 032035 | Survey | Tombol power tidak bisa ditekan, baterai sudah tidak mengisi; Cover/housing alat dan baterai rusak | Pending |  | Adnan
6000992 | 2026-01-09 | RSUD Cilacap | USG | Mindray/M9 | CCI-51006326 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Dwi
18000661 | 2026-01-08 | RS Gunung Maria Tomohon | Infant Warmer | Draeger/Baby Thermal | ASPL0108 | Survey | Error 35. Heater rusak | Pending |  | Irfan R
19000556 | 2026-01-08 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0087 | Service | Breathing circuit test fail, kondisi barang kurang bagus. Penggantian breathing circuit, leakage 0% . Unit OK | Ready to use |  | Nanda
8000787 | 2026-01-08 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASHL0063 | Survey | Gas analyser fail, fresh gas delivery fail; Minimodule pump rusak | Pending |  | Amna
17000225 | 2026-01-08 | RSUD Cut Meutia | ESWL | Medispec/EM1000 | 0119127 | Service | Penggantian HVPS 1pcs, selang air. Pembersihan EMH | Ready to use |  | Bahar, Imam
12000299 | 2026-01-08 | RS GMIM Bethesda | Ventilator | Draeger/Savina 300 | ASTJ0024 | Service | Error leakage, bocor breathing circuit. Penggantian breathing circuit | Ready to use |  | Marshall
11000427 | 2026-01-08 | RS Siloam Makassar | Hemodialisa | Bbraun | 250198 | Service | Chamber bocor. Cek sensor level air, adjsut NSVB | Ready to use |  | Nasrul
12000288 | 2026-01-07 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Service | Tidak bisa daily check, film tidak terbaca. Test daily ulang, setting ulang film printer. Unit sudah normal | Ready to use |  | Marshall
7000931 | 2026-01-07 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0144 | Survey | O2 sensor dan gas analyser failure, pump faulty. Mini module pump rusak | Pending |  | Wahyu
7000930 | 2026-01-07 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASJH0125 | Service | Cockpit restart. Penggantian service board cockpit | Ready to use |  | Wahyu
18000659 | 2026-01-07 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Survey | Daily check gagal. Melakuakn daily check sesuai prosedur | Ready to use |  | Irfan R, Marshall
19000555 | 2026-01-07 | RSUD Simo Boyolali | Mesin Anasthesia | Draeger/Fabius Plus XL | ASTC0003 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Nanda
8000785 | 2026-01-07 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0115 | Survey | O2 sensor fail, check o2 sensor dan system check; O2 capsule expired | Pending |  | Amna
8000785 | 2026-01-07 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASII0148 | Survey | O2 sensor fail, check o2 sensor dan system check; O2 capsule expired | Pending |  | Amna
8000785 | 2026-01-07 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0049 | Survey | O2 sensor fail, check o2 sensor dan system check; O2 capsule expired | Pending |  | Amna
8000785 | 2026-01-07 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0048 | Survey | O2 sensor fail, check o2 sensor dan system check; O2 capsule expired | Pending |  | Amna
9000696 | 2026-01-07 | RSUD Jailolo | UPS | OAXAKA/AP3330 | 110045034F6539360001 | Survey | Koneksikan power listrik dan kabel output dari UPS ke panel distrubusi. Unit belum aktif karna kabel panel distribusi belum ada dari RS | Pending |  | Ruslan
11000426 | 2026-01-07 | RSUP Tadjuddin Chalid | Mesin Anasthesia | Draeger/Primus | ASJL0023 | Survey | O2 sensor failure. Flow power dan O2 sensor peru diganti | Pending |  | Nasrul
6000982 | 2026-01-06 | RSUD Ajibarang | USG | Mindray/DC30 | GP-18009329 | Survey | Crystal probe convex 3C5P mati dari sisi tengah ke samping. | Pending | Disarankan penggantian probe dan UPS minimal 2000 watt | Dwi
14000397 | 2026-01-06 | RSUP Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum | G247L0146 | Service | Console tidak menyala. Fuse putus, penggantian fuse | Ready to use |  | Fahim
8000781 | 2026-01-06 | RSUD Kayen Pati | Patient Monitor | Draeger/M540 | 5615811970 | Service | Unit tidak menyala. Pengecekan power supply. Power supply rusak | Pending |  | Amna
6000982 | 2026-01-06 | RSUD Ajibarang | USG | Mindray/ DC-30 | GP-18009329 | Survey | Crystal probe convex 3C5P mati dari sisi tengah ke samping. | Pending |  | Dwi
19000553 | 2026-01-06 | RS St Elizabeth Semarang | Mesin Anasthesia | Draeger/Fabius Tiro | ASHA0114 | Service | Unit tidak bisa menyala, penggantian baterai | Ready to use |  | Nanda
9000689 | 2026-01-06 | RSUD Dr. H. Chasan Boesoirie | MOT | SYAS | - | Service | Alarm low pressure, pengecekan kebocoran unit outdoor dan cek HVAC. Tidak ada kebocoran, unit normal, alarm karna pintu dibuka terlalu lama dan mempengaruhi suhu ruangan | Ready to use |  | Ruslan
9000688 | 2026-01-06 | RSUD Dr. H. Chasan Boesoirie | Lampu Operasi | SYAS | OL9680/80XUZY00012 | Service | Setting posisi lengan lampu | Ready to use |  | Ruslan
10000258 | 2026-01-06 | RS Restu Ibu Balikpapan | Hemodialisa | Bbraun/Dialog+ | 300008 | Service | Temperature too high; Pembersihan sensor suhu dan kalibrasi sensor | Ready to use |  | Adnan
11000425 | 2026-01-05 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum | G248L0145 | service | Konek PCAS dengan unit; setting IP Address | Ready to use |  | Nasrul
21000001 | 2026-01-05 | RSUD H. Boejasin | Laparascopy | Aesculap | 497628 | Survey | Pengecekan alat karna unit lama tidak digunakan | Ready to use |  | Habibi
9000686 | 2026-01-05 | RSUD Dr. H. Chasan Boesoirie | MOT | SYAS | - | Service | Monitor film viewer X-ray mati. Perbaiki posisi kabel power dan restart PC | Ready to use |  | Ruslan
18000660 | 2026-01-05 | RS Mata Manado | Mesin Anasthesia | Draeger/Fabius | ASNE0014 | Service | Flow sensor gagal kalibrasi. Cek unit. Flow sensor gagal. Kalibrasi ulang, unit sudah normal | Ready to use |  | Irfan R
9000686 | 2026-01-05 | RSUD Dr. H. Chasan Boesoirie | MOT | SYAS | - | Service | Pintu sliding tidak bisa terbuka (mati), periksa MCB pada trafo isolasi dan ubah settingan pada monitor panel. Unit sudah berfungsi | Ready to use |  | Ruslan
22000004 | 2026-01-05 | SIloam Hospital Manado | Hemodialisa | Bbraun | 216630 | Service | Bicarbonat mixing. Perbaikan bodypump bicarbonat | Ready to use |  | Ari
20000467 | 2026-02-01 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XV2U00068 | Survey | LCD control mati. Pengecekan lcd kontrol, power lampu, reconnect LCD. Unit perlu penggantian lcd control | Pending |  | Sujana
9000692 | 2025-12-31 | RSUD Sanana | Pendant Anasthesia | SYAS | SY-PA2501008 | Survey | Unit belum diinstal karna struktur pendant belum dibuat | Pending |  | Ruslan
6000979 | 2025-12-31 | RSUD Panembahan Senopati | Bubble CPAP | Draeger/Oxymixer | QM00230153 | Survey | Tidak lolos kalibrasi; Cek alarm balancing, replacing membran dan kalibrasi O2 flow; Unit sudah berfungsi normal | Ready to use |  | Dwi
6000979 | 2025-12-31 | RSUD Panembahan Senopati | Bubble CPAP | Draeger/Oxymixer | QM20230154 | Survey | Tidak lolos kalibrasi; Cek alarm balancing, replacing membran dan kalibrasi O2 flow; Unit sudah berfungsi normal | Ready to use |  | Dwi
3000925 | 2025-12-31 | RSUD Dr. R. Soedjati Soemodiarjo | Digital Radiografi Fluoroskopi | Apelem/Platinum | 251025003 | Survey | Training user | Ready to use |  | Imam, Bahar
19000552 | 2025-12-30 | RS JIH Solo | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0125 | Maintenance | Maintenance Service contract | Ready to use |  | Nanda
19000551 | 2025-12-30 | RS JIH Solo | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Maintenance | Maintenance Service contract | Ready to use |  | Nanda
11000424 | 2025-12-30 | RS Regional dr. Hasri Ainun Habibie | Ventilator | Draeger/Babylog VN800 | ASSM0042 | Survey | FIO2 error,Neoflow sensor fail. Flow sensor perlu diganti | Pending |  | Nasrul
18000664 | 2025-12-30 | RSUP Prof. Dr. R. D. Kandou Manado | Bor tulang | Aesculap/Elan4 | - | Service | Handpiece tidak berputar. Dibawa untuk diperbaiki ke bbraun. SN Handpiece: 003101 | Pending |  | Irfan Rusli
6000972 | 2025-12-30 | RSUD Dr. Soeselo Slawi | MOT | SYAS | - | Service | Plafon MOT bocor. Pemasangan galvalum 5 lembang pada bagian atas MOT dan sealing bagian lubang/celah MOT plafon dan penggantian lighting room. | Ready to use |  | Dwi
20000435 | 2025-12-30 | RSUP Makassar | MOT | SYAS | - | Service | HVAC panas; Perlu dilakukan pencucian AC | Ready to use |  | Sujana
12000292 | 2025-12-30 | RSUD ODSK Provinsi Sulawesi Utara | Hemodialisa | Bbraun | 518850 | Service | Penggantian filter endoksin | Ready to use |  | Marshall
9000690 | 2025-12-29 | RSUD Sanana | Baby Bed | SYAS | 250526211000400056 | Instalasi | Unit berfungsi dengan baik dan normal | Ready to use |  | Ruslan
9000690 | 2025-12-29 | RSUD Sanana | Baby Bed | SYAS | 250526211000400059 | Instalasi | Unit berfungsi dengan baik dan normal | Ready to use |  | Ruslan
9000690 | 2025-12-29 | RSUD Sanana | Baby Bed | SYAS | 250526211000400060 | Instalasi | Unit berfungsi dengan baik dan normal | Ready to use |  | Ruslan
8000783 | 2025-12-29 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Atlan A300 XL | ASTJ0118 | Service | Upgrade fitur indicator gas setting, display gas cinsumption, mu X co2 trend | Ready to use |  | Amna
6000978 | 2025-12-29 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASLD0042 | Maintenance | Preventive Maintenance Contract Service | Ready to use |  | Amna
5000389 | 2025-12-27 | M&P Dental Specialists | Dental Unit | Planmeca/ Compact i5 | UC16191427 | Instalasi | Unit berfungsi dengan baik dan normal | Ready to use |  | Farhat, Ipan
11000423 | 2025-12-26 | RSKD Dadi | C-arm | Apelem/Evo R+ | 55120063849 | Service | Penggantian modul inverter | Ready to use |  | Nasrul
18000662 | 2025-12-24 | RS GMIM Pancaran | Ventilator | Draeger/Savina 300 | ASTJ0020 | Service | Valve merah saat device check. Pembersihan valve dan membran | Ready to use |  | Irfan R
7000928 | 2025-12-24 | RSUP dr. Sardjito | Sternotormy | Aesculap | - | Instalasi | Unit berfungsi dan aksesoris lengkap | Ready to use |  | Wahyu
6000977 | 2025-12-24 | RSU An Ni'mah | Ventilator | Draeger/Babylog VN600 | ASSK0026 | Survey | Flow sensor fail. Tarik error. Kit air filter harus ganti | Pending |  | Dwi
3000924 | 2025-12-24 | RSUD Malingping | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/131/C2 | Survey | Kalibrasi colimator. Daily check gagal karna tegangan drop dan suhu ruangan tidak dingin | Pending | Akan dilanjutkan jika suhu ruangan sudah dingin dan tegangan stabil | Imam, Bahar
8000782 | 2025-12-24 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0138 | Service | LCD TFT Fail; Penggantian LCD TFT 6,5 | Ready to use |  | Amna
9000684 | 2025-12-24 | RSUD Dr. H. Chasan Boesoirie | Mesin Anasthesia | Draeger/Perseus A500 | ASSK0059 | Survey | SPO2 sensor fail. Sensor SPO2 tidak menyala. Sensor SPO2 rusak | Pending |  | Ruslan
9000685 | 2025-12-24 | RSUD Dr. H. Chasan Boesoirie | MOT | SYAS | - | Service | Pintu passbox tidak bisa dibuka; Setting posisi sensor | Ready to use |  | Ruslan
19000565 | 2025-12-23 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASUM0457 | Service | Penukaran vaporizer demo dengan vaporizer bawaan unit | Ready to use |  | Nanda
19000565 | 2025-12-23 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASUM0057 | Service | Penukaran vaporizer demo dengan vaporizer bawaan unit | Ready to use |  | Nanda
20000434 | 2025-12-23 | RSUD I Lagaligo | Hemodialisa | Bbraun/Dialog + | 226699 | Service | Pemasangan motor FPA | Ready to use |  | Sujana
14000395 | 2025-12-23 | RSUP Surabaya | C-arm | Apelem/Evo R+ | 551240231085 | Service | Keyboard monitor tidak berfungsi; Perbaikan connector sambungan keyboard | Ready to use |  | Fahim
9000681 | 2025-12-23 | RSUD Weda | Mesin Anasthesia | Draeger/Primus | ASNE0150 | Survey | Pengecekan unit. Disarankan penggantian service set | Pending |  | Ruslan
9000682 | 2025-12-23 | RSUD Weda | Dental Unit | Kavo/1058 Life | 10209459 | Survey | Alarm bunyi saat unit digunakan. Cek sistem dan jalur air. Tendon air bersih kosong dan jalur pipa terputus. Mainboard rusak | Pending |  | Ruslan
6000975 | 2025-12-23 | RSUD Cilacap | Vaporizer | Draeger/Vapor 2000 | ASUM0552 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
7000797 | 2025-12-23 | RSUD Kota Yogyakarta | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Service | Flow meter pecah. Penggantian flow meter | Ready to use |  | Wahyu
7000796 | 2025-12-23 | RSUD Kota Yogyakarta | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Maintenance | Preventinve Maintenance Total Care | Ready to use |  | Wahyu
8000781 | 2025-12-23 | RSUD RA KArtini | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0118 | Service | LCD TFT Fail; Penggantian LCD TFT 6,5 | Ready to use |  | Amna
18000747 | 2025-12-22 | RS GMIM Pancaran | Ventilator | Draeger/Savina 300 | ASTJ0018 | Survey | Refresh Training | Ready to use |  | Irfan R, Marshall
19000549 | 2025-12-22 | RSUP dr. kariadi | Baby Incubator | Draeger/Caleo | ARUC0050 | Service | Air humidity bocor. Cek selenoid water heater, sensor level air. Sensor level air rusak sementara diperbaiki dengan dilem | Ready to use |  | Nanda
6000974 | 2025-12-19 | RSUD Cilacap | USG | Mindray/Nuewa 19 | KV3-52000073 | Service | Installasi demo unit | Ready to use |  | Dwi
8000780 | 2025-12-19 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Atlan | ASUJ0153 | Service | Setting AGSS Scavenging unit | Ready to use |  | Amna
8000780 | 2025-12-19 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Atlan | ASUJ0152 | service | Setting AGSS Scavenging unit | Ready to use |  | Amna
9000678 | 2025-12-19 | RSUD Dr. H. Chasan Boesoirie | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/089/C1 | Survey | Film tidak keluar saat mode print; Perbaiki posisi tray dan setting dicom | Ready to use |  | Ruslan
9000679 | 2025-12-19 | RSUD Dr. H. Chasan Boesoirie | Central Monitor | Draeger/CMS Vista 300 | CASUC009 | Service | Koneksi patient monitor ke CMS sering putus. Restrat PC dan router | Ready to use |  | Ruslan
9000680 | 2025-12-19 | RSUD Dr. H. Chasan Boesoirie | MOT | SYAS | - | Service | Lemari instrument tidak bisa mengunci. Penggantian part kunci lemari instrument | Ready to use |  | Ruslan
11000422 | 2025-12-19 | RSUD Tenriawaru | Hemodialisa | Bbraun/Dialog + | 241701 | Service | Error DF pressure; Cek mainboard , cek koneksi, cek VSDEK 2 dan kalibrasi | Ready to use |  | Nasrul
11000422 | 2025-12-19 | RSUD Tenriawaru | Hemodialisa | Bbraun/Dialog + | 264404 | Service | Error DF pressure; Cek mainboard , cek koneksi, cek VSDEK 2 dan kalibrasi | Ready to use |  | Nasrul
8000779 | 2025-12-18 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHA0113 | Survey | Leakage dan breathing system fail. System check, breathing system check. Expiratory valve check | Ready to use |  | Amna
19000545 | 2025-12-18 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Service | Chamber untuk test batu ginjal rusak; Lem chamber | Ready to use |  | Nanda
3000923 | 2025-12-18 | RSUD Pagelaran | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/136/C2 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Imam
14000394 | 2025-12-18 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Survey | Unit tidak bisa menembak. HVG rusak | Pending |  | Fahim
20000433 | 2025-12-18 | RSUP Dr. Wahidin Sudirohusodo | Laparascopy | Aesculab | 251125 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana, Nasrul
20000432 | 2025-12-17 | RSUP Makassar | MOT | SYAS | - | Service | Digital timer error; Power lost, perbaikan digital timer. Unit sudah kembali berfungsi | Ready to use |  | Sujana
8000777 | 2025-12-17 | RSI PKU Muhammadiyah Pekajangan | Mesin Anasthesia | Draeger/Atlan A300 | ASUM0315 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
6000973 | 2025-12-17 | RSUD Dr. Soeselo Slawi | Hospital Furniture | SYAS | - | Instalasi | Installasi Premium bed 3 motor 10 unit, overbed table 10 unit, bedside cabinet 10 unit, wheel chair 4 unit, trolley emergency 2 unit, trolly instrument 5 unit, trolley dressing 2 unit | Ready to use |  | Dwi
19000544 | 2025-12-17 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Service | Pointer tidak sesuai; Setting dan adjustment | Ready to use |  | Nanda
5000387 | 2025-12-17 | Smile Work Dental Care | Dental Unit | Kavo/E50 | 30212529 | Service | Keluar bunyi saat kursi naik turun. Bongkar bawah kursi. Setting lampu | Ready to use |  | Farhat
13000250 | 2025-12-17 | RS Otak Moh. Hatta | C-arm | Apelem/Evo R+FP | 0102-24-003-012 | Survey | Layar monitor ada kerusakan pada black screen (Monitor BENQ) | Pending | Perlu penggantian layar monitor | Irfan Izzudin
18000666 | 2025-12-18 | RSUD Aloei Saboe | Trolley Emergency | SYAS | 251169501000400165 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan Rusli
18000666 | 2025-12-18 | RSUD Aloei Saboe | Trolley Emergency | SYAS | 251169501000400162 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan Rusli
19000564 | 2025-12-16 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASUJ0096 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
9000675 | 2025-12-16 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0011 | Survey | Gas supply unit fail; Blower rusak | Pending | Proses claim warranty blower unit | Ruslan
9000675 | 2025-12-16 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0012 | Survey | Gas supply unit fail; Blower rusak | Pending | Proses claim warranty blower unit | Ruslan
7000926 | 2025-12-16 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0092 | Survey | Cockpit restart; Coba service board cockpit dan reload software, unit masih belum bisa | Pending |  | Wahyu
3000922 | 2025-12-16 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | UPS console BEP. Perbaiki kabel power tegangan input ke ups | Ready to use |  | Imam
20000431 | 2025-12-16 | RSUP Dr. Wahidin Sudirohusodo | Hemodialisa | Bbraun/Dialog + | 554457 | Service | Heater fail. Penggantian heater | Ready to use |  | Sujana, Nasrul
10000255 | 2025-12-16 | Siloam Hospital Balikpapan | Hemodialisa | Bbraun/Dialog + | - | Survey | Pengecekan alat | Ready to use |  | Adnan
10000254 | 2025-12-15 | RS Hermina Balikpapan | Hemodialisa | Bbraun | - | Survey | Disinfectan routine | Ready to use |  | Adnan
10000253 | 2025-12-15 | RS Hermina Balikpapan | Syringe Pump | Bbraun/Perfusor Compact Plus | 76050 | Service | LCD intermitten. Kadang ngeblank hitam kadang tidak; Perbaiki kabel fleksible LCD | Ready to use |  | Adnan
19000563 | 2025-12-15 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASUJ0096 | Service | Setting unit untuk training user dan pemeriksaan | Ready to use |  | Nanda
5000386 | 2025-12-15 | FKG Dr. Moestopo | Dental Simulator | Kavo/DSE Compact | 00000855 | Service | Lampu patah; Oenggantian part lampu | Ready to use |  | Farhat, Bahar
16000305 | 2025-12-15 | RSU Haji Medan | Dental Unit | Kavo/1058 Life | 10207192 | Service | Penggantian main board dan update software | Ready to use |  | Ipan
8000777 | 2025-12-15 | RSI PKU Muhammadiyah Pekajangan | Mesin Anasthesia | Draeger/Atlan A300 | ASUM0315 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
20000430 | 2025-12-15 | RSUP Dr. Wahidin Sudirohusodo | Hemodialisa | Bbraun/Dialog + | 554829 | Service | Error water leak. Supply air bocor, dan filter lepas; Unit sudah diperbaiki dan berfungsi dengan baik | Ready to use |  | Sujana
20000430 | 2025-12-15 | RSUP Dr. Wahidin Sudirohusodo | Hemodialisa | Bbraun/Dialog + | 554869 | Service | Error water leak. Supply air bocor, dan filter lepas; Unit sudah diperbaiki dan berfungsi dengan baik | Ready to use |  | Sujana
12000289 | 2025-12-15 | RS Advent Manado | Bor tulang | Aesculap | 008036 | Survey | Kabel power sobek, setelah di cek tidak ada yang putus. Unit masih bisa digunakan | Ready to use |  | Marshall
12000290 | 2025-12-15 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1500 | Survey | Unit tidak nyala. Cek baterai, cek input supply power. Baterai rusak tidak bisa di charge | Pending |  | Marshall
19000562 | 2025-12-12 | RSUP Surakarta | USG | Mindray/DP-20T | LK7B36001327 | Survey | Training user
6000968 | 2025-12-12 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASUM00001 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
6000969 | 2025-12-12 | RSUD Panti Nugroho | Infant Warmer | Draeger/Resuscitaire | MV11418 | Survey | Knop On/off gas supply bocor, pemeriksaan valve control knop. Valve control mini 5 way perlu diganti | Pending |  | Dwi
6000970 | 2025-12-12 | RSUD Dr. Soeselo Slawi | EyeMag | Zeiss/Eyemag Pro F | 6905522656 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
13000252 | 2025-12-12 | RSUP Dr. M. Djamil | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0029 | Survey | System check, leak test tidak complete. Pengecekan pada APL Valve dan breathing circuit; APL Valve rusak | Pending |  | Irfan Izzudin
5000388 | 2025-12-12 | RSUD Cibinong | Dressing Trolley | SYAS/SCD1 | 251169201000400160 | Instalasi | Alat berfungsi dengan baik dan benar | Ready to use |  | Farhat
5000388 | 2025-12-12 | RSUD Cibinong | Dressing Trolley | SYAS/SCD1 | 251169201000400159 | Instalasi | Alat berfungsi dengan baik dan benar | Ready to use |  | Farhat
5000388 | 2025-12-12 | RSUD Cibinong | Dressing Trolley | SYAS/SCD1 | 251169201000400161 | Instalasi | Alat berfungsi dengan baik dan benar | Ready to use |  | Farhat
5000388 | 2025-12-12 | RSUD Cibinong | Dressing Trolley | SYAS/SCD1 | 251169201000400162 | Instalasi | Alat berfungsi dengan baik dan benar | Ready to use |  | Farhat
5000388 | 2025-12-12 | RSUD Cibinong | Dressing Trolley | SYAS/SCD1 | 251169201000400163 | Instalasi | Alat berfungsi dengan baik dan benar | Ready to use |  | Farhat
5000388 | 2025-12-12 | RSUD Cibinong | Dressing Trolley | SYAS/SCD1 | 251169201000400164 | Instalasi | Alat berfungsi dengan baik dan benar | Ready to use |  | Farhat
12000291 | 2025-12-12 | RSUD ODSK Provinsi Sulawesi Utara | Hemodialisa | Bbraun | 518512 | Survey | Unit bocor | Ready to use |  | Marshall
10000252 | 2025-12-11 | RS Restu Ibu Balikpapan | Hemodialisa | Bbraun/Dialog+ | 504212 | Survey | Pengecekan mesin RO | Ready to use |  | Adnan
11000421 | 2025-12-11 | RSUP Makassar | Ventilator | Draeger/Evita V300 | ASPH0033 | Instalasi | Alat berfungsi dengan baik dan benar | Ready to use |  | Nasrul
11000421 | 2025-12-11 | RSUP Makassar | Ventilator | Draeger/Evita V300 | ASPH0032 | Instalasi | Alat berfungsi dengan baik dan benar | Ready to use |  | Nasrul
14000391 | 2025-12-11 | RSUD Grati Pasuruan | Wheel chair | SYAS | 251231313000401367 | Instalasi | Telah diinstall 15 unit. Alat berfungsi dengan baik | Ready to use |  | Fahim
7000924 | 2025-12-11 | RSUP dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0109 | Service | Maintenance kit 3 years | Ready to use |  | Wahyu
7000922 | 2025-12-11 | RSUP dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0109 | Survey | O2 sensor faulty. Cek O2 sensor, housing sensor. Kabel O2 housing rusak | Pending |  | Wahyu
6000976 | 2025-12-11 | RSUD Cilacap | Mesin Anasthesia | Draeger/Perseus A500 | ASUS0001 | Service | Konfigurasi SIMRS dengan gateaway draeger | Ready to use |  | Dwi
3000921 | 2025-12-11 | RSUD Alimuddin Umar | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0402 | Survey | Generator tidak bisa on; IDB rusak | Pending |  | Imam, Bahar
5000385 | 2025-12-11 | RSKO Cibubur | Mesin Anasthesia | Draeger/Atlan A1000 XL | ASUH0427 | instalasi | Unit berfungsi dengan baik | Ready to use |  | Farhat, Ipan, Wahyu
8000775 | 2025-12-11 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Atlan A300 XL | ASTJ0118 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
9000673 | 2025-12-11 | RSD TIdore | Bubble CPAP | Draeger/Oxymixer | 212270 | Service | Penggantian outlet konektor O2 karna tidak cocok dengan outlet milik RS | Ready to use |  | Ruslan
12000287 | 2025-12-11 | RS GMIM Bethesda | Ventilator | Draeger/Savina 300 | ASTJ0022 | Service | Penggantian mixer | Ready to use |  | Marshall, Irfan R
10000251 | 2025-12-10 | Siloam Hospital Balikpapan | Hemodialisa | Bbraun/Dialog+ | - | Survey | Pengecekan RO milik RS | Ready to use |  | Adnan
12000286 | 2025-12-10 | RSUP Ratatotok Buyat | Laparascopy | Bbraun | 2530000 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Marshall, Irfan R
8000776 | 2025-12-10 | RSUD Dr. R. Soedjati Soemodiarjo | UPS | OAXAKA/AP33100 | 110045491FA3998400001 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000776 | 2025-12-10 | RSUD Dr. R. Soedjati Soemodiarjo | UPS | OAXAKA/AP33880 | 110048341FA3998100001 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
6000980 | 2025-12-10 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHK0151 | Survey | Valve ventilator fail, compliance test dan system test. Distrubuting valve perlu diganti | Pending |  | Dwi
6000981 | 2025-12-10 | RSUD Brebes | Ventilator | Draeger/Babylog VN800 | ASPH0003 | Survey | Unit mati total. Cek PSU, PSU rusak perlu diganti | Pending |  | Dwi
19000542 | 2025-12-10 | RSUD Budi Rahayu | ventilator | Draeger/Evita V600 | ASTD0042 | Maintenance | Cleaning unit, circuit test, running. Unit OK | Ready to use |  | Nanda
11000420 | 2025-12-10 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | ;0720017 | Survey | Suara EMH tidak normal. Kadang bunyi kadang tidak. Konektor EMH-HVG rusak | Pending |  | Nasrul
7000929 | 2025-12-09 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0092 | Survey | Cokcpit rusak. Reset M16, unit masih belum bisa | Pending | Coba ajukan peminjaman ke draeger pada board cockpit | Wahyu
7000923 | 2025-12-09 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASHL0055 | Service | Sensor check failure. Penggantian PBA VGC Analog 2, distrubuting valve, primus service set 2 tahun | Ready to use |  | Wahyu
9000676 | 2025-12-09 | RSUD Dr. H. Chasan Boesoirie | Pendant Surgery | SYAS | SY-PS2501008 | Service | Klip cover pada body pendant pecah. Penggantian klep baru | Ready to use |  | Ruslan
14000390 | 2025-12-09 | RSUD Gambiran Kediri | C-arm | Apelem/Evo R+ | 55121001850 | Survey | Unit error. Ada board yang terbakar | Pending |  | Fahim
17000224 | 2025-12-05 | RSK Dharmais | Digital Radiografi Fluoroskopi | Apelem/Platinum | L216F0664 | Service | Detector sering reboot, limit switch activated. Cek komunikasi detektor, cek pengaturan console. Unit sudah berfungsi kembali | Ready to use |  | Bahar
9000672 | 2025-12-05 | RSD TIdore | Bubble CPAP | Draeger/Oxymixer | 212270 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
18000665 | 2025-12-05 | RSUP Prof. Dr. R. D. Kandou Manado | Phototherapy | Draeger/Phototheraphy 4000 | ASJB0042 | Service | Penggantian lampu blue light | Ready to use |  | Irfan R
19000543 | 2025-12-05 | RSUP dr. Kariadi | Jaundice Meter | Draeger/JM-105 | 3505909 | service | Error 05. Kalibrasi dan penggantian baterai | Ready to use |  | Nanda
18000674 | 2025-12-05 | RSUD ODSK Provinsi Sulawesi Utara | Phototherapy | Draeger/ PT4000 | ASJB0042 | service | Penggantian lampu blue light | Ready to use |  | Irfan R
20000429 | 2025-12-05 | RS Regional dr. Hasri Ainun Habibie | Bubble CPAP | Draeger/Oxymixer | QM20246035 | service | Error 87. Reconnect power board dan control board. Unit sudah berfungsi kembali | Ready to use |  | Sujana
12000285 | 2025-12-04 | RSUD Ainun Habibie | Ventilator | Draeger/Evita V300 | ASLD0172 | service | Penggantian kit gas inlet counter | Ready to use |  | Marshall
18000673 | 2025-12-04 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista120S | VGSFE0108 | service | Penggantian part PCBA Mainboard | Ready to use |  | Irfan R
3000920 | 2025-12-04 | RSUP Dr. Mohammad Hoesin (KSO) | X-ray Digital | Apelem/Camargue DR | IMA704028 | service | Penggantian hardisk, backup data pasien | Ready to use |  | Imam, Sutan
18000669 | 2025-12-04 | RSUP Prof. Dr. R. D. Kandou Manado | Infant Warmer | Draeger/Resuscitaire | JY90107 | Survey | Error 05. Heater rusak | Pending |  | Irfan R
18000675 | 2025-12-04 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120S | VGSSFE0108 | service | Penggantian main PCBA dan display drive PCBA | Ready to use |  | Irfan R
12000284 | 2025-12-03 | RSUD Ainun Habibie | Bubble CPAP | Draeger/Oxymixer | QM20255071 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Marshall
11000419 | 2025-12-03 | RSUP Tadjuddin Chalid | Ventilator | Draeger/Savina 300 | ASJL0030 | Survey | Baterail internal hilang, VT tiadk terbaca. Device check; Disrankan oenggantian service set 6 tahun | Pending |  | Nasrul
6000967 | 2025-12-03 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0048 | service | O2 kalibrasi fail, power fail, flow sensor fail; Penggantian O2 sensor, spirolog flow sensor | Ready to use |  | Dwi
19000541 | 2025-12-03 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASUJ0096 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
18000672 | 2025-12-03 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN800 | ASRK0009 | Survey | Nebulizer error; Pressure regulator rusak, diagntikan dengan pressure regulator bekas punya RS | Ready to use | Sementara unit bisa digunakan | Irfan R
8000772 | 2025-12-03 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Atlan A300 | ASUJ0153 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000772 | 2025-12-03 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Atlan A300 | ASUJ0152 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
11000418 | 2025-12-02 | RSUD Makassar | Bubble CPAP | Draeger/Oxymixer | QM20246056 | service | Error E87. Pengecekan mainboard, electrical board dan sensor. Unit sudah berfungsi kembali | Ready to use |  | Nasrul
7000920 | 2025-12-02 | RSUD Merah Putih | Ventilator | Draeger/Babylog VN800 | ASTK0008 | service | Touchscreen tidak responsif. Hit software, kalibrasi touchscreen, cek kabel lcd. Reconnect kabel lcd | Ready to use |  | Wahyu
19000540 | 2025-12-02 | RSUP dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0322 | Survey | Pengukuran tekanan tidak akurat. Kalibrasi airway pressure; Running test | Ready to use |  | Nanda
19000540 | 2025-12-02 | RSUP dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0300 | Survey | Pengukuran tekanan tidak akurat. Kalibrasi airway pressure; Running test | Ready to use |  | Nanda
9000670 | 2025-12-02 | RSUD Dr. H. Chasan Boesoirie | MOT | SYAS | - | service | Pintu sliding macet. Setting posisi motor ke rel motor | Ready to use |  | Ruslan
20000428 | 2025-12-02 | RSUP Makassar | MOT | SYAS | - | service | Kabel power putus; Perbaikan kabel power kontrol sliding door | Ready to use |  | Sujana
19000562 | 2025-12-01 | RSUP Surakarta | USG | MIndray/DP20T | LK&B36001327 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
11000417 | 2025-12-01 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum | G248L0145 | service | Meja dan generator tidak bisa menyala. Cek tegangan dan kontrol panel. Pengecekan MCB 1. Unit sudah berfungsi normal | Ready to use |  | Nasrul
14000386 | 2025-12-01 | RSUD Dr. Moh. Saleh | Meja mayo | SYAS | 2511335113000400278 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Fahim
14000386 | 2025-12-01 | RSUD Dr. Moh. Saleh | Meja mayo | SYAS | 2511335113000400279 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Fahim
14000389 | 2025-12-01 | RSUD Dr. Moh. Saleh | Meja Trolley | SYAS | 2511691010000400678 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Fahim
14000389 | 2025-12-01 | RSUD Dr. Moh. Saleh | Meja Trolley | SYAS | 2511691010000400679 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Fahim
14000388 | 2025-12-01 | RSUD Dr. Moh. Saleh | Bed Patient | SYAS/ 3 crank | 251062213000400138 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Fahim
14000387 | 2025-12-01 | RSUD Dr. Moh. Saleh | Trolley Anasthesi | SYAS | 250936607000400024 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Fahim
18000663 | 2025-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0411 | Survey | Layar LCD ada flek hitam. Unit masih bisa berfungsi namun tetap disarankan diganti | Ready to use |  | Irfan Rusli
8000771 | 2025-12-01 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASHD0034 | Survey | Leakage system test fail. Cek Co2y, cek breathing system. Device ok | Ready to use |  | Amna
16000304 | 2025-12-01 | Signature Dental Care | Dental Unit | Planmeca/ Compact i5 | UPCV574532 | service | Kursi pasien tidak bisa naik. Pengecekan kabel sensor; Sensor aman dan perbaikan jalur kabel sensor | Ready to use |  | Ipan
19000548 | 2025-12-01 | RSUD Ibu Fatmawati Soekarno | Mesin Anasthesia | Draeger/Atlan A100 XL | ASUD0161 | service | Unit TSB15; Pasang motor piston | Ready to use |  | Nanda
6000948 | 2025-12-01 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0072 | service | Penggantian pressure regulator O2, air, N2O, valve | Ready to use |  | Dwi
6000966 | 2025-12-01 | RSUD Prambanan | Bubble CPAP | Draeger/Oxymixer | QM20230153 | service | Perbaikan kit kalibrasi | Ready to use |  | Dwi
5000404 | 2025-11-29 | RS Restu Ibu Balikpapan (KSO) | Hemodialisa | Bbraun/Dialogist Evolight | 54212 | service | Bicarbonate catridge tidak terpasang dengan benar. Bersihkan dan ganti filter jalur bicarbonate | Ready to use |  | Farhat, Adnan
9000663 | 2025-11-29 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671867 | Survey | Training user | Ready to use |  | Ruslan, Irfan
9000664 | 2025-11-29 | RSUD Dr. H. Chasan Boesoirie | Infus Pump | Bbraun/Infuse Somat Space P | 964927 | Survey | Training user | Ready to use |  | Ruslan, Irfan
9000669 | 2025-11-29 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASHJ0180 | service | Penggantian baterai internal | Ready to use |  | Ruslan
9000669 | 2025-11-29 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASHJ0179 | service | Penggantian baterai internal | Ready to use |  | Ruslan
9000669 | 2025-11-29 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASHJ0178 | service | Penggantian baterai internal | Ready to use |  | Ruslan
18000667 | 2025-11-29 | RSUD ODSK Provinsi Sulawesi Utara | Hemodialisa | Bbraun | 518856 | Survey | Error filter expiration; Penggantian filter endoroksin dan reset filter | Ready to use |  | Irfan R
18000667 | 2025-11-29 | RSUD ODSK Provinsi Sulawesi Utara | Hemodialisa | Bbraun | 518851 | Survey | Error filter expiration; Penggantian filter endoroksin dan reset filter | Ready to use |  | Irfan R
18000671 | 2025-11-28 | RS GMIM Bethesda | Ventilator | Draeger/Savina 300 | ASTJ-0025 | Survey | Error O2; Penggantian mixer. Unit masih belum bisa digunakan | Pending | Menunggu teknisi draeger untuk di cek | Irfan R
5000403 | 2025-11-28 | RSUP IKN | Instrument | Bbraun | - | Instalasi | Barang dan diterima dan sudah sesuai surat jalan. Kekurangan menunggu pengiriman | Ready to use |  | Farhat
7000919 | 2025-11-28 | RSUD Merah Putih | X-ray Digital | Apelem/Camargue DR | 703979 | Survey | Lampu kolimasi redup, pengecekan board power kolimator. Unit perlu penggantian board dan lampu colimator | Pending |  | Wahyu
14000385 | 2025-11-28 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 0720015 | service | Power supply rusak; Penggantian board power supply | Ready to use |  | Fahim
14000384 | 2025-11-27 | RSUD Gambiran Kediri | C-arm | Apelem/Evo R+ | 55121001850 | Survey | Error unavailable inverter power supply; PSU dibawa untuk di cek | Pending |  | Fahim
19000547 | 2025-11-27 | RSUD Ibu Fatmawati Soekarno | Mesin Anasthesia | Draeger/Atlan A100 XL | ASUD0161 | Survey | Training alat | Ready to use |  | Nanda
7000918 | 2025-11-27 | RSUD Nyi Ageng Serang | CT Scan | Siemens/Somatom | 698122072 | Survey | pc server lemot. bersihkan komponen driver graphic dan ram, cek hardisk. Unit masih error | Pending | Perlu pengecekan teknisi Siemens | Wahyu
8000768 | 2025-11-27 | RSI PKU Muhammadiyah Mayong | Ventilator | Draeger/Babylog VN600 | ASSK0028 | Survey | Alarm minute volume dan baterai check. System check; Penggantian neonatal flow sensor | Ready to use |  | Amna
11000412 | 2025-11-28 | RSUD Tenriawaru | Hemodialisa | Bbraun/Dialog + | 264402 | Service | Syringe holder rusak, DF pressure. Error water supply; Penggantian pressure regulator | Ready to use |  | Nasrul
20000426 | 2025-12-26 | RSUD Andi Makkasau | ESWL | Medispec/EM1000 | 0622173 | Service | Daya tembak lemah. Pengecekan connector EMH. Perbaikan konektor, unit sudah bisa digunakan kembali | Ready to use |  | Sujana
18000676 | 2025-12-26 | RS Mata Manado | Mesin Anasthesia | Draeger/Fabius Plus | ASME0021 | Maintenance | Preventive Maintenance Total Care | Ready to use |  | Irfan R
12000282 | 2025-11-26 | RSUD Aloei Saboe | Bed Patient | SYAS/Premium 3 crank | 251162445000400484 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Marshall
12000282 | 2025-11-26 | RSUD Aloei Saboe | Bed Patient | SYAS/Premium 3 crank | 251162445000400484 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Marshall
6000965 | 2025-11-26 | RSUD Cilacap | USG | Mindray/M9 | CCI-42005879 | Service | Penarikan unit demo untuk pengembalian ke Mindray | Ready to use |  | Dwi
3000918 | 2025-11-26 | RSUD Gabriel Manek | Digital Radiografi Fluoroskopi | Apelem/Platinum | E226F0723 | Service | Hasil gambar terdapat artefak; Penggantian detector | Ready to use | Detector menggunakan dari unit Papua | Imam, bahar
8000765 | 2025-11-26 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Primus | ASHM0112 | Survey | Boating failure, o2 sensor dan gas failure. Cek modul PGM dan Mobi, part tsb rusak | Pending |  | Amna
7000917 | 2025-11-26 | RSUP dr. Soeradji Tirtonegoro | Infant Warmer | Draeger/Babytherm8010 | ASMM0046 | Survey | Suhu tidak sesuai. Cek d18 d19 aman. Heater dan relay bagus. Post test | Ready to use |  | Wahyu
7000916 | 2025-11-26 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0091 | Survey | Cockpit restart. Cek error log; Penggantian baterai cmos dan update software cockpit | Ready to use |  | Wahyu
19000538 | 2025-11-26 | RSUP dr. Kariadi | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0116 | service | Expiration port leakage; Cleaning Co2y dan ganti breathing circuit | Ready to use |  | Nanda
5000402 | 2025-11-25 | RSUD Abdoel Wahab Sjahranie | Syringe Pump | Bbraun/Perfusor Compact Plus | 001323 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Farhat, Adnan
5000402 | 2025-11-25 | RSUD Abdoel Wahab Sjahranie | Syringe Pump | Bbraun/Perfusor Compact Plus | 001322 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Farhat, Adnan
5000402 | 2025-11-25 | RSUD Abdoel Wahab Sjahranie | Syringe Pump | Bbraun/Perfusor Compact Plus | 001324 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Farhat, Adnan
5000401 | 2025-11-25 | RSUD Abdoel Wahab Sjahranie | Infus Pump | Bbraun/Infusomat Compact Plus | 001283 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Farhat, Adnan
5000401 | 2025-11-25 | RSUD Abdoel Wahab Sjahranie | Infus Pump | Bbraun/Infusomat Compact Plus | 001284 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Farhat, Adnan
5000401 | 2025-11-25 | RSUD Abdoel Wahab Sjahranie | Infus Pump | Bbraun/Infusomat Compact Plus | 001282 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Farhat, Adnan
90000668 | 2025-11-25 | RSUD Sanana | Pendant Anasthesia | SYAS | SY-PA2501008 | Survey | Pengecekan kelengkapan unit baru. Alat sesuai packing list | Pending | Unit belum diinstall karna ruangan belum siap | Ruslan
16000303 | 2025-11-25 | Smileworks Dental Care | Dental Unit | Kavo/E50 | 30217530 | Survey | Handpiece Prophy (punya klinik diluar bawaan) rusak saat dicoba menggunakan bubuk pearls | Ready to use | Sudah diinfokan ke dokter | Ipan
19000536 | 2025-11-25 | RS Banyumanik 2 | Ventilator | Draeger/Savina 300 | ASPJ0110 | Survey | Cek unit | Pending | Perlu penggantian maintenance kit 6 tahun | Nanda
7000913 | 2025-11-25 | RSUD Wonosari | Ventilator | Draeger/Babylog VN800 | ASPF0004 | service | Internal gas sensor calibration failed. Device check; Kompressor rusak | Pending |  | Wahyu
20000425 | 2025-11-25 | RSUP Mata Makassar | Mesin Anasthesia | Draeger/Atlan A350 | ASUL0192 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana, Nasrul
15000325 | 2025-11-24 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-Ray | Apelem/Rafale EV30 | 009/15/01095 | service | Filament fault; Pembersihan seluruh board dengan contact cleaner | Ready to use |  | Sutan
7000910 | 2025-11-24 | RSUD Sleman | Ventilator | Draeger/Evita V800 | ASUL0018 | Instalasi | Alat berfungsi dengan baik | ready to use |  | Wahyu
8000769 | 2025-11-24 | RSI Arafah Rembang | Mesin Anasthesia | Draeger/Fabius Plus | ASUA0091 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
11000416 | 2025-11-24 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina 300 | ASJJ0409 | Survey | VT tidak sesuai setting. Cek flow sensor, cek kabel harnest; Kabel harnest rusak | Pending |  | Nasrul
18000670 | 2025-11-24 | RS Siloam Manado | Hemodialisa | Bbraun | 214217 | service | DF pressure error; Penggantian tinse bridge dan kopling | Ready to use |  | Irfan R
19000534 | 2025-11-24 | RSUD dr. Adhyatma | Vaporizer | Draeger/Vapor 2000 | ASHM1356 | Survey | Inlet bocor. O ring inlet ada sobekan | Pending | Perlu penggantian O ring | Nanda
19000535 | 2025-11-24 | RSUD dr. Adhyatma | Vaporizer | Draeger/Vapor 2000 | ASHM0137 | Service | Nilai settingan tidak sesuai. Cek flow analyser. Setting flow anaylser. Nilai setting sudah sesuai | ready to use |  | Nanda
11000415 | 2025-11-23 | RSKD Dadi | C-arm | Apelem/Evo R+ | 55120063849 | Service | Workstation tidak bisa hidup atau booting terus; Pengecekan koneksi, cpu dan bios. Unit sudah berfungsi kembali | Ready to use |  | Nasrul
9000662 | 2025-11-23 | RSUD Dr. H. Chasan Boesoirie | Laparascopy | Bbraun | 1023784 | Survey | Training user | ready to use |  | Ruslan, Irfan
7000909 | 2025-11-23 | RSUP dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | service | Patient gas measurement failed, lcd blank; Penggantian flow power dan conv kit display | ready to use |  | Wahyu
6000960 | 2025-11-22 | RSUD Brebes | USG | Mindray/Pesona I9 | KH9C3900026ST | Maintenance | Preventive Maintenance | Ready to use |  | Dwi
4000528 | 2025-11-21 | RSUP Dr. Hasan Sadikin | Bor tulang | Aesculap/ELAN4 | 2910 | Maintenance | Putaran bor tidak normal | Pending | Sudah di info teknisi bbraun untuk di cek lebih lanjut | Satria, Ipan
6000958 | 2025-11-21 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASNE0151 | Survey | Sensor check, leak system, compliance leak, O2 sensor fail | Pending | O2 sensor perlu diganti, flow sensor, sample line | Dwi
3000917 | 2025-11-21 | RSUD Dr. R. Soedjati Soemodiarjo | X-ray Digital | Apelem/DR Camargue | 251025003 | Service | Adjust kolimasi dan uji paparan. Pintu masih ada kebocoran | Ready to use |  | Imam, Bahar
3000916 | 2025-11-21 | RSUD Dr. R. Soedjati Soemodiarjo | Digital Radiografi Fluoroskopi | Apelem/Platinum | H248L0156 | Service | Adjust kolimasi dan uji paparan. Ruang operator tidak ada kebocoran | Ready to use |  | Imam, Bahar
4000528 | 2025-11-21 | RSUP Dr. Hasan Sadikin | Bor Tulang | Bbraun/GB891R | 2910 | Maintenance | Preventive Maintenance. Putaran bor tidak normal, charging unit dan baterai berfungsi dengan baik | Pending |  | Satria, Ipan
4000528 | 2025-11-21 | RSUP Dr. Hasan Sadikin | Bor Tulang | Bbraun/GB884R | 2946 | Maintenance | Preventive Maintenance. Putaran bor tidak normal, charging unit dan baterai berfungsi dengan baik | Pending |  | Satria, Ipan
4000528 | 2025-11-21 | RSUP Dr. Hasan Sadikin | Bor Tulang | Bbraun/GA344 | 3044 | Maintenance | Preventive Maintenance. Putaran bor tidak normal, charging unit dan baterai berfungsi dengan baik | Pending |  | Satria, Ipan
4000528 | 2025-11-21 | RSUP Dr. Hasan Sadikin | Bor Tulang | Bbraun/GB894R | 3241 | Maintenance | Preventive Maintenance. Putaran bor tidak normal, charging unit dan baterai berfungsi dengan baik | Pending |  | Satria, Ipan
8000767 | 2025-11-21 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/SBW4 | 250879101260400103 | Service | Side guard tidak berfungsi; Pemasangan pematuk side guard | Ready to use |  | Amna
19000533 | 2025-11-21 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0125 | Survey | NIBP leak. Cek alat, ganti manset dan unit masih bermasalah. | Pending |  | Nanda
15000324 | 2025-11-20 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Tube tidak bisa bergerak; Pembersihan sensor grid dan penurunan CB3 | Ready to use |  | Sutan
19000532 | 2025-11-20 | RSUD Dr. Gunawan Mangunkusomo Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSMK0125 | Service | Notif NIBP simulation menyala terus, tanggal central monitor; Setting alat, turn off NIBP simulation | Ready to use |  | Nanda
6000957 | 2025-11-20 | RSUD Dr.(H.C) Ir. Soekarno | USG | Mindray/ DC-30 | JM3-48006338 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
8000762 | 2025-11-19 | RSI Arafah Rembang | Mesin Anasthesia | Draeger/Fabius Plus | ASUA0091 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
6000956 | 2025-11-19 | RSUD Cilacap | USG | Mindray/M9 | CCI-51006326 | Service | Baterai terdeteksi terkunci, tidak bisa digunakan. Sudah ganti baterai namun tetap rusak; Penggantian mainboard backup | Ready to use |  | Dwi
7000915 | 2025-11-19 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0092 | Service | Setting tanggal dan update software | Ready to use |  | Wahyu
19000546 | 2025-11-19 | RSUD Ibu Fatmawati Soekarno | Mesin Anasthesia | Draeger/Atlan A100 XL | ASUD0161 | Instalasi | Unit berfungsi dengan baik dan aksesoris lengkap | Ready to use |  | Nanda
11000414 | 2025-11-18 | RSUP Dr. Wahidin Sudirohusodo | Bor tulang | Aesculap/ELAN4 | - | Survey | Ketebalan potongan tidak sesuai setting. Alat perlu ditarik untuk di service bbraun | Pending |  | Nasrul
11000413 | 2025-11-18 | RSUP Dr. Wahidin Sudirohusodo | Hemodialisa | Bbraun/Dialog+ | 85049 | Service | Blood leak; Melakukan disinfeksi prosedur | Ready to use |  | Nasrul
4000527 | 2025-11-18 | RSUD dr. Chasbullah Abdulmadjid | Laser Surgical | JenaSurgical/Multipulse HoPlus | J191150379 | Service | Demo alat | Ready to use |  | Satria, Ipan
7000908 | 2025-11-18 | RSUD Kota Yogyakarta | Ventilator | Draeger/Babylog VN800 | ASRK0005 | Maintenance | Preventive Maintenance | Ready to use | Penggantian kit air filter | Wahyu
9000666 | 2025-11-18 | RSUD Dr. H. Chasan Boesoirie | MOT | SYAS | - | Service | Lemari MOT tida bisa tertutup rapat, pengecekan switch pintu lemari dan terdapat salah satu switch rusak; Sementara menggunakan 1 switch | Ready to use | Menunggu switch baru dari KSS | Ruslan
18000680 | 2025-11-18 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0411 | Service | Penggantian oxy trace, O2 sensor, magnet valve dan main switch | Ready to use |  | Irfan R
18000680 | 2025-11-18 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0415 | Service | Penggantian oxy trace, O2 sensor, magnet valve dan main switch | Ready to use |  | Irfan R
18000678 | 2025-11-18 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0415 | Survey | Nebulizer, O2 sensor dan main switch rusak, penggantian main switch dan O2 sensor. Namun ada kerusakan pada baterai dan pressure regulator | Pending |  | Irfan R
12000278 | 2025-11-18 | RSUD Toto Kabila | MOT | SYAS | - | Service | Penggantian vigilohm | Ready to use | Penggantian spare part dilakukan oleh tim IPSRS | Marshall
4000527 | 2025-11-18 | RSUD dr. Chasbullah Abdulmadjid | Laser Surgical | JenaSurgical/Multipulse HoPlus | J191150379 | Survey | Pendampingan demo alat ke pasien | Ready to use |  | Satria, Ipan
8000760 | 2025-11-18 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHA0113 | Survey | Unloading backup spare part | Ready to use |  | Amna
9000661 | 2025-11-18 | RSUD Dr. H. Chasan Boesoirie | Laparascopy | Bbraun | 2507CE0162 | Survey | Training user | Ready to use |  | Ruslan
9000665 | 2025-11-18 | RSUD Dr. H. Chasan Boesoirie | Pendant Surgery | SYAS | SY-PS2501009 | Service | clip pada cover pendant patah/pecah; Penggantian clip/clem yang baru | Ready to use |  | Ruslan
19000531 | 2025-11-18 | RS Umi Barokah | Baby Incubator | Draeger/Isolette C2000 | DSRL0029 | Maintenance | Cleaning unit, running. Unit OK | Ready to use |  | Nanda
19000531 | 2025-11-18 | RS Umi Barokah | Baby Incubator | Draeger/Isolette C2000 | DSSK0008 | Maintenance | Cleaning unit, running. Unit OK | Ready to use |  | Nanda
200000421 | 2025-11-18 | RSUP Makassar | MOT | SYAS | - | Service | Penggantian lampu | Ready to use |  | Sujana
14000382 | 2025-11-17 | RSUD Cerati | Meja Mayo | SYAS/SCY1 | 25113511300400270 | Instalasi | Telah diinstall 8 unit | Ready to use |  | Fahim
7000907 | 2025-11-17 | RSUD R.A.A Tjokronegoro | Bubble CPAP | Draeger/Oxymixer | QM20255065 | Service | Humidifier C87; Voltage calibration | Ready to use |  | Wahyu
8000770 | 2025-11-17 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0027 | Instalasi | Installasi demo unit | Ready to use |  | Amna
6000955 | 2025-11-17 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASUJ0058 | Survey | Breathing system failure, flow sensor failure; Perlu penggantian Pba thermal control unit | Pending |  | Dwi
18000679 | 2025-11-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0165 | Service | Penggantian baterai set, main switch dan pressure regulator | Ready to use |  | Irfan R
18000679 | 2025-11-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0167 | Service | Penggantian baterai set, main switch dan pressure regulator | Ready to use |  | Irfan R
12000277 | 2025-11-14 | RSUD Hasri Ainun Habibie | MOT | SYAS | - | Survey | Pemantauan unit HVAC. Suhu -19 hingga 21, humidity 45-60% | Ready to use |  | Marshall
6000954 | 2025-11-14 | RS JIH | Jaundice Meter | Draeger/JM-105 | 3507946 | Service | Kalibrasi unit | Ready to use |  | Dwi
6000954 | 2025-11-14 | RS JIH | Jaundice Meter | Draeger/JM-105 | 3507945 | Service | Kalibrasi unit | Ready to use |  | Dwi
18000681 | 2025-11-14 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASEL0164 | Service | Penggantian baterai set, oxytrace/O2 sensor dan filter service set 1 tahun | Ready to use |  | Irfan R
18000681 | 2025-11-14 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASEL0166 | Service | Penggantian baterai set, oxytrace/O2 sensor dan filter service set 1 tahun | Ready to use |  | Irfan R
7000906 | 2025-11-13 | RSUP dr. SArdjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Service | Penggantian Conv kit display set | Ready to use |  | Wahyu
12000275 | 2025-11-13 | RSUD Otonaha | Baby Incubator | Draeger/Babylog VN800 | ASTE0041 | Survey | Setting EtCo2 lupa password, touchscreen delay; Reset password dan kalibrasi touchscreen | Ready to use |  | Marshall
12000276 | 2025-11-13 | RSUD Toto Kabila | MOT | SYAS | - | Survey | Vigilohm error; Pengajuan claim warranty | Pending |  | Marshall
3000915 | 2025-11-13 | RSUD Dr. R. Soedjati Soemodiarjo | Digital Radiografi Fluoroskopi | Apelem/Platinum | H248L0156 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Imam, Bahar, Nanda, Amna
3000914 | 2025-11-13 | RSUD Dr. R. Soedjati Soemodiarjo | X-ray Digital | Apelem/Camargue DR | 251025003 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Imam, Bahar, Nanda, Amna
6000964 | 2025-11-13 | RSUD dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASNB0071 | Survey | Sensor check failed, flow sensor failed; Cek mixer, pressure sensor. Penggantian flow sensor namun sensor check masih failed | Pending |  | Dwi
9000659 | 2025-11-12 | RSUD Dr. H. Chasan Boesoirie | Laparascopy | Bbraun | 2507CE0162 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan, Irfan
12000271 | 2025-11-12 | RSUD Hasri Ainun Habibie | MOT | SYAS | - | Service | Penambahan freon dan pemantauan unit. Suhu stabil 19 - 21, humidity 45-60% | Ready to use |  | Marshall
16000302 | 2025-11-12 | Klinik Gigi Drg. Maria | Handpiece | Kavo/637B | 07-1019721 | Service | Cap pada handpiece seret; Penggantian cap handpiece | Ready to use | Penggantian spare part di kantor | Ipan
12000274 | 2025-11-12 | RSUD Hasri Ainun Habibie | MOT | SYAS | - | Service | Perapihan wiring kabel kompressor dan pemantauan unit. Suhu stabil 19-21, dan humidity stabil 45-60% | Ready to use |  | Marshall
7000904 | 2025-11-12 | RSUP dr. Sardjito | Mesin Anasthesia | Draeger/Atlan A300 | ASTK0060 | Service | Penggantian ventilator kit assembly | Ready to use |  | Wahyu
11000411 | 2025-11-12 | RSUP Dr. Wahidin Sudirohusodo | Infant Warmer | Draeger/Babyleo TN500 | ASTH0140 | Service | Skin sensor error; Kalibrasi sensor | Ready to use |  | Nasrul
11000410 | 2025-11-11 | RSUP Makassar | Mobile X-Ray | Apelem/Rafale BDR | 87124084483 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Nasrul
9000657 | 2025-11-11 | RSUD Dr. H. Chasan Boesoirie | Bor Tulang | Bbraun/ELAN4 | 003988 | Survey | Handpiece tidak berfungsi dna banyak karatan. Pengecekan handpiece L7 dan eraniotoma; Handpiece dibawa untuk di cek bbraun | Pending |  | Ruslan, Irfan
8000759 | 2025-11-11 | RSUD Dr. R. Soedjati Soemodiarjo | USG | Mindray/Z50T | HH3E58000152 | Survey | Training alat | Ready to use |  | Amna
12000273 | 2025-11-11 | RSUD Hasri Ainun Habibie | MOT | SYAS | - | Survey | Pemantauan suhu dan humidity | Ready to use |  | Marshall
6000953 | 2025-11-11 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASFL0121 | Survey | Display mati total; Pengecekan mainboard, LCD, power supply. | Pending | Disarankan penggantian mainboard | Dwi
6000953 | 2025-11-11 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASFL0120 | Survey | Display mati total; Pengecekan mainboard, LCD, power supply. | Pending | Disarankan penggantian mainboard | Dwi
19000525 | 2025-11-11 | RSUP dr. Kariadi | Ventilator | Draeger/Savina 300 | ARXL0178 | Survey | Internal baterai failed, O2 failed. Cek flow sensor | Pending | Disarankan penggantian O2 sensor dan service set 6 tahun | Nanda
9000656 | 2025-11-10 | RSUD Dr. H. Chasan Boesoirie | Central Monitor | Draeger/CMS Vista 300 | CASUC0009 | Service | Unit tidak bisa konek CMS. Setting koneksi IP dan restart router | Ready to use |  | Ruslan, Irfan
6000952 | 2025-11-10 | RSU An Ni'mah | Bubble CPAP | Draeger/Oxymixer | QM20255060 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
6000952 | 2025-11-10 | RSU An Ni'mah | Bubble CPAP | Draeger/Oxymixer | QM20255074 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
7000903 | 2025-11-10 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASKE0069 | Survey | Cockpit restart. Cek error log. Internal baterai rusak | Pending |  | Wahyu
19000523 | 2025-11-10 | RSUP dr. Kariadi | Jaundice Meter | Draeger/JM-105 | H3505909 | Service | Penarikan unit untuk di kalibrasi dan penggantian baterai | Ready to use |  | Nanda
11000409 | 2025-11-07 | RSUP Dr. Wahidin Sudirohusodo | Digital Radiografi Fluoroskopi | Apelem/Platinum | G248L0145 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Nasrul
19000522 | 2025-11-07 | RSUD dr. Adhyatma | Baby Incubator | Draeger/Caleo | ASHK0030 | Service | Udara lembab tidak mengalir, conve heater water kit rusak; Ganti dengan unit ruang lain (ASHK0035) | Ready to use |  | Nanda
3000913 | 2025-11-07 | RSUP H. Adam Malik | Bone Mineral Densitometer | Medilink/Medix DR | J17016M280 | Survey | Generator fault. Cek tegangan trafo dan inverter, restrat password pc workstation; X-ray monoblock dan inverter rusak | Pending |  | Imam
17000223 | 2025-11-07 | RSUP Jayapura | Digital Radiografi Fluoroskopi | Apelem/Platinum | H24L0157 | Service | Artefak detector, DAP tidak terbaca; Mengganti detector baru dan generator control board, kalibrasi detector | Ready to use | Unit sudah berfungsi normal | Bahar, Eddy
6000951 | 2025-11-07 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0072 | Survey | Valve/Delivery sytem, internal system error. Cek monitor ptank, prp, adjustment O2. Pressure regulator rusak | Pending |  | Dwi
18000682 | 2025-11-07 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120S | S2SRF0253 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R, Ruslan
18000682 | 2025-11-07 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120S | S2SRF0243 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R, Ruslan
8000758 | 2025-11-07 | RSUD RAA Soewondo Pati | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/054/C1 | Survey | Pendampingan koneksi PACS ke unit | Ready to use |  | Amna
7000902 | 2025-11-07 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Evita V300 | ASPC0205 | Survey | Gas supply sensor calibration failed. Cek error log, HIT software. | Pending | Unit disarankan penggantian kit proportional valve, kit air filter, internal baterai NIMH | Wahyu
19000527 | 2025-11-06 | RSUD Ibu Fatmawati Soekarno | Mesin Anasthesia | Draeger/Atlan A100 XL | ASUD0161 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
7000901 | 2025-11-06 | RSUD Prambanan | Jaundice Meter | Draeger/JM-105 | H3500432 | Survey | Display rusak. Pengecekan LCD | Pending | Unit perlu penggantian LCD | Wahyu
3000912 | 2025-11-06 | RSU Haji Medan | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/033/C0 | Survey | Motor colimator rusak. Cek tegangan dan cek kolimator; Motorkolimator rusak dan tegangan drop saat expose | Ready to use | Unit secara fungsi masih bisa digunakan dengan stabilizer 15kVa | Imam, Sutan
3000911 | 2025-11-06 | RSU Haji Medan | Bone Mineral Densitometer | Medilink/Medix DR | D197H056 | Service | Penggantian monoblock dan kalibrasi filament | Ready to use |  | Imam, Sutan
8000757 | 2025-11-06 | RSUD Ki Ageng Selo | Infant Warmer | Draeger/Resuscitaire | AC13516 | Service | Error 05; Penggantian heater | Ready to use |  | Amna
20000419 | 2025-11-06 | RSUP Dr. Wahidin Sudirohusodo | Laser Surgical | JenaSurgical/Multipulse HoPlus | J1911501700 | Service | Pengecekan unit; Laser optic perlu diganti | Pending |  | Sujana, Nasrul
11000408 | 2025-11-05 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | SYAS/Premium Bed | 2407624040000400439 | Service | Tidak bisa naik turun. Cek engkol. Terlepas; Pemasangan engkol | Ready to use |  | Nasrul
19000530 | 2025-11-05 | RS Paru Ario Wirawan | Syringe Pump | Bbraun/Perfusor Compact Plus | 001199 | service | Unit error. Perbaikan oleh tim teknisi bbraun | Ready to use |  | Nanda
19000530 | 2025-11-05 | RS Paru Ario Wirawan | Syringe Pump | Bbraun/Perfusor Compact Plus | 001197 | service | Unit error. Perbaikan oleh tim teknisi bbraun | ready to use |  | Nanda
6000950 | 2025-11-05 | RS Dadi Keluarga | Ventilator | Draeger/Babylog VN800 | ASPF0030 | Survey | O2 supply, air supply, pressure sensor, O2 sensor tidak bisa. Kalibrasi O2, masih tidak bisa; Perlu penggantian PATO O2 sensor | Pending |  | Dwi
17000222 | 2025-11-05 | RSUP Jayapura | Digital Radiografi Fluoroskopi | Apelem/Platinum | H24L0157 | Service | DAP tidak terbaca, terdapat artefak pada detector; Remote oleh Apelem, unit masih belum ada perbaikan | Pending |  | Bahar, Eddy
16000301 | 2025-11-05 | RS Gading Pluit | Dental Unit | Kavo/1058 Life | 2004-1002866 | Survey | Suction tidak menyedot. Pengecekan selang suction; Selang suction mengkerut dibagian dalam | Pending | Selang suction perlu penggantian | Ipan, Farhat
12000267 | 2025-11-05 | RSUD Hasri Ainun Habibie | MOT | SYAS | - | Survey | Pemantauan suhu dan humidity | Ready to use |  | Marshall
8000756 | 2025-11-05 | RSUD RAA Soewondo Pati | Patient Monitor | Draeger/Vista 300 | VSSTE0128 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
7000794 | 2025-11-05 | RS Panti Rapih | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ0084 | Service | Keypad stuck; Penggantian front panel keypad | Ready to use |  | Wahyu
12000260 | 2025-11-04 | RSUD Hasri Ainun Habibie | MOT | SYAS | - | Service | Pembersihan Pro filter dan outdoor unit | Ready to use |  | Marshall
6000949 | 2025-11-04 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0010 | Service | Problem electronic; Penggantian Pba mainboard, waterlock dan sample line | Ready to use |  | Dwi
19000521 | 2025-11-04 | RSUD dr. Adhyatma | Vaporizer | Draeger/Vapor 2000 | ASHM1356 | Survey | Vapo bocor bagian inlet, kemungkinan seal rusak | Pending | Unit belum bisa digunakan | Nanda
19000520 | 2025-11-04 | RSUD dr. Adhyatma | Mesin Anasthesia | Draeger/Primus | ASHH0122 | Service | APL valve failed. Cek unit, tarik error, device oke | Ready to use |  | Nanda
15000322 | 2025-11-04 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Tube tidak bisa bergerak. Bersihkan sensor pada panel tegangan | Ready to use |  | Sutan
8000754 | 2025-11-04 | RSUD Dr. R. Soedjati Soemodiarjo | USG | Mindray/Z50T | HH3E58000152 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
20000416 | 2025-11-03 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0145 | Service | Penggantian push button | Ready to use |  | Sujana
5000384 | 2025-11-03 | Signature Dental Care | Dental Unit | Kavo/1058 Life | 10207191 | Service | Pengambilan scaler, selang, dan socket dental unit untuk backup rsud jailolo | Pending |  | Farhat
7000793 | 2025-11-03 | RSUP dr. Sardjito | Bor tulang | Bbraun | 7463 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Wahyu
6000963 | 2025-11-03 | RSUD CIlacap | USG | Mindray/M9 | CCI-42005879 | Survey | Probe tidak terdeteksi. Reset konektor probe, reset default setting | Ready to use |  | Dwi
19000519 | 2025-11-03 | RS Telogorejo Semarang | Mesin Anasthesia | Draeger/Primus | ASDK0223 | Service | Sensor check failure. Tarik error, sensor PTank rusak. Ganti sensor PTank yang baru | Ready to use |  | Nanda
3000910 | 2025-11-03 | RS Immanuel Bandung | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F045 | Maintenance | Preventive Maintenance | Ready to use | Ada kerusakan pada joystick console lifter colimator | Imam
8000753 | 2025-11-03 | RSJD Amino Gondohutomo | Ventilator | Draeger/Savina 300 | ASMK0306 | Survey | Nebulizer tidak bisa, inp/exp tidak berfungsi. System check | Pending | Unit perlu penggantian savina service set 1 tahun | Amna
12000266 | 2025-11-03 | RSUD Hasri Ainun Habibie | MOT | SYAS | - | Survey | Pemantauan HVAC 1x24 jam | Ready to use |  | Marshall
13000206 | 2025-11-03 | RSUP Dr. M. Djamil | MOT | SYAS | - | Service | Perbaikan karet lengan pendant OK 8 | Ready to use |  | Irfan Izzudin
11000407 | 2025-11-03 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum | G248L0145 | Service | Console tidak bisa ON. Cek console, UPS, pc console. Fuse putus; Penggantian fuse console | Ready to use |  | Nasrul, Sujana
13000205 | 2025-11-02 | RSUP Dr. M. Djamil | MOT | SYAS | - | Survey | MOT Pendant OK 8 ada karet lengan atas kiri yang rusak (bocor) | Pending |  | Irfan Izzudin
12000263 | 2025-11-01 | RSUD Hasri Ainun Habibie | MOT | SYAS | - | Service | Ac tidak dingin; Pembersihan pre filter, outdoor unit dan penyesuaian freon sesuai standar | Ready to use | Suhu 20 dan humidity 50-55% | Marshall
20000427 | 2025-11-01 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASND0141 | Survey | Neonatal flow error; Perlu ganti baterai, kit proportional valve, maintenance kit | Pending |  | Sujana
20000427 | 2025-11-01 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASEF0068 | Survey | Unit error; Perlu penggantian kit proportional valve, maintenance kit dan mainetanance expiration valve | Pending |  | Sujana
5000383 | 2025-10-31 | FKG Trisakti | Handpiece | Kavo/Smart Tourque Lux | 1045783 | Survey | Mata bur tidak berfungsi. Perlu penggantian bearing | Pending |  | Farhat, Ipan
8000751 | 2025-10-31 | RSUD Dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Atlan A350 | ASUK0231 | Service | Pendampingan ke pasien dan pemasangan konektor central gas | Ready to use |  | amna
9000654 | 2025-10-31 | RSD Tidore | Ventilator | Draeger/Babylog VN500 | ASJK0042 | Service | Device check fail, air supply fail, O2 supply fail, sensor O2 fail; Penggantian kit proportional valve, Pato SA, neonat, flow sensor | Ready to use |  | Ruslan
19000518 | 2025-10-31 | RSUD Dr. Soeratno Gemolong | Jaundice Meter | Draeger/JM-105 | H3503803 | Service | Baterai internal error 05. Penggantian baterai | Ready to use |  | Nanda
20000415 | 2025-10-30 | RSUD Madani Palu | Ventilator | Draeger/Savina 300 | ASSL0013 | Maintenance | Preventive Maintenance | Ready to use |  | Sujana
20000414 | 2025-10-30 | RSUD Madani Palu | Ventilator | Draeger/Babylog VN800 | ASTA0015 | Service | Error supply O2. Self test, system check. Kalibrasi; Unit sudah bisa digunakan kembali | Ready to use |  | Sujana
5000382 | 30 Oktober 2025 | Smileworks Dental Care | Dental Unit | Kavo/E50 | 30212530 | Service | Tukar handpiece backup dan pemindahan posisi threeway syringe | Ready to use |  | Farhat, Ipan
19000526 | 2025-10-30 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000526 | 2025-10-30 | RS JIH | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0125 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
12000262 | 2025-10-30 | RSUD Hasri Ainun Habibie | MOT | SYAS | - | Survey | HVAC dipantau di ruang OK Mata. Suhu dingin, tegangan normal 384, high pross 260 dan low pross 43 | Ready to use | Unit berfungsi normal | Marshall
3000908 | 2025-10-30 | RSUD Dr. R. Soedjati Soemodiarjo | Digital Radiografi Fluoroskopi | Apelem/Platinum | H248L0156 | Instalasi | Unit sudah terinstall, namun belum uji fungsi karna UPS belum diinstall dan ruang operator masih belum siap | Pending |  | Imam, Satria, Amna
17000221 | 2025-10-30 | RS Otak Moh. Hatta | C-arm | Apelem/Evo R+ FP | 010224003012 | Service | Install demo alat untuk training. Hasil gambar kurang jelas, kalibrasi detektor | Ready to use |  | Bahar, Irfan I
07000792 | 2025-10-30 | RSUP dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Service | Patient gas measurement fault; Penggantian flow power | Ready to use |  | Wahyu
06000945 | 2025-10-30 | RSUD Dr. Soeselo Slawi | Lampu Operasi | SYAS/Lumina | OL9680/80240100055 | Service | Handle panel lampu terlepas karna kendur baut, panel display monitor kendur; Pengencangan baut, penggantian pasbox 2 pcs dan restart communication panel layar | Ready to use |  | Dwi
11000406 | 2025-10-30 | RSUP Makassar | Mobile X-Ray | Apelem/Rafale BDR | 97124084483 | Service | Detector disconnected; Cek dan setting WAP, d-link dan cek detector | Ready to use | Unit sudah berfungsi normal kembali | Nasrul
11000405 | 2025-10-29 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0034 | Service | Sensor failure, engsel patah. Cek sensor board. Unit bisa berfungsi dengan kondisi engsel patah | Ready to use |  | Nasrul
12000261 | 2025-10-29 | RSUD Hasri Ainun Habibie | MOT | SYAS | - | Survey | AC tidak dingin, MCB kompresor gak berfungsi kabel RST dan UVW terbakar | Pending |  | Marshall
07000791 | 2025-10-29 | RSUD Saras Adyatma | Patient Monitor | Draeger/Vista 120 | VGSTC0078 | Service | Unit mati total; Penggantian power PCBA | Ready to use |  | Wahyu
06000943 | 2025-10-29 | RSUD Ajibarang | Ventilator | Draeger/Babylog VN500 | ASMK0054 | Survey | O2 supply, air supply, O2 sensor nebulizer failed. Self test dan hit software. | Pending | Disarankan penggantian PATO2, pressure regulator, baterai pack, proportional valve | Dwi
06000941 | 2025-10-29 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASHA0842 | Survey | Fail flow sensor; Penggantian flow sensor adult | Ready to use |  | Dwi
8000743 | 2025-10-29 | RSUD Dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Atlan A350 | ASUK0231 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
9000653 | 2025-10-29 | RSUD Dr. H. Chasan Boesoirie | Bor | Aesculap/ELAN4 | 003988 | Survey | Ring handpiece L7 dan craniotomic rusak berkarat; Handpiece dikirim ke bbraun untuk diperbaiki | Pending |  | Ruslan
18000690 | 2025-10-29 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Fabius Plus | ASFM0036 | Service | Batterai drop dan O2 sensor expired; Penggantian battery pack dan o2 sensor | Ready to use |  | Irfan R
18000690 | 2025-10-29 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Fabius Plus | ASJK0016 | Service | Batterai drop dan O2 sensor expired; Penggantian battery pack dan o2 sensor | Ready to use |  | Irfan R
16000300 | 2025-10-28 | FKG Moestopo | Dental Simulator | Kavo/DSE Compact | 2019-11-00000855 | Survey | Lampu patah. Sambung antara lampu dan lengan yang patah sementara | Pending |  | Farhat, Ipan
12000259 | 2025-10-28 | RSUD Tani dan Nelayan | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/181/C0 | Survey | Test menyala dan expose ukur tegangan karna high kv tidak bisa; Tegangan rs tidak stabil butuh stabilizer 20 kVa | Ready to use |  | Marshall
07000790 | 2025-10-28 | RSUP dr. Soeradji Tirtonegoro | Jaundice Meter | Draeger/JM-105 | B3502776 | Service | Kalibrasi internal, error 05; Penggantian baterai JM105 | Ready to use |  | Wahyu
07000789 | 2025-10-28 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | V3SHK0157 | Service | NIBP failed; Penggantian X2 NIBP dan SPO2 module | Ready to use |  | Wahyu
06000944 | 2025-10-28 | RSUD Brebes | Baby Incubator | Draeger/Isolette C200 | DSNM0005 | Service | Pemasangan evaportizer valve 230 V | Ready to use |  | Dwi
07000795 | 2025-10-27 | RSUD Kota Yogyakarta | Ventilator | Draeger/Babylog VN800 | ASRK0005 | Service | O2 measurement failed; Penggantian O2 sensor | Ready to use |  | Wahyu
9000652 | 2025-10-27 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | 2510220561 | Service | Buckystand turun saat unit mati dan ada bunyi saat naik turun; Menambahkan pemberat dan lapisi karet roda pemberat | Ready to use |  | Ruslan
18000692 | 2025-10-27 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1598 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000692 | 2025-10-27 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1467 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000692 | 2025-10-27 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1479 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000692 | 2025-10-27 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1473 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000692 | 2025-10-27 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1489 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000692 | 2025-10-27 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1500 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000692 | 2025-10-27 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1591 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000692 | 2025-10-27 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1545 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000692 | 2025-10-27 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1470 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
19000517 | 2025-10-27 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Perseus A500 | ASNK0040 | Service | Ganti board PCB. Cek software, reload. Cek touchscreen. Unit sudah dikalibrasi touchscreen | Ready to use |  | Nanda
19000516 | 2025-10-27 | RSJ Prof Dr Soerojo Magelang | Incubator Transport | Draeger/TI500 | CH12771 | Survey | PSU fuse putus terus. Cek psu, gsnti fuse namun unit tetap mati; PSU dan baterai internal rusak | Pending |  | Nanda
20000413 | 2025-10-27 | RSUP Makassar | Mobile X-Ray | Apelem/Rafale BDR | 97124084483 | Service | Penyambungan unit dengan PACS dan RIS | Ready to use |  | Sujana, Nasrul
3000909 | 2025-10-26 | RSUD Dr. R. Soedjati Soemodiarjo | X-ray Digital | Apelem/Camargue DR | - | Pra Install | Besi sudah terpasang | Pending | Menunggu unit datang | Imam, Satria, Amna
19000537 | 2025-10-25 | RSUD Ibu Fatmawati Soekarno | Hospital Furniture | SYAS | - | Instalasi | Installasi 8 unit emergency trolley, 20 instrument trolley, 4 unit medicine trolley, food trolley 2 unit, kursi roda 6 unit, mayo stand 4 unit, standard examination 3 unit | Ready to use |  | Nanda
14000380 | 2025-10-24 | RSUP Surabaya | MOT | SYAS | - | Survey | Pengecekan unit untuk pelaporan karna scrub belum terinstall | Pending |  | Fahim
16000299 | 2025-10-24 | Kusumah Dental Care | Dental Unit | Kavo/E30TM | 10111157 | Survey | Nilai rpm pada mikromotor; Cek unit. Mikromotor normal dan nilai rpm ada di display | Ready to use | Unit tidak ada kendala | Ipan
8000742 | 2025-10-24 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | service | Appnea pressure ventilator fail; Cek unit, running test | Ready to use |  | Amna
9000720 | 2025-10-24 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0011 | Survey | Gas supply fail, air supply fail. System check dan tarik error; Board E-set PCB dan firmware GS500 bermasalah | pending | Diajukan claim warranty | Ruslan
9000720 | 2025-10-24 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0012 | Survey | Gas supply fail, air supply fail. System check dan tarik error; Board E-set PCB dan firmware GS500 bermasalah | pending | Diajukan claim warranty | Ruslan
18000683 | 2025-10-24 | RSUP Prof. Dr. R. D. Kandou Manado | Bor tulang | Aesculap/ELAN4 | 008036 | service | Wadah handpiece terlepas, terdapat baut yang hilang; Pemasangan baut baru | Ready to use |  | Irfan R
15000321 | 2025-10-24 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Software duet tidak bisa terbuka; ganti hardisk yang sudah ada software | Ready to use |  | Sutan
20000424 | 2025-10-24 | RSUP Makassar | Lampu Operasi | SYAS | - | Service | Camera error; Re connect adaptor | Ready to use |  | Sujana
13000204 | 2025-10-23 | RSUP Dr. M. Djamil | MOT | SYAS | - | Service | Vacuum suction tidak menyedot. Memberikan udara tekan pada saluran pipa vacuum yang tersumbat. Vacuum suction sudah berfungsi dengan baik | Ready to use |  | Irfan Izzudin
4000526 | 2025-10-23 | RSUD Welas Asih | Bor Ortho | Bbraun/ Elan4 | 008095 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Satria, Bahar
07000788 | 2025-10-23 | RSUD Muntilan | Mesin Anasthesia | Draeger/Perseus A500 | ASRH0184 | Survey | Gas measurement not available, Device check dan check errir; PA Sensor sudah lemah | Pending | Disarankan penggantian PA sensor, battery pack, breathassy service set 2 tahun | Wahyu
2000209 | 2025-10-23 | RSUD Gabriel Manek | Digital Radiografi Fluoroskopi | Apelem/Platinum | E226F073 | Survey | Artefak image detector. Cek software, cek image detector, kalibrasi detector; Update software version Adam 1.8 namun detector tetap ada artefak | Pending | Disarankan penggantian detector | Eddy
06000942 | 2025-10-23 | RSUD Cilacap | USG | Mindray/M9 | CCI-51006326 | Survey | Pengecekan dengan penggantian mainboard M9 dan unit berfungsi | Pending | Unit akan claim warranty mainboard | Dwi
8000746 | 2025-10-23 | RSUD Dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Atlan A350 | ASUK0231 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
9000750 | 2025-10-23 | RSUD Jailolo Halmahera Barat | UPS | OAXAKA/AP3330 | 1100045034F65393600001 | Instalasi | Unit sudah diinstall namun belum diuji fungsi karna panel induk belum ada | Pending |  | Ruslan
11000404 | 2025-10-23 | RSUP Tadjuddin Chalid | Infant Warmer | Draeger/Babytherm 8004 | ASDH0038 | Survey | Error 55. Cek board, heater dan keypad; Keypad night light rusak | Pending |  | Nasrul
19000514 | 2025-10-23 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Perseus A500 | ASNK0040 | Survey | Layar mati. Cek lcd, cek touchscreen. Kalibrasi touchscreen | Pending | Unit belum bisa | Nanda
19000513 | 2025-10-22 | RSUD Tidar Magelang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0496 | Service | Meja tidak bisa naik turun. Penggantian resistor break | Ready to use |  | Nanda
18000684 | 2025-10-22 | RSU GMIM Pancaran Kasih Manado | Ventilator | Draeger/Savina 300 | ASTJ0019 | service | Flow sensor sering tidak terbuka; Pembersihan expirasi valve | Ready to use |  | Irfan R
19000512 | 2025-10-21 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | service | Selang sering bocor. Cek sambungan selang dan ganti selang ke yang lebih keras dan standar pabrik | Ready to use |  | Nanda
07000780 | 2025-10-21 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0091 | Survey | Nebulizer failed, self test; Pressure regulator lemah | Pending | Perlu penggantian pressure regulator | Wahyu
8000740 | 2025-10-21 | RSI PKU Muhammadiyah Pekajangan | Mesin Anasthesia | Draeger/Fabius Plus XL | ASUE0013 | Survey | Training alat | Ready to use |  | Amna
18000685 | 2025-10-21 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0182 | Survey | Unit mati sendiri. Pengecekan power supply dan baterai; Power suppy dan baterai drop, perlu diganti | Pending |  | Irfan r
15000320 | 2025-10-21 | RSUP Dr. Mohammad Hoesin (KSO) | Printer | Fuji/Dryfix | 66431571 | service | Tidak bisa print, di cek ada film yang nyangkut | Ready to use |  | Sutan
13000203 | 2025-10-21 | RSUP Dr. M. Djamil | MOT | SYAS | - | Survey | Vacuum suction MOT pendant OK 2,3,4,7 kenadla dalam menyedot; Penggantian Pipa | Pending | Masih dalam tahap perbaikan | Irfan Izzudin
13000202 | 2025-10-20 | RSUP Dr. M. Djamil | MOT | SYAS | - | Survey | Pengecekan gas system, server, hepafilter, HVAC dan panel listrik. Unit berfungsi dengn baik | Ready to use |  | Irfan Izzudin
19000511 | 2025-10-20 | RSUD Ibu Fatmawati Soekarno | Infant Warmer | Draeger/Resuscitaire | DSNL0046 | service | Error 05. Cek heater, ternyata putus; Sementara heater disambung | Ready to use |  | Nanda
11000403 | 2025-10-20 | RSUP Makassar | Ventilator | Draeger/Evita V300 | ASDM0031 | Service | VTE tidak tampil, cek setting parameter; Edukasi user untuk setting parameter yang benar | Ready to use |  | Nasrul
06000939 | 2025-10-20 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASUJ0058 | Instalasi | Unit berfungsi dengan baik dan aksesoris lengkap | Ready to use |  | Dwi
18000691 | 2025-10-20 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1598 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000691 | 2025-10-20 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1467 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000691 | 2025-10-20 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1479 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000691 | 2025-10-20 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1473 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000691 | 2025-10-20 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1489 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000691 | 2025-10-20 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1500 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000691 | 2025-10-20 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1591 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000691 | 2025-10-20 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1545 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000691 | 2025-10-20 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | SBSNF1470 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R
14000379 | 2025-10-18 | RSUP Surabaya | MOT | SYAS | - | Service | Pintu tidak berfungsi dengan baik, roda anjlok keluar rel; Perbaikan roda dan jalur rel | Ready to use |  | Fahim
14000378 | 2025-10-18 | RSUP Surabaya | MOT | SYAS | - | Service | Passbox tidak berfungsi interlock nya; Penggantian program installasi | Ready to use |  | Fahim
15000319 | 2025-10-17 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Tidak bisa expose. Cek kabel tombol expose, ada yang lepas. Solder kabel yang terlepas | Ready to use |  | Sutan
8000739 | 2025-10-17 | RSI PKU Muhammadiyah Pekajangan | Mesin Anasthesia | Draeger/Fabius plus XL | ASUE0013 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
9000748 | 2025-10-17 | RSD Tidore | UPS | OAXAKA/EPTRAFO-3120 | 110045765F80312600001 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000748 | 2025-10-17 | RSD Tidore | UPS | OAXAKA/EPTRAFO-3120 | 110045765F80312600002 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
20000409 | 2025-10-17 | RSUP Dr. Wahidin Sudirohusodo | Syringe Pump | Bbraun/Perfusor Space | 716608 | Survey | Baterai error; Perlu penggantian baterai | Pending |  | Sujana
20000409 | 2025-10-17 | RSUP Dr. Wahidin Sudirohusodo | Syringe Pump | Bbraun/Perfusor Space | 716609 | Survey | Baterai error; Perlu penggantian baterai | Pending |  | Sujana
16000298 | 2025-10-16 | Smileworks Dental Care | Dental Unit | Kavo/E50 | 30212530 | Service | Setting RPM dan tekananan angin handpiece, dari 300.000 RPM (3bar) menjadi 3.4 Bar | Ready to use |  | Ipan
16000297 | 2025-10-16 | Smileworks Dental Care | Dental Unit | Kavo/E50 | 30212529 | Service | Setting RPM dan tekananan angin handpiece, dari 300.000 RPM (3bar) menjadi 3.4 Bar | Ready to use |  | Ipan
20000408 | 2025-10-16 | RS Mata Makassar | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0125 | Instalasi | Unit sudah diinstallasi namun belum uji fungsi | Pending |  | Sujana, Nasrul
04000524 | 2025-10-16 | Signature Dental Care | Dental Simulator | EDE/SimEx | SO270003 | Service | Installasi alat demo | Ready to use |  | Satria
18000693 | 2025-10-15 | RSUD ODSK Provinsi Sulawesi Utara | Bubble CPAP | Draeger/Oxymixer | QMHF2022208431 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000693 | 2025-10-15 | RSUD ODSK Provinsi Sulawesi Utara | Bubble CPAP | Draeger/Oxymixer | QMHF2022208433 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
12000258 | 2025-10-15 | RSUD Samratulangi Tondano | Dental Unit | Kavo/1058 Life | 10205322 | Survey | Board dentist rusak | Pending |  | Marshall
19000505 | 2025-10-14 | RSUD Dr. Soediran Mangun Sumarso | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/134/C2 | Survey | Kolimasi tidak pas; Setting kolimasi | Ready to use |  | Nanda, Imam
6000938 | 2025-10-14 | RSU An Ni'mah | Bubble CPAP | Draeger/Oxymixer | QM20255060 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
6000938 | 2025-10-14 | RSU An Ni'mah | Bubble CPAP | Draeger/Oxymixer | QM20255074 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
20000407 | 2025-10-14 | RSUP Tadjuddin Chalid | ESWL | Medispec/EM1000 | 0922014 | Survey | Pengecekan bubble pada EMH unit. Pengecekan sirkulasi; Pengurasan air pada EMH | Ready to use |  | Sujana
11000401 | 2025-10-13 | RSUD Labuang Baji | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0494 | Service | Meja tidak bisa tidur dan software loading lambat. Cek meja dan joystick & hapus database | Ready to use |  | Nasrul
07000786 | 2025-10-13 | RSUP dr. Sardjito | Mesin Anasthesia | Draeger/Atlan A300 | ASTK0060 | Survey | Apnea flow failed, Cek error log dan HIT software | Ready to use |  | Wahyu
19000510 | 2025-10-13 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Service | Penggantian EMH | Ready to use |  | Nanda
07000785 | 2025-10-13 | RSA UGM | Bubble CPAP | Draeger/Oxymixer | QM20215539 | Survey | Fraksi oksigen tidak sesuai, cek gas supply dan valve serta top tube; Top tube rusak | Pending | Disarankan penggantian top tube | Wahyu
6000946 | 2025-10-13 | RSUD Cilacap | USG | Mindray/M9 | CCI-51006326 | Survey | Baterai unit error; Pemasangan baterai warranty. Unit masih error | Pending | Menunggu jadwal visit teknisi Mindray | Dwi
8000737 | 2025-10-13 | RS Mardi Rahayu Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | Service | Alarm ventilator fail, piston tidak bekerja. Cek APL, CO2y, system check leakage, kalibrasi motor pump; Unit berfungsi kembali | Ready to use |  | Amna
20000406 | 2025-10-13 | RSUP makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum | G248L0145 | Service | Perbaikan AEC | Ready to use |  | Sujana, Nasrul, Bahar, Satria
18000688 | 2025-10-13 | RSUP Prof. Dr. R. D. Kandou Manado | Bor tulang | Bbraun/ELAN4 | 008317 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan R, Marshall
6000936 | 2025-10-11 | RSUD Hj. Anna Lesmanah | Surgical Microscope | Zeiss/Opmi Lumera 700 | 6634144809 | Survey | Lampu source microscope mati, fitting lampu terbakar meleleh | Pending | Disrankan penggantian fitting lampu 2 pcs | Dwi
20000405 | 10 Ocotber 2025 | RSUD Kota Makassar | Bone Mineral Densitometer | Medilink/Medix DR | 460624 | Survey | Generator fault; Block inverter dan monoblock rusak | Pending |  | Sujana, Nasrul
07000784 | 2025-10-10 | RSUP dr. Soeradji Tirtonegoro | Infant Warmer | Draeger/Resuscitaire | KH12932 | Survey | Resuscitaire bocor, cek control valve; Mini 5 way bocor | Pending | Perlu penggantian Gantle valve mini 5 way | Wahyu
19000504 | 2025-10-10 | RS Indriati Boyolali | Ventilator | Draeger/Savina 300 | ASNL0485 | Survey | Mode pv vt tidak naik. Self test, cek membrane ada sobek | Pending | Disarankan penggantian membrane | Nanda
6000937 | 2025-10-10 | RSUD Hj. Anna Lesmanah | Mesin Anasthesia | Draeger/Primus | ASLD0042 | Maintenance | Preventive Maintenance | Ready to use |  | Dwi
20000404 | 2025-10-09 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum | G248L0145 | Survey | Error monitor gagal booting. Pengecekan pc unit | Ready to use |  | Sujana
11000402 | 2025-10-09 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0039 | Survey | PAW tidak terbaca. Cek kebocoran, system test, cek mainboard | Pending | perlu penggantian mainboard dan pressure sensor | Nasrul
07000783 | 2025-10-09 | RSUD Muntilan | Mesin Anasthesia | Draeger/Perseus A500 | ASRH0184 | Service | Verification O2 sensor failed, Cek sample line dan waterlock; Penggantian sample line dan waterlock | ready to use |  | Wahyu
16000296 | 2025-10-09 | Signature Dental Care | Dental Unit | Planmeca Compact 1 | UPCV574531 | Service | Air handpiece sering nyala mati sendiri dan operating light berat; Setting air handpiece agar foot control tidak menyala mati otomatis dan setting operating light arm | Ready to use |  | Ipan
3000906 | 2025-10-09 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Imam
20000403 | 2025-10-08 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Primus | ASHE0186 | Survey | Error ETCO2. Tarik error. ILCA rusak | Pending |  | Sujana, Nasrul
19000508 | 2025-10-08 | RSUD Dr. Gunawan Mangunkusomo Ambarawa | Incubator Transport | Draeger/TI500 | KY11400 | Survey | Baterai internal drop. Cek supply baterai, hanya 9.5VDC harusnya 12 VDC | Pending | Disarankan penggantian baterai | Nanda
11000450 | 2025-10-08 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | SYAS/Premium 3 motor | 240762445000400476 | Survey | Remote rusak; Penggantian remote | Ready to use |  | Nasrul
05000381 | 2025-10-08 | Smileworks Dental Care | Dental Unit | Kavo/E50 | 30212529 | Survey | Handpiece kurang kencang; Pengecekan compressor dan handpiece lain. | Ready to use |  | Farhat, Ipan, Satria
6000935 | 2025-10-08 | RSUD dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASNB0071 | Survey | Power fail error, baterai capacity low | Pending | Disarankan penggantian service set 3 tahun dan breathassy kit 2 tahun | Dwi
11000449 | 2025-10-07 | RSUD Kota Makassar | Mesin Anasthesia | Draeger/Perseus A500 | ASJD0033 | Survey | Gas sensor fail, O2 sensor fail; System test, hit osftware. Sensor ILCA rusak | Pending | Disarankan penggantian sensor ILCA | Nasrul
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1162 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1163 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1164 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1165 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1166 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1167 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1168 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1169 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1170 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1171 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1172 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1173 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1174 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000689 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Infus Pump | Bbraun/Infusomat Compact | 1175 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000687 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Syringe Pump | Bbraun/Perfusor Compact | 1148 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000687 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Syringe Pump | Bbraun/Perfusor Compact | 1149 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000687 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Syringe Pump | Bbraun/Perfusor Compact | 1150 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000687 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Syringe Pump | Bbraun/Perfusor Compact | 1151 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000687 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Syringe Pump | Bbraun/Perfusor Compact | 1152 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000687 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Syringe Pump | Bbraun/Perfusor Compact | 1153 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000687 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Syringe Pump | Bbraun/Perfusor Compact | 1154 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000687 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Syringe Pump | Bbraun/Perfusor Compact | 1156 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000687 | 2025-10-07 | RSUP Prof. Dr. R. D. Kandou Manado | Syringe Pump | Bbraun/Perfusor Compact | 1157 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
07000782 | 2025-10-07 | RSUD Merah Putih | Ventilator | Draeger/Babylog VN800 | ASTK0008 | Survey | Training user | Ready to use |  | Wahyu
05000380 | 2025-10-07 | Smileworks Dental Care | Dental Unit | Kavo/E50 | 30212529 | Service | Penampungan air bocor; Penampungan dibersihkan dan ditutup ulang, cek dan keringkan body yang rembes air | Ready to use |  | Farhat, Ipan
6000934 | 2025-10-07 | RS Raffa Medika | Ventilator | Draeger/Babylog VN600 | ASTE0067 | Survey | Pengecekan unit | Ready to use |  | Dwi
8000732 | 2025-10-07 | RSUD Dr. R. Soedjati Soemodiarjo | Patient MOnitor | Draeger/Vista 120 | VGSNF2515 | Survey | Software booting. Mencoba reload software namun unit masih booting | Pending | Perlu penggantian 9645 mainboard | Amna
07000781 | 2025-10-06 | RSUP dr. Sardjito | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0026 | Service | Penggantian service set 3 tahun | Ready to use |  | Wahyu
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1079 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1065 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1067 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1074 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1068 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1064 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1069 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1073 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1072 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1071 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1075 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1081 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1070 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1070 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1078 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1066 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1080 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1082 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1063 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000501 | 2025-10-06 | RSUP dr. Kariadi | Syringe Pump | BBraun/Perfusor Compact Plus | 1077 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nanda
9000747 | 2025-10-07 | RSUD Dr. H. Chasan Boesoirie | Meja Operasi | SYAS/OP850 | OP850XUZY000009 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000747 | 2025-10-07 | RSUD Dr. H. Chasan Boesoirie | Meja Operasi | SYAS/OP850 | OP850XUZY000006 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000740 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Pendant Surgery | SYAS | SY-PS2501009 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000740 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Pendant Surgery | SYAS | SY-PS2501008 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000740 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Pendant Anasthesia | SYAS | SY-PA2501006 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000740 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Pendant Anasthesia | SYAS | SY-PA2501007 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000742 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Lampu Operasi | SYAS/Lumina | OL9680/80XUZY00012 | Instalasi | Unit sudah berfungsi dengan baik dan sudah training user | Ready to use |  | Ruslan
9000742 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Lampu Operasi | SYAS/Lumina | OL9680/80XUZY00010 | Instalasi | Unit sudah berfungsi dengan baik dan sudah training user | Ready to use |  | Ruslan
9000743 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Infus Pump | Bbraun/Infusomat Space P | 964923 | Service | Penggantian baterai | Ready to use |  | Ruslan
9000743 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Infus Pump | Bbraun/Infusomat Space P | 964924 | Service | Penggantian baterai | Ready to use |  | Ruslan
9000743 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Infus Pump | Bbraun/Infusomat Space P | 964951 | Service | Penggantian baterai | Ready to use |  | Ruslan
9000743 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Infus Pump | Bbraun/Infusomat Space P | 964926 | Service | Penggantian baterai | Ready to use |  | Ruslan
9000743 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Infus Pump | Bbraun/Infusomat Space P | 964925 | Service | Penggantian baterai | Ready to use |  | Ruslan
9000743 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Infus Pump | Bbraun/Infusomat Space P | 814166 | Service | Penggantian baterai | Ready to use |  | Ruslan
9000743 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Infus Pump | Bbraun/Infusomat Space P | 814169 | Service | Penggantian baterai | Ready to use |  | Ruslan
9000743 | 2025-10-04 | RSUD Dr. H. Chasan Boesoirie | Infus Pump | Bbraun/Infusomat Space P | 964927 | Service | Penggantian baterai | Ready to use |  | Ruslan
11000448 | 2025-10-03 | RSUP Makassar | Mobile X-Ray | Apelem/Rafale BDR | 97124084483 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Nasrul
18000686 | 2025-10-03 | RSUP Prof. Dr. R. D. Kandou Manado | Fototerapi | Draeger/Phototherapy | ASJB0043 | Service | Penggantian blue light | Ready to use |  | Irfan R
12000268 | 2025-10-03 | RSUD Hasri Ainun Habibie | MOT | SYAS | - | Service | AC tidak dingin. Pengecekan unit, ada alarm highpress sehingga mcb kompressor trip dan ukuran froon berlebihan; Pengurangan froon pada batas normal | Ready to use | Unit sudah berfungsi normal kembali | Marshall
05000379 | 2025-10-03 | GiGi Dental House | Dental Unit | Kavo/E50 | 30211887 | Service | Suction tidak nyedot, air masuk ke motor suction; Perbaikan fan motor suction dan pasang separator pada motor | Ready to use |  | Farhat
14000376 | 2025-10-03 | RSUP Vertikal Surabaya | Meja Operasi | SYAS/EPIOne OP850 | OP850XVZY0003 | Service | Meja operasi pada bagian kepala turun sendiri dalam waktu 2 hari; Perbaikan selenoid valve, cek level oil | Ready to use |  | Eddy, Fahim
9000745 | 2025-10-03 | RSUD Sanana | Pendant | SYAS/MOT | SY-PA2501006 | Survey | Ada kebocoran pada pendant; Penggantian selang angin | Ready to use |  | Ruslan
11000447 | 2025-10-02 | RSUP Tadjuddin Chalid | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/133/C2 | Service | Daily check fail. Cek tegangan dan ups, restart unit dan daily check ulang. Unit sudah berfungsi normal | ready to use |  | Nasrul
18000656 | 2025-10-02 | RS Budi Setia Langowan | Ventilator | Draeger/Savina 300 | ASTK0072 | Survey | Refresh training | Ready to use |  | Irfan R
12000270 | 2025-10-02 | RSUD Hasri Ainun Habibie | Lampu Operasi | Draeger |  | Service | Arm lampu tidak bisa diputar 360 derajat; Pembersihan karatan pada arm lampu | Ready to use | Unit sudah bisa diputar normal | Marshall
12000252 | 2025-10-02 | RSUD Otohana | Patient MOnitor | Draeger/Vista 120S | S2SNM0821 | Survey | SPO2 tidak terbaca. Sensor SPO2 rusak | Pending | Akan dilakukan claim warranty | Marshall
8000736 | 2025-10-02 | RS Assuyuthiyyah | Mesin Anasthesia | Draeger/Fabius Plus | ASUA0003 | Survey | O2 sensor calibration fail. Cek O2 sensor, calibration o2 sensor | Ready to use |  | Amna
07000779 | 2025-10-02 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/ Babylog VN500 | ASJN0023 | Maintenance | Maintenance Total care | Ready to use |  | Wahyu
07000778 | 2025-10-02 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/ Babylog VN500 | ASKL0092 | Maintenance | Maintenance Total care | Ready to use |  | Wahyu
05000378 | 2025-10-02 | RS Dr. Suyoto | Dental Unit | Kavo/1058 Life | 10207195 | Survey | Penampungan air suction bocor; Pembersihan penambungan air | Ready to use |  | Farhat
05000378 | 2025-10-02 | RS Dr. Suyoto | Dental Unit | Kavo/1058 Life | 10207194 | Survey | Penampungan air suction bocor; Pembersihan penambungan air | Ready to use |  | Farhat
04000523 | 2025-10-02 | RSUD Wonosari | C-arm | Apelem/Evo R+ | 55121060911 | Service | Nilai kolimasi terlalu besar, tidak lolos uji; Adjustment kolimasi | Ready to use |  | Satria, Imam
04000522 | 2025-10-02 | RSUD Wonosari | ESWL | Medispec/E3000 | 0922005 | Service | Nilai kolimasi terlalu besar, tidak lolos uji; Adjustment kolimasi | Ready to use |  | Satria, Imam
17000220 | 2025-10-02 | RSUD Gambiran Kediri | C-arm | Apelem/Evo R+ | 55121001850 | Survey | Refresh training | Ready to use |  | Bahar
14000375 | 2025-10-02 | RSUP Vertikal Surabaya | C-arm | Apelem/Evo R+ | 551240231085 | Service | Unit tidak bisa menyala, cek board distribusi (IDB), cek power supply dan console; Ada problem di console dan sudah diperbaiki | Ready to use |  | Eddy, Fahim
9000730 | 2025-10-02 | RSUD Sanana | Central Monitor | Draeger/CMS Vista 300 | CASUC0009 | Instalasi | Unit belum uji fungsi karna belum bisa konek ke patient monitor | Pending |  | Ruslan, Ipan
9000729 | 2025-10-01 | RSUD Sanana | X-ray Digital | Apelem/Camargue DR | 25102312 | Maintenance | Preventive Maintenance Warranty dan pemasangan DAP | Ready to use |  | Ruslan, Ipan
6000933 | 2025-10-01 | RS JIH Purwokerto | Jaundice Meter | Draeger/JM-105 | H350745 | Service | Kalibrasi Jaudience Meter | Ready to use |  | Dwi
14000374 | 2025-10-01 | RSUP Vertikal Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0146 | Service | Gambar gelap saat mode AEC; Kalibrasi tube, adjust EC dan backup data | Ready to use |  | Eddy, Bahar, Fahim
05000377 | 2025-10-01 | Gigi Dental House | Dental Unit | Kavo/E50 | 30211887 | Survey | Suction tidak menyedot; Pengecekan dental unit dan motor suction | Pending |  | Farhat
19000496 | 2025-10-01 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Service | Pemindahan eswl, setting unit | Ready to use |  | Nanda
18000655 | 2025-10-01 | RSUD ODSK Provinsi Sulawesi Utara | Patient MOnitor | Draeger/Vista 120 | VGSSF0108 | Survey | Unit tidak bisa menyala. Cek power supply dan board power supply; Board LCD rusak | Pending | Pengajuan claim warranty | Irfan R
19000497 | 2025-09-30 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Survey | Pendampingan ukes C-arm | Ready to use |  | Nanda
15000318 | 2025-09-30 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-Ray | Apelem/Rafale | 009/15/01095 | Service | Filament fault; Pembersihan board menggunakan contact cleaner dan pengecekan fuse | Ready to use |  | Sutan
07000777 | 2025-09-30 | RSUD R.A.A Tjokronegoro | Bubble CPAP | Draeger/Oxymixer | QM20255065 | Survey | Humidity E87; Install humidity backup, humidifier dibawa untuk di cek | Ready to use |  | Wahyu
05000376 | 2025-09-30 | Signature Dental Care | Dental Unit | Kavo/1058 Life | 10207194 | Survey | Three way syringe bocor; Kencangkan socket three way syringe | Ready to use |  | Farhat
09000728 | 2025-09-30 | RSUD Sanana | Mesin Anasthesia | Draeger/Atlan A300 | ASSD0233 | Service | Pemasangan monting arm dan penyambungan unit ke patient monitor | Ready to use |  | Ruslan, Ipan
09000727 | 2025-09-30 | RSUD Sanana | Dental Unit | Kavo/E30 | 101111050 | Survey | Handpiece tidak berfungsi; Periksa sambungan selang pada dentist element dan selenoid dan pengecekan sumber air. Micromotor, scaller tidak berfungsi karna board dentist element, board scaller dan handpiece scaler bermasalah | Ready to use | Unit masih bisa berfungsi dengan catatan micromotor dan handpiece masih belum berfungsi | Ruslan, Ipan
8000752 | 2025-09-30 | RSUP Dr. R. Soedjati Purwodadi | Digital Radiografi Fluoroskopi | Apelem/Platinum | - | Survey | Survey ruangan untuk installasi unit | Ready to use |  | Amna
18000695 | 2025-09-29 | RSUD Provinsi Sulawesi Utara | Bubble CPAP | Draeger/Oxymixer | QMHF20208430 | Service | Pemasangan monting arm dan penyambungan unit ke patient monitor | Ready to use |  | Irfan, Marshall
18000695 | 2025-09-29 | RSUD Provinsi Sulawesi Utara | Bubble CPAP | Draeger/Oxymixer | QMHF20208431 | Service | Pemasangan monting arm dan penyambungan unit ke patient monitor | Ready to use |  | Irfan, Marshall
8000735 | 2025-09-29 | RS PKU Muhammadiyah Pekajangan | Mesin Anasthesia | Draeger/Atlan A300 XL | ASMM0060 | Service | Install backup unit | Ready to use |  | Amna
06000923 | 2025-09-29 | RSU An Ni'mah | USG | Mindray/Z60 Expert | HG5-38000272 | Survey | Body pecah retak, trackball tidak respon, baterai drop; Pemeriksaan lanjutan dengan remote teknisi mindray, menunggu info spare part | Pending |  | Dwi
07000776 | 2025-09-29 | RSUD Merah Putih | Ventilator | Draeger/Babylog VN800 | ASTK0008 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Wahyu
011000446 | 2025-09-29 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum | G248L0145 | Service | Monitor/Komputer mati; Cek tegangan, cek CPU | Ready to use | Unit sudah bisa digunakan kembali | Nasrul
09000724 | 2025-09-29 | RSUD Jailolo Halmahera Barat | Dental Unit | Kavo/E30 | 10117281 | Instalasi | Unit berfungsi dengan scaller backup | Ready to use |  | Ruslan, Ipan
19000495 | 2025-09-29 | RST Bhakti Wira Tamtama | Bubble CPAP | Draeger/Oxymixer | QM20555073 | Service | Air compressor masuk ke unit; Pemasangan pressure regulator dan water trap | Ready to use |  | Nanda
19000495 | 2025-09-29 | RST Bhakti Wira Tamtama | Bubble CPAP | Draeger/Oxymixer | QM20255076 | Service | Air compressor masuk ke unit; Pemasangan pressure regulator dan water trap | Ready to use |  | Nanda
14000372 | 2025-09-28 | RSUD Gambiran Kediri | C-arm | Apelem/Evo R+ | 55121001850 | Survey | Pergerakan naik turun c-arm tidak bisa, gambar tidak tersimpan, tombol putar error; Cek motor kolimator dan tombol emergency baik | Ready to use |  | Fahim
05000375 | 2025-09-26 | Signature Dental Care | Dental Unit | Kavo/1058 Life
011000445 | 2025-09-26 | RS Regional dr. Hasri Ainun Habibie | Ventilator | Draeger/Babylog VN800 | ASSM0042 | Service | FIO2 error; Cek supply O2, supply air, system test, tank error log, hit software. Konsentrasi O2 supply rendah (hanya 86%), menyebabkan setting FIO2 error pada saat settingan tinggi (diatas 40%) | Pending | Settingan FIO2 rendah tidak dapat digunakan (Diatas 40%) | Nasrul
3000905 | 2025-09-26 | RSUD Kota Kotamobagu | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/043/C0 | Survey | Power supply detector error, power supply pc console error, software console error | Pending |  | Imam, Marshall
3000904 | 2025-09-26 | RSUD Kota Kotamobagu | C-arm | Apelem/Evo R+ | 55118031654 | Survey | Power supply error motor, software console error; Board PSU 24V DC dan board B5 GSC rusak | Pending |  | Imam, Marshall
8000733 | 2025-09-26 | RS PKU Muhammadiyah Pekajangan | Mesin Anasthesia | Draeger/Atlan A300 XL | ASMM0060 | Service | Packing backup unit | Ready to use |  | Amna
06000930 | 2025-09-26 | RSUD Dr. Soeselo Slawi | UPS | Oaxaka/AP3380 | 110045791F80311900001 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
05000375 | 2025-09-26 | Signature Dental Care | Dental Unit | Kavo/1058 Life | 10207191 | Survey | Peminjaman board scaller, selang dan scaller tip scaller | Ready to use | Untuk backup E30 RSUD Jailolo | Farhat
07000775 | 2025-09-26 | RSUD Kota Yogyakarta | Mesin Anasthesia | Draeger/Atlan A350 | ASUF0273 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Wahyu
15000317 | 2025-09-26 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156P0142 | Service | Table top tidak bisa bergerak dan push button console patah; Kalibrasi table top | Ready to use |  | Sutan
19000494 | 2025-09-25 | RSUD Pandan Arang Boyolali | Infant Warmer | Draeger/Resuscitaire | CF10509 | Service | Error 05. Heater putus | Pending | Disarankan penggantian baterai | Nanda
19000507 | 2025-09-26 | RSUP Surakarta | USG | Mindray/Nuewa 19 Exp | KV3-52000073 | Survey | Installasi unit demo | Ready to use |  | Nanda
19000493 | 2025-09-25 | RSUD Pandan Arang Boyolali | Incubator Transport | Draeger/TI500 | BH12345 | Survey | Baterai drop, baterai mengembung | Pending | Disarankan penggantian baterai | Nanda
12000255 | 2025-09-25 | RSUD Ainun Habibi Gorontalo | Bor | Aesculap/ELAN4 | 003075 | Survey | Training user | Ready to use |  | Marshall, Irfan
14000371 | 2025-09-25 | RSUD Dr. Moh. Saleh | Bed Patient | SYAS | 240679101260400127 | Survey | Bed tidak bisa bergerak; Pengecekan remote dan tombol bed. Unit sudah berfungsi kembali | Ready to use |  | Fahim
8000734 | 2025-09-25 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/Noble 4 | 250879101260400103 | Survey | Sideguard tidak bisa terkunci; Pengecekan pegas, dan pelatuk side guard | Pending |  | Amna
5000374 | 2025-09-24 | Smileworks Dental Care | Dental Unit | Kavo/E50 | 30212529 | Service | Penampungan air suction bocor; Cek O'ring pada penampungan dan penambahan sealtip | Ready to use |  | Farhat
5000374 | 2025-09-24 | Smileworks Dental Care | Dental Unit | Kavo/E50 | 30212530 | service | Scaller bermasalah; Pemasangan selang scaller backup | Ready to use |  | Farhat
011000444 | 2025-09-24 | RSUP Dr. Wahidin Sudirohusodo | Infant Warmer | Draeger/Babyleo TN500 | ASTH0140 | Service | Parameter skin peripheral skin central tidak terbaca; Penggantian skin sensor | Ready to use |  | Nasrul
18000654 | 2025-09-24 | RSUD Zainal Umar Sidiki | Ventilator | Draeger/Savina ID | ASUH3017 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
18000654 | 2025-09-24 | RSUD Zainal Umar Sidiki | Ventilator | Draeger/Savina ID | ASUH3015 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
19000492 | 2025-09-24 | RS St Elizabeth Semarang | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Survey | General check. Unit OK | Ready to use |  | Nanda
19000491 | 2025-09-24 | RS St Elizabeth Semarang | Mesin Anasthesia | Draeger/Fabius Tiro | ASHA0114 | Survey | Bootloop dan unit tidak menyala. Unit check; Baterai rusak | Pending | Disarankan penggantian baterai dan housing O2 | Nanda
20000412 | 2025-09-24 | RSUP Makassar | MOT | SYAS | - | Service | Scrubcink (SYLA2501009) bocor; Penampah sealtip pada sambungan | Ready to use |  | Sujana
8000731 | 2025-09-23 | RSUD Dr. Loekmono Hadi Kudus | Mesin Anasthesia | Draeger/Perseus A500 | ASUS0046 | Maintenance | Device check, general check, cleaning valve | Ready to use |  | Amna
14000373 | 2025-09-23 | RSUP Vertikal Surabaya | C-arm | Apelem/Evo R+ | 551240231085 | Survey | Lampu standby tidak menyala dan alat tidak dapat digunakan. Pengecekan fungsi tombol dan output powersupply. Board B5 tidak ada output | Pending | Menunggu pengecekan pada board | Fahim
04000520 | 2025-09-23 | RSUD Subang | Bor Ortho | Bbraun/Acculan4 Orthopedi Optima | 2511700 | Instalasi | Install unit | Ready to use | Alat berfungsi dengan baik dan traning sudah dilakukan | Satria, Bahar
05000372 | 2025-09-23 | Smileworks Dental Care | Dental Unit | Kavo/E50 | 30212530 | Survey | Pengecekan scaller coba tukar selang dan scallernya pada unit lain | Pending | Selang/konektor scaller bermasalah dan selang scaller di bawa ke kantor (warranty) | Farhat
06000929 | 2025-09-23 | RSUD Dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Atlan A350 | ASRK0304 | Survey | Leakage bag, preventice maintenance; Cek self test. Penggantian sample line, water lock | Ready to use | Unit masih bisa digunakan namun direkomendasikan untuk servoce set 2tahun | Dwi
06000928 | 2025-09-23 | RSUD Dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Fabius Tiro | ASCH0077 | Survey | O2 fail; O2 calibration, cek flow meter. | Pending | Rekomendasi penggantian O2 sensor capsule, flow meter dan fabius service set 3 tahun | Dwi
6000927 | 2025-09-23 | RSUD Dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Primus | ASKJ0050 | Survey | Analisa gas gagal, O2 sensor fail, valve ventilator fail; Cek SC monitorng gas dengan PCM, cek Mv2. | Pending | Rekomendasi penggantian ILCA 3 dual, flow power, 2/2 port distributing valve dan primus service set 3 tahun | Dwi
19000506 | 2025-09-22 | RS JIH Solo | Ventilator | Draeger/Babylog VN600 | ASSK0027 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Nanda
11000443 | 2025-09-22 | RSUD Kota Makassar | Bone Mineral Densitometer | Medilink/Medix DR | 460624 | Service | Re-install, kalibrasi filament gagal | Pending | Disarankan penggantian inverter dan filament board | Nasrul
05000373 | 2025-09-22 | Smileworks Dental Care | Dental Unit | Kavo/E50 | 30212530 | Service | Setting air bowl, adjust lampu posisi, setting air handpiece | Ready to use |  | Farhat
05000373 | 2025-09-22 | Smileworks Dental Care | Dental Unit | Kavo/E50 | 30212529 | Service | Setting air bowl, adjust lampu posisi, setting air handpiece | Ready to use |  | Farhat
6000926 | 2025-09-22 | RSUD Cilacap | USG | Mindray/M9 | CCI-51006326 | Survey | Claim warranty baterai |  |  | Dwi
8000727 | 2025-09-22 | RSUD RA Kartini Jepara | Ventilator | Draeger/Savina 300 | ASEK0197 | Service | Alat mati total; Penggantian baterai | Ready to use |  | Amna
05000371 | 2025-09-19 | Smileworks Dental Care | Motor Suction | Durr | 5506976019 | Instalasi | Install unit | Ready to use | Hanya pemasangan motor suction (pembelian di luar gratia) | Farhat, Ipan
05000370 | 2025-09-19 | Smileworks Dental Care | Dental Unit | Kavo/Estetica E50 Life | 30212530 | Instalasi | Install unit | Ready to use | Unit telah terpasang, air, angin, listrik sudah ok, suction not ready (menunggu motor suction) | Farhat, Ipan
05000369 | 2025-09-19 | Smileworks Dental Care | Dental Unit | Kavo/Estetica E50 Life | 30212529 | Instalasi | Install unit | Ready to use | Unit telah terpasang, air, angin, listrik sudah ok, suction not ready (menunggu motor suction) | Farhat, Ipan
03000903 | 2025-09-19 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Service | Setting IP WSI | Ready to use | WSI I sudah konek ke console, sudah bisa kirim image | Imam
07000773 | 2025-09-19 | RSUD R.A.A Tjokronegoro | Bubble CPAP | Draeger/Oxymixer | QM20255065 | Service | Pemasangan pressure regulator & refresh training | Ready to use |  | Wahyu
09000738 | 2025-09-19 | RSUD Dr. H. Chasan Boesoirie | Lampu Operasi | SYAS/Lumima | OL9680/8-XUZY00010 | Survey | Arm lampu kurang tinggi; Setting arm | Ready to use |  | Ruslan
09000738 | 2025-09-19 | RSUD Dr. H. Chasan Boesoirie | Lampu Operasi | SYAS/Lumima | OL9680/80XUZY00012 | Survey | Arm lampu kurang tinggi; Setting arm | Ready to use |  | Ruslan
19000488 | 2025-09-19 | RSUD Bung Karno Surakarta | USG | Mindray/Resona 7 | 7V-OB002253 | Maintenance | general check. Unit OK | Ready to use |  | Nanda
19000487 | 2025-09-18 | BKIM Provinsi Jawa Tengah Semarang | Patient Monitor | Draeger/Vista 120 SC | SBSNF0650 | Survey | Alat mati ketika listrik dicabut. Baterai ada yang hilang; Sementara gunakan baterai datri unit lain | Ready to use |  | Nanda
12000249 | 2025-09-18 | RSUD Totokabila | Ventilator | Draeger/Evita 300 | ASND0131 | Survey | Tombol rotari knob error; Cek dengan rotari knob lain masih oke. Kerusakaan di main board base | Pending |  | Marshall
17000219 | 2025-09-18 | RSUD Kab. Temanggung | C-Arm | Apelem/Evo R+ | 551-15-011-469 | Service | Mengganti plat kolimasi dengan buatan baru, adjust Iris untuk mengatur dosis dan uji kesesuain | Ready to use | Dosis kolimasi dan UKES memenuhi persyaratan/lolos, harus ganti kolimator karena kolimator tidak bisa diatur di posisi tertentu | Bahar, Imam
011000442 | 2025-09-18 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | SYAS/Premium SB 3 motor | 240762445004400476 | Survey | Remote tidak berfungsi; Cek kontrol box dan cek motor. Remote control rusak | Pending | Pengajuan claim warranty | Nasrul
18000696 | 2025-09-18 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASED0171 | Survey | Device check fail, sensor mixer fail, gas analyzer fail; Cek mixer dan sensor high dan low | Pending | Direkomendasikan penggantian ILCA2, sensor P.DIFF | Irfan
20000401 | 2025-09-17 | RSUP Tadjudin Chalid | ESWL | Medispec/EM1000 | 0922014 | Service | Terdapat angin di balon EMH; Pengosongan angin di EMH dan penambahan air | Ready to use |  | Sujana
12000254 | 2025-09-17 | RSUD Totokabila | Lampu Operasi | Draeger/Polaris | ASML0272 | Service | Arm lampu operasi hampir terlepas; Setting ulang arm dan pasang lebih kencang | Ready to use |  | Marshall
03000902 | 2025-09-17 | RSUD Tidar Magelang | Bone Mineral Densitometer | Medilink/Medix DR | M15016M258 | Survey | Ganti PSU (disediakan pihak RS) | Ready to use |  | Imam, Bahar
09000735 | 2025-09-17 | RSUD Dr. H. Chasan Boesoirie | Central Monitor | Draeger/CMS Vista 300 | CASUE0002 | Instalasi | Unit berfungsi baik | Ready to use |  | Ruslan
18000629 | 2025-09-17 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Perseus A500 | ASJM0178 | Service | Gas O2 fail, mixer fail dan baterai drop; Penggantian part PBA Optimix pressure, differential dan baterai | Ready to use |  | Irfan
12000247 | 2025-09-16 | RSUD Ainun Habibi Gorontalo | MOT | SYAS | - | Survey | Cover monolitik lepas; Pemasangan kembali cover dan dilem ulang | Ready to use |  | Marshall
12000246 | 2025-09-16 | RSUD Totokabila | MOT | SYAS | - | Service | Ac tidak dingin; Ada kerusakan pada kontraktor HVAC, dan sudah diperbaiki | Ready to use |  | Marshall
14000370 | 2025-09-16 | RSUP Vertikal Surabaya | MOT | SYAS | - | Survey | Pintu tidak dapat menutup | Pending | Penjadwalan dengan teknisi KSS | Fahim
19000482 | 2025-09-16 | RSUD Ibu Fatmawati Surakarta | Infant Warmer | Draeger/Resuscitaire | DSNL-0048 | Survey | Cek alat, cek unit, kontrol mini valve 5 way rusak | Pending | Disarankan ganti control mini 5 way | Nanda
19000450 | 2025-09-16 | RSUD Ibu Fatmawati Surakarta | Infant Warmer | Draeger/Resuscitaire | DSNL-0046 | Survey | Cek unit, cek probe, ada kerusakan pada sensor probe suhu, kit corner blok patah | Ready to use | Unit sementara dapat digunakan, disarankan ganti sensor probe dan kit corner blok | Nanda
17000218 | 2025-09-16 | RSUD Kab. Temanggung | C-Arm | Apelem/Evo R+ | 551-15-011-469 | Service | Cek iris di II, Cek sesuai instruksi apelem | Ready to use | Dosis pada pengujian input II (POV 23,15,11) sudah ok, pengujian dosis di mode typical dan maksimum belum ok karena iris di kolimator bermasalah | Bahar, Imam
07000772 | 2025-09-16 | RSUD Bagas Waras | USG | Mindray/DC8 | DE-76005148 | Survey | Pengecekan populasi | Ready to use | Unit berfungsi baik | Wahyu
6000931 | 2025-09-16 | RSUD Ajibarang | USG | Mindray/DC30 | 9D-18009329 | Survey | Pendampingan hands on mindray di poli obgyn | Ready to use |  | Dwi
09000736 | 2025-09-16 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Evita V600 | ASTE0093 | Survey | Alarm maintenance battery dan humidifier error; Penggantian kabel heater wire yang baru & maintenance battery | Ready to use |  | Ruslan
15000316 | 2025-09-15 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156P0142 | Service | Hapus data base pasien karna software lambat | Ready to use |  | Sutan
15000315 | 2025-09-15 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-Ray | Apelem/Rafale | 009/15/01095 | Service | Filament fault; Pembersihan board menggunakan contact cleaner dan pengecekan fuse | Ready to use |  | Sutan
11000551 | 2025-09-15 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0396 | Survey | Expiration flow sensor fail | Pending | Direkomendasikan penggantian flow sensor dan service set 2 tahun | Nasrul
6000924 | 2025-09-15 | RSUD PREMBUN KAB. KEBUMEN | Bubble CPAP | Draeger/Oxymixer | [5 empty] | Dwi
9000721 | 2025-09-15 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 ID Prime | ASUH3003 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000721 | 2025-09-15 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 ID Prime | ASUH3004 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
9000721 | 2025-09-15 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 ID Prime | ASUH3005 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/SBW 4 Bed Noble | 250879101260400108 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/SBW 4 Bed Noble | 250879101260400105 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/SBW 4 Bed Noble | 250879101260400103 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/SBW 4 Bed Noble | 250879101260400110 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/SBW 4 Bed Noble | 250879101260400104 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/SBW 4 Bed Noble | 250879101260400107 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/SBW 4 Bed Noble | 250879101260400102 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/SBW 4 Bed Noble | 250879101260400109 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/SBW 4 Bed Noble | 250879101260400106 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Bed Patient | SYAS/SBW 4 Bed Noble | 250879101260400101 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Strecher | SYAS/SST-2 | 250968303000400182 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Strecher | SYAS/SST-3 | 250968303000400179 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Strecher | SYAS/SST-4 | 250968303000400181 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Strecher | SYAS/SST-5 | 250968303000400175 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Strecher | SYAS/SST-6 | 250968303000400177 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Strecher | SYAS/SST-7 | 250968303000400180 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Strecher | SYAS/SST-8 | 250968303000400176 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
8000725 | 2025-09-15 | RSUD Dr. R. Soedjati Soemodiarjo | Strecher | SYAS/SST-9 | 250968303000400178 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
18000697 | 374633 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0180 | Service | Ada kebocoran O2 di dalam unit; Setelah di cek kerusakan pada pressure regulator O2. Penggantian pressure regulator O2 milik RS | Ready to use |  | Irfan
20000398 | 2025-09-15 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASEF0068 | Service | Gas supply O2 failure, pengecekan gas supply dan tarik error; Perlu penggantian pressure regulator dan mainteannce kit | Pending |  | Sujana
07000771 | 2025-09-13 | RSUP dr. Sardjito | Mesin Anasthesia | Dareger/Perseus A500 | ASKF0076 | Maintenance | Penggantian maintenance kit breathy assy 2y | Ready to use |  | Wahyu
6000924 | 2025-09-13 | RSUD PREMBUN KAB. KEBUMEN | Bubble CPAP | Draeger/Oxymixer | QM20230154 | Survey | Gagal lolos kalibrasi; Check dengan Flow analyser; Nilai FIO2 dan flow tidak laik | Pending |  | Dwi
6000924 | 2025-09-13 | RSUD PREMBUN KAB. KEBUMEN | Bubble CPAP | Draeger/Oxymixer | QM20230153 | Survey | Gagal lolos kalibrasi; Check dengan Flow analyser; Nilai FIO2 dan flow tidak laik | Pending |  | Dwi
07000771 | 2025-09-13 | RSUP dr. Sardjito | Mesin Anasthesia | Dareger/Perseus A500 | ASKF0077 | Maintenance | Penggantian maintenance kit breathy assy 2y | Ready to use |  | Wahyu
11000440 | 2025-09-12 | RSUP Tadjudin Chalid | Ventilator | Draeger/Babylog VN800 | ASSJ0037 | Service | Flow sensor putus dan VT tidak terbaca; Penggantian flow sensor | Ready to use |  | Nasrul
8000724 | 2025-09-12 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0093 | Survey | System check neo measurement failed; System check, check insert neo flow, cek neo flow cable. Neo flow cable rusak | Pending | Perlu penggantian neo flow cable | Amna
12000245 | 2025-09-12 | RSUD Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Survey | Gambar tidak terkirim ke workstation, switch hub rusak; Ganti switch hub milik RS | Ready to use |  | Marshall
20000397 | 2025-09-12 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 Plus | DSTH0095 | Service | Heater failed, pengecekan heater; Heater rusak dan perlu diganti | Pending |  | Sujana
8000730 | 2025-09-11 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 | ASTK0258 | Survey | Alarm pressure low connect; Cek dengan service mode, cek blower, cek expiratory valve | Ready to use |  | Amna
14000367 | 2025-09-11 | RSUD Gambiran Kediri | C-Arm | Apelem/Evo R+ | 551-21-001-850 | Service | Penggantian power supply 24V, Penggantian SSD ITB dan RAM TDR4 16 GB (Include software) | Ready to use |  | Fahim, Imam, Bahar
04000519 | 2025-09-11 | RSUP dr. Ben Mboi | Digital Radiography Fluoroscopy | Apelem/Platinum | F226F0736 | Maintenance | Preventive Maintenance dan Refresh Training | Ready to use | Alat berfungsi dengan baik, kualitas gambar bagus, semua pergerakan aman, refresh training telah dilakukan pada tanggal 10 September 2025 | Satria
6000922 | 2025-09-11 | RSUD Cilacap | USG | Mindray/M9 | CC1-51006326 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
14000369 | 2025-09-10 | RSUP Vertikal Surabaya | Pass Box | Lokal |  | Survey | Cek Kuncian Pass box | Pending | Pass box tidak dapat tertutup dengan baik, perbaikan dijadwalkan oleh tim KSS | Fahim
03000901 | 2025-09-10 | RSUD Kab. Temanggung | C-Arm | Apelem/Evo R+ | 551-15-011-469 | Service | Reload software Dis 1000 (upgrade SSD ITB), adjust kolimasi (ok), adjust dose (belum sempurna) menunggu instruksi dari apelem | Ready to use |  | Imam, Bahar
19000449 | 2025-09-10 | RS Indriati Boyolali | Ventilator | Draeger/Savina 300 | ASNL-0485 | Survey | Cek unit, device check ulang, good, running, ok | Ready to use | Unit usia sudah 5 tahun; disarankan penggantian service set 6 tahun include baterai | Nanda
07000767 | 2025-09-10 | RSUD R.A.A Tjokronegoro | Bubble CPAP | Draeger/Oxymixer | QM20255065 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
15000314 | 2025-09-10 | RSUP Dr. Mohammad Hoesin (KSO) | Printer | Fuji/Dry Pix | 66431571 | Service | Printer tidak bisa print; Setting ulang komputer menyesuaikan jumlah film | Ready to use |  | Sutan
8000729 | 2025-09-10 | RSI PKU Muhammadiyah Pekajangan | Mesin Anasthesia | Draeger/Atlan A300 XL | ASMM0060 | Service | Installasi backup unit | Ready to use |  | Amna
18000698 | 2025-09-10 | RSUD Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | VGSSF0108 | Service | Baterai tidak mengisi; Baterai tidak terpasang dengan baik, dan sudah dipasang kembali dengan baik | Ready to use |  | Irfan
8000723 | 2025-09-09 | RS Mitra Bangsa Pati | Mesin Anasthesia | Draeger/Atlan A300 XL | ASMM0060 | Service | Packing demo unit | Ready to use |  | Amna
15000313 | 2025-09-09 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Table top tidak bisa bergerak; Pengencangan rantai table top | Ready to use |  | Sutan
19000448 | 2025-09-09 | RST Bhakti Wira Tamtama | CPAP | Draeger/Seattle PAP | QM20255073 | Survey | Cek alat, cek unit, kompresor berair | Ready to use | Disarankan pakai water trap dan pressure regulator | Nanda
05000368 | 2025-09-09 | RSUD Dr. Drajat Prawiranegara | Dental Unit | Kavo/Primus 1058 | 2018-11-10205461 | Survey | Pengecekan pada sensor plat base ada kemungkinan sensor kursi rusak, air pada kopling highspeed tidak keluar | Pending | Unit belum bisa digunakan dan suction belum bisa dapat digunakan | Farhat, Ipan
04000518 | 2025-09-09 | RSUP dr. Ben Mboi | Laser Surgical | Jena/Multipulse Hoplus | J191150295 | Maintenance | Preventive Maintenance | Ready to use | Alat berfungsi dengan baik, energi tembakan masih tinggi | Satria
04000517 | 2025-09-09 | RSUP dr. Ben Mboi | Mobile X-Ray | Apelem/Rafale BDR 40 | 971-23-004-385 | Maintenance | Preventive Maintenance | Ready to use | Alat berfungsi dengan baik, semua pergerakan aman bisa melakukan ekspose | Satria
07000766 | 2025-09-09 | RSUP dr. Sardjito | Patient Monitor | Draeger/Vista 120 | VTSKF0032 | Survey | MIBP bising; Cek module dan pompa NIBP. Module NIBP rusak | Ready to use | Unit bisa digunakan namun terbatas. Direkomendasikan penggantian module NIBP | Wahyu
07000765 | 2025-09-09 | RSUP dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASTD0075 | Survey | LCD muncul lightsaber; Cek LCD. Penggantian LCD | Ready to use |  | Wahyu
9000737 | 2025-09-09 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 ID | ASUA3009 | Survey | Tidak ada mode O2 therapy; Cek sn dan versi software untuk claim penambahan mode | Pending |  | Ruslan
9000737 | 2025-09-09 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 ID | ASUA3008 | Survey | Tidak ada mode O2 therapy; Cek sn dan versi software untuk claim penambahan mode | Pending |  | Ruslan
9000737 | 2025-09-09 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 ID | ASUA3007 | Survey | Tidak ada mode O2 therapy; Cek sn dan versi software untuk claim penambahan mode | Pending |  | Ruslan
9000737 | 2025-09-09 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 ID | ASUA3006 | Survey | Tidak ada mode O2 therapy; Cek sn dan versi software untuk claim penambahan mode | Pending |  | Ruslan
09000734 | 2025-09-08 | RSUD Dr. H. Chasan Boesoirie | Pendant | SYAS | SY-PA2501007 | Instalasi | Unit sudah diuji fungsi dengan kelistrikan namun gas medik belum tersambung | Pending |  | Ruslan
09000734 | 2025-09-08 | RSUD Dr. H. Chasan Boesoirie | Pendant | SYAS | SY-PS2501008 | Instalasi | Unit sudah diuji fungsi dengan kelistrikan namun gas medik belum tersambung | Pending |  | Ruslan
09000734 | 2025-09-08 | RSUD Dr. H. Chasan Boesoirie | Pendant | SYAS | SY-PA2501006 | Instalasi | Unit sudah diuji fungsi dengan kelistrikan namun gas medik belum tersambung | Pending |  | Ruslan
09000734 | 2025-09-08 | RSUD Dr. H. Chasan Boesoirie | Pendant | SYAS | SY-PA2501006 | Instalasi | Unit sudah diuji fungsi dengan kelistrikan namun gas medik belum tersambung | Pending |  | Ruslan
04000516 | 2025-09-08 | RSUP dr. Ben Mboi | ESWL | Medispec/EM1000 | 0922170 | Maintenance | Preventive Maintenance | Ready to use | Alat berfungsi dengan baik, test nembak aman | Satria
07000764 | 2025-09-08 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASJH0121 | Service | Layar blank; Penggantian LCD | Ready to use |  | Wahyu
6000921 | 2025-09-08 | RSUD Prof. Dr. Margono Soekarjo | Jaundice Meter | Draeger/JM-105 | H3503802 | Survey | Internal kalibrasi. Unit ready digunakan | Ready to use |  | Dwi
11000435 | 2025-09-08 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Kinerja motor actuator berat; Cek water chamber. Penggantian water chamber dan konketor kuningan emh | Ready to use |  | Nasrul
08000721 | 2025-09-08 | RSUD RA Kartini Jepara | USG | Mindray/DP20 | LK7B36001200 | Survey | cek software, cek probe, pembersihan trackball | Ready to use |  | Amna
08000721 | 2025-09-08 | RSUD RA Kartini Jepara | USG | Mindray/DP20 | LK7B36001370 | Survey | cek software, cek probe, pembersihan trackball | Ready to use |  | Amna
08000721 | 2025-09-08 | RSUD RA Kartini Jepara | USG | Mindray/DP20 | LK7B36001070 | Survey | cek software, cek probe, pembersihan trackball | Ready to use |  | Amna
08000721 | 2025-09-08 | RSUD RA Kartini Jepara | USG | Mindray/DP20T | GR87005634 | Survey | cek software, cek probe, pembersihan trackball | Ready to use |  | Amna
08000721 | 2025-09-08 | RSUD RA Kartini Jepara | USG | Mindray/DP20T | GR870056 | Survey | cek software, cek probe, pembersihan trackball | Ready to use |  | Amna
18000628 | 2025-09-08 | RS Advent Manado | Bor | Aesculap/ELAN4 | 00806 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Irfan, Marshall
20000411 | 2025-09-08 | RSUP Makassar | C-arm | Apelem/Evo R+ | 551240241086 | Survey | Refresh training | Ready to use |  | Sujana
011000438 | 2025-09-04 | RSUP Dr. Wahidin Sudirohusodo | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/087/C1 | Service | Monitor console mati; Cek tegangan supply dan koneksi kabel | Ready to use | Unit sudah berfungsi dengan baik | Nasrul
14000366 | 2025-09-04 | RSUD Dr. Moh. Saleh | Dressing Trolley | Syas/SCDI | 250869201000400073 | Instalasi | Install unit | Ready to use | Alat berfungsi dengan baik | Fahim
14000366 | 2025-09-04 | RSUD Dr. Moh. Saleh | Dressing Trolley | Syas/SCDI | 250869201000400072 | Instalasi | Install unit | Ready to use | Alat berfungsi dengan baik | Fahim
14000365 | 2025-09-04 | RSUD Dr. Moh. Saleh | Foot Step | Syas/Sesi 1 | 250861301000400502 | Instalasi | Install unit | Ready to use | Alat berfungsi dengan baik | Fahim
14000365 | 2025-09-04 | RSUD Dr. Moh. Saleh | Foot Step | Syas/Sesi 1 | 250861301000400501 | Instalasi | Install unit | Ready to use | Alat berfungsi dengan baik | Fahim
19000447 | 2025-09-04 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Survey | Cek unit, coba reconnect dan by pass kabel startech, running, ok | Ready to use | Unit sementara di running dulu | Nanda
03000399 | 2025-09-04 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Fluoroscopy | Apelem/Platinum | L156F0142 | Service | Penggantian Motor lateral | Ready to use |  | Imam, Sutan
07000763 | 2025-09-04 | RSUP dr. Soeradji Tirtonegoro | Infant Warmer | Draeger/Babytherm 8004 | ASDL0042 | Survey | E35; Cek error log, test D18 & D19, heater mati. Penggantian heating element | Ready to use |  | Wahyu
8000722 | 2025-09-04 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLN0009 | Service | Touchscreen stuck; Calibrated touchscreen | Ready to use |  | Amna
12000253 | 2025-09-04 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 10131182 | Survey | General check; Penggantian baterai CMOS komputer | Ready to use |  | Marshall
20000402 | 2025-09-04 | RS Universitas Hasanuddin | Pendant | Draeger/Movita | ASCM0308 | Service | Penggantian spring pin & Scryp screw | Ready to use |  | Sujana
20000402 | 2025-09-04 | RS Universitas Hasanuddin | Pendant | Draeger/Movita | ASCM0309 | Service | Penggantian spring pin & Scryp screw | Ready to use |  | Sujana
20000402 | 2025-09-04 | RS Universitas Hasanuddin | Pendant | Draeger/Movita | ASCM0310 | Service | Penggantian spring pin & Scryp screw | Ready to use |  | Sujana
20000402 | 2025-09-04 | RS Universitas Hasanuddin | Pendant | Draeger/Movita | ASCM0311 | Service | Penggantian spring pin & Scryp screw | Ready to use |  | Sujana
20000396 | 2025-09-03 | RSUD Andi Makassau | ESWL | Medispec/EM1000 | 0622173 | Service | Monitor tidak connect ke console, pengecekan monitor; Penggantian kabel USB | Ready to use |  | Sujana
19000446 | 2025-09-03 | RSUD dr Adhyatma MPH Semarang | Mesin Anasthesia | Draeger/Primus | ASHH-0122 | Survey | Cek unit, calibrating APL valve, good, running, ok | Ready to use | Unit aman tetapi user minta kepastian untuk di tarik error nya | Nanda
6000932 | 2025-09-03 | RSUD Dr. Soeselo Slawi | MOT | SYAS | - | Service | Surface mounted led panel ceilling mati, swing door mati, panel display mati; Penggantian lampu led panel, adjustment swing door, repair ganti fuse panel | Ready to use |  | Dwi
07000770 | 2025-09-03 | RS Pratama Yogyakarta | Jaundice Meter | Draeger/JM-105 | H3510353 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Wahyu
2000410 | 2025-09-02 | RSUP Makassar | Laparascopy | Aesculap/MultiFlow Plus | 2403CE0798/P6145 | Service | Selang Co2 patah, pengecekan connector selang; Perlu penggantian selang Co2 | Pending |  | Sujana
12000242 | 2025-09-02 | RSUD Provinsi Sulawesi Utara | Bubble CPAP | Draeger/Oxymixer | QMHF20208433 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Marshall, Irfan
12000242 | 2025-09-02 | RSUD Provinsi Sulawesi Utara | Bubble CPAP | Draeger/Oxymixer | QMHF20208431 | iNstalasi | Unit berfungsi dengan baik | Ready to use |  | Marshall, Irfan
09000732 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Pendant | SYAS/Double Arm | SY-PI2406004 | Instalasi | Unit sudah terpasang namun belum uji fungsi karna listrik dan gas medik belum tersambung | Pending |  | Ruslan
09000732 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Pendant | SYAS/Double Arm | SY-PI2506003 | Instalasi | Unit sudah terpasang namun belum uji fungsi karna listrik dan gas medik belum tersambung | Pending |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0018 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0027 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0095 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0034 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0025 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0031 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0035 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0017 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0032 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0029 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0026 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000712 | 2025-09-02 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF0097 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
19000445 | 2025-09-02 | RSUD Dr. Soediran Mangun Sumarso | Infant Warmer | Draeger/Resuscitaire | AC13490 | Service | Cek alat, cek valve flow patient 2, unit ok | Ready to use |  | Nanda
19000444 | 2025-09-02 | RSUD Dr. Soediran Mangun Sumarso | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/134/C2 | Maintenance | Maintenance unit mammography, reset password, daily check expose, running, ok | Ready to use |  | Nanda
14000368 | 2025-09-02 | RSUP Vertikal Surabaya | Meja Operasi | Syas/Epione OP850 |  | Service | Cek valve aktif, cek hidrolik aktif, cek kebocoran oli | Pending | Meja operasi bagian kepala turun sendiri ketika tidak digunakan | Fahim
16000295 | 2025-09-02 | Signature Dental Care | Dental Unit | Planmeca Compact 1 | UPC574532 | Service | Pasang separator dan rinsing durr agar motor suction central tidak banjir | Ready to use | Unit bisa digunakan | Ipan, Farhat, Satria
07000762 | 2025-09-02 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina 300 | ASPJ0062 | Survey | Electronic function test, prenumatic test, unternal battery test | Ready to use | Unit masih bisa digunakan namun direkomendasikan penggantian service set 2 tahun | Wahyu
07000762 | 2025-09-02 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina 300 | ASPJ0063 | Survey | Electronic function test, prenumatic test, unternal battery test | Ready to use | Unit masih bisa digunakan namun direkomendasikan penggantian service set 2 tahun | Wahyu
07000762 | 2025-09-02 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina 300 | ASPJ0064 | Survey | Electronic function test, prenumatic test, unternal battery test | Ready to use | Unit masih bisa digunakan namun direkomendasikan penggantian service set 2 tahun | Wahyu
07000762 | 2025-09-02 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina 300 | ASPJ0065 | Survey | Electronic function test, prenumatic test, unternal battery test | Ready to use | Unit masih bisa digunakan namun direkomendasikan penggantian service set 2 tahun | Wahyu
11000437 | 2025-09-02 | RSUP Makassar | Ventilator | Draeger/Babylog VN800 | ASTM0103 | Service | VT dan RR tidak terbaca; Penggantian flow sensor | Ready to use |  | Nasrul
8000720 | 2025-09-02 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHA0113 | Service | Alarm battery fail, baterai tidak mengisi; M73 plus rusak | Pending |  | Amna
07000760 | 2025-09-01 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Survey | Refresh training dan penggantian water filter | Ready to use |  | Wahyu
11000436 | 2025-09-01 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 07220017 | Service | Penggantian EMH | Ready to use |  | Nasrul, Sujana
05000367 | 2025-08-31 | RSUP Dr. M. Djamil Padang | Lemari dan 2 Unit Passbox | Syas | - | Instalasi | Install unit | Pending | Masih belum dicoba karena belum ada listrik | Farhat, Sutan
05000366 | 2025-08-31 | RSUP Dr. M. Djamil Padang | Surgery | SYAS/1422 S | SY-PS2501003 | Instalasi | Install unit | Pending | Unit sudah terinstall namun adapart yang bermasalah perlu diganti | Farhat, Sutan
05000366 | 2025-08-31 | RSUP Dr. M. Djamil Padang | Syas Pegasus Pendant Anasthesia | SYAS/1422 A | SY-PA2501002 | Instalasi | Install unit | Pending | Unit sudah terinstall namun adapart yang bermasalah perlu diganti | Farhat, Sutan
15000312 | 2025-08-30 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Fluoroscopy | Apelem/Platinum | L156F0142 | Service | Bersihkan sensor detektor, cek motor detektor OK | Ready to use | Tube tidak bergerak namun sudah di perbaiki, alat dapat digunakan kembali | Sutan
12000241 | 2025-08-29 | RSUD Totokabila | Ventilator | Draeger/Evita 300 | ASND-0131 | Service | Penggantian Touchscreen | Ready to use | LCD touchscreen rusak, telah dilakukan penggantian touchscreen, dan unit berfungsi dengan baik | Marshall
09000719 | 2025-08-29 | RSUD Jailolo Halmahera Barat | Instrument | Bbraun/Aesculap |  |  |  |  | Alat belum tiba di tempat | Ruslan
06000919 | 2025-08-29 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL-0016 | Survey | Cek hit software,cek error log, cek status voltage supply | Pending | menunggu sparepart compare JV dengan teknisi draeger pusat | Dwi
07000769 | 2025-08-28 | RSUD Kota Yogyakarta | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Maintenance | Total care; Penggantian MV2 | Ready to use |  | Wahyu
07000768 | 2025-08-28 | RSUD Kota Yogyakarta | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Maintenance | Total care; Penggantian MV2 | Ready to use |  | Wahyu
12000235 | 2025-08-28 | RSUD Ainun Habibi Gorontalo | Bor Tulang | Bbraun/Acculan 4 | 003075 | Service | Pergantian baterai backup dan uji fungsi unit | Ready to use | Acculan 4 sudah selesai, unit berfungsi dengan baik | Marshall
14000364 | 2025-08-28 | RS TK IV DKT Kediri | Kursi Roda | SYAS/SW01-55 | 250831313000400669 | Instalasi | Pra Install Kursi roda | Ready to use | Alat berfungsi dengan baik | Fahim
14000364 | 2025-08-28 | RS TK IV DKT Kediri | Kursi Roda | SYAS/SW01-55 | 250831313000400668 | Instalasi | Pra Install Kursi roda | Ready to use | Alat berfungsi dengan baik | Fahim
14000364 | 2025-08-28 | RS TK IV DKT Kediri | Kursi Roda | SYAS/SW01-55 | 250831313000400672 | Instalasi | Pra Install Kursi roda | Ready to use | Alat berfungsi dengan baik | Fahim
14000364 | 2025-08-28 | RS TK IV DKT Kediri | Kursi Roda | SYAS/SW01-55 | 250831313000400671 | Instalasi | Pra Install Kursi roda | Ready to use | Alat berfungsi dengan baik | Fahim
14000364 | 2025-08-28 | RS TK IV DKT Kediri | Kursi Roda | SYAS/SW01-55 | 250831313000400670 | Instalasi | Pra Install Kursi roda | Ready to use | Alat berfungsi dengan baik | Fahim
08000718 | 2025-08-28 | RS Mitra Bangsa Pati | Mesin Anasthesia | Draeger/Atlan A300 XL | ASMM-0060 | Instalasi | Install Demo Unit | Ready to use | Device ready to use | Amna
07000759 | 2025-08-28 | RSUD Kota Yogyakarta | Jaundice Meter | Draeger JM -105 | B3503416 | Maintenance | Kalibrasi dan adjustmen di workshop draeger surabaya | Ready to use |  | Wahyu
07000759 | 2025-08-28 | RSUD Kota Yogyakarta | Jaundice Meter | Draeger JM -105 | B3503376 | Maintenance | Kalibrasi dan adjustmen di workshop draeger surabaya | Ready to use |  | Wahyu
02000208 | 2025-08-28 | RSUD Raden Mattaher | Mammography | Metaltronica/Helianthus DBT | HELIDBT/026/C0 | Survey | Perlu penggantian grid | Pending | Akan ditawarkan part grid helianthus DBT | Eddy
02000207 | 2025-08-28 | RSUD Raden Mattaher | Digital Radiography Fluoroscopy | Apelem/Platinum | G176F0223 | Service | Cek kerusakan touch screen dan hardware/software | Pending | Perlu penggantian part LCD display set touch screen, akan dibuat penawaran perbaikan | Eddy
11000433 | 2025-08-27 | RSUP Makassar | Ventilator | Draeger/Babylog VN800 | ASTM0102 | Maintenance | System check, device check | Ready to use |  | Nasrul
11000433 | 2025-08-27 | RSUP Makassar | Ventilator | Draeger/Babylog VN800 | ASTM0103 | Maintenance | System check, device check | Ready to use |  | Nasrul
03000398 | 2025-08-27 | RSUP Jayapura | Digital Radiography Fluoroscopy | Apelem/Platinum | H248L0157 | Maintenance | Preventive maintenance | Ready to use |  | Imam
02000206 | 2025-08-27 | RSUD Raden Mattaher | X-ray Digital | Apelem/Camargue | IMA 703863 | Service | Penggantian part console board dan display LCD modified | Ready to use | Perbaikan sesuai penawaran, console xray camargue sudah ok | Eddy
18000627 | 2025-08-26 | RSU GMIM Bethesda | Ventilator | Draeger/Savina 300 | ASTJ-0024 | Service | setelah di cek, ada kerusakan pada flow sensor | Ready to use | Sudah ganti flow sensor baru yang cadangan | Irfan
09000718 | 2025-08-26 | RSUD Sanana | UPS | Yoshiga/GPI 80 KVA | 501212003370N3800001 | Service | Ganti konektor SR | Ready to use | Unit sudah berfungsi normal | Ruslan
07000758 | 2025-08-26 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Babylog VN800 | ASUA-0006 | Instalasi | Menjelaskan cara menggunakan alat | Ready to use | device ready to use | Wahyu
08000837 | 2025-08-26 | RSUD dr. Adhyatma | Bed Patient | Draeger/Infinity Acute C500 | ASUA0117 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
18000613 | 2025-08-25 | RS Mata Manado | Mesin Anasthesia | Draeger/Perseus A500 | ASJM-0178 | Service | Unit garansi kontrak service total care | Ready to use | penggantian baterai pack | Irfan, Marshall
08000717 | 2025-08-25 | RSUD Jepara | Ventilator | Draeger/Babylog VN 500 | ASMK-0029 | Survey | Cek sistem, cek insert neo flow, cek expiratory valve, running test | Ready to use |  | Amna
08000717 | 2025-08-25 | RSUD Jepara | Ventilator | Draeger/Babylog VN 500 | ASMK-0028 | Survey | Cek sistem, cek insert neo flow, cek expiratory valve, running test | Ready to use |  | Amna
07000757 | 2025-08-25 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | V3SHK0157 | Survey | Post test, cek setting MBP, cek module NIBP, Cek NIBPPump, NIBP pump leak | Ready to use | Unit perlu penggantian X2 NIBP module | Wahyu
04000515 | 2025-08-25 | RS Al-Islam Bandung | Infus Pump | Bbraun/Infusomat Compact Plus | 11335 | Maintenance | Preventive maintenance ke 1 | Ready to use | Alat berfungsi dengan baik | Satria, Bahar
04000515 | 2025-08-25 | RS Al-Islam Bandung | Syringe Pump | Bbraun/Perfusor Compact Plus | 102716 | Maintenance | Preventive maintenance ke 1 | Ready to use | Alat berfungsi dengan baik | Satria, Bahar
05000365 | 2025-08-25 | RS Suyoto | Dental Unit | KAVO/1058 Life | 2021-01-10207193 | Survey | Pengecekan pada display error messages dan muncul pesan oxygenal empty | Ready to use | Untuk sekarang menggunakan stock RS dan sudah ditawarkan pembelian oxygenal 6 | Farhat, Ipan
20000386 | 2025-08-24 | RSUP Makassar | C-arm | Apelem/Evo R+ | 551240241086 | Survey | Camera disconnected, pengecekan connection console dengan IP Address; Penggantian RAM Console | Ready to use |  | Sujana
09000709 | 2025-08-23 | RSD Tidore | UPS | OAXAKA/Powerindo UP-6 | 410044367F65225600001 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000709 | 2025-08-23 | RSD Tidore | UPS | OAXAKA/Powerindo UP-6 | 410044367F65225600003 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000709 | 2025-08-23 | RSD Tidore | UPS | OAXAKA/Powerindo UP-6 | 410044367F65225600002 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000709 | 2025-08-23 | RSD Tidore | UPS | OAXAKA/Powerindo UP-6 | 410044367F65225600004 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
06000916 | 2025-08-22 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASLD-0042 | Maintenance | Preventive Maintenance | Ready to use |  | Dwi
19000486 | 2025-08-22 | RSUP dr Kariadi | USG | Mindray/Resona A20 | NX7-57000142 | Instalasi | Install Unit demo | Ready to use | Unit sudah terinstall dengan baik | Nanda
12000244 | 2025-08-22 | RSUD Ainun Habibi Gorontalo | MOT | Syas Tech | - | Survey | Pendampingan BPFK untuk pemeriksaan | Ready to use | MOT Syas telah selesai, unit berfungsi dengan baik | Marshall
09000708 | 2025-08-22 | RSD Tidore | UPS | OAXAKA/Powerindo UP-1 | 310044379F65063400006 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000708 | 2025-08-22 | RSD Tidore | UPS | OAXAKA/Powerindo UP-1 | 310044379F65063400004 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000708 | 2025-08-22 | RSD Tidore | UPS | OAXAKA/Powerindo UP-1 | 310044379F65063400001 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000708 | 2025-08-22 | RSD Tidore | UPS | OAXAKA/Powerindo UP-1 | 310044379F65063400003 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000708 | 2025-08-22 | RSD Tidore | UPS | OAXAKA/Powerindo UP-1 | 310044379F65063400002 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
09000708 | 2025-08-22 | RSD Tidore | UPS | OAXAKA/Powerindo UP-1 | 310044379F65063400005 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
08000704 | 2025-08-22 | RS Mitra Bangsa Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ-0041 | Survey | General check, system check, check O2 sensor, replace flow sensor, running test | Ready to use |  | Amna
07000756 | 2025-08-22 | RSUD Merah Putih | Baby Incubator | Draeger/Isolette C2000 | DSRL-0028 | Service | Post test, kelembapan rendah, evaporator humidifier, evaporator berkerak, bersihkan evaporator humidifier, post test | Ready to use |  | Wahyu
11000420 | 2025-08-21 | RSUP Dr. Wahidin Sudirohusodo | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/087/C1 | Service | Daily check failed; Cek tegangan dan kalibrasi detector. Restart unit | Ready to use |  | Nasrul
07000755 | 2025-08-21 | RSUD Wonosari | Ventilator | Draeger/Babylog VN800 | ASPF-0003 | Service | Device check, cek exp valve, cek mebran exp valve, cek INSP valve, bersihkan exp valve, membran exp valve, INSP valve,device check, post test | Ready to use |  | Wahyu
06000917 | 2025-08-21 | RSUD Prof. Dr. Margono Soekarjo | Power Supply | Draeger Univ Power Supply | MS18284 | Instalasi | Instalation | Ready to use | Power supply ok | Dwi
19000443 | 2025-08-20 | RS Telegorejo Semarang | Mesin Anasthesia | Draeger/Primus | ASDK-0223 | Survey | Mixer ada error, coba exchange dengan sensor ptank unit demo draeger, unit sedang running test | Ready to use | unit sedang di running dengan sensor ptank dari PT Draeger medical indonesia | Nanda
15000311 | 2025-08-20 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Fluoroscopy | Apelem/Platinum | L156F0142 | Service | Cek kabel exphose | Ready to use | Kabel ada yang putus, sudah dilakukan penyolderan, alat dapat digunakan kembali | Sutan
12000232 | 2025-08-20 | RSUD Zainal Umar Sidiki | Foot Step | Syas Tech | - | Instalasi | Install, uji fungsi, dan training user | Ready to use | Install foot step 2 unit | Marshall
12000232 | 2025-08-20 | RSUD Zainal Umar Sidiki | Wheel Chair | Syas Tech | - | Instalasi | Install, uji fungsi, dan training user | Ready to use | Install wheel chair 3 unit | Marshall
12000232 | 2025-08-20 | RSUD Zainal Umar Sidiki | Examination Table | Syas Tech | - | Instalasi | Install, uji fungsi, dan training user | Ready to use | Install examination table 2 unit | Marshall
12000232 | 2025-08-20 | RSUD Zainal Umar Sidiki | Infused Stand | Syas Tech | - | Instalasi | Install, uji fungsi, dan training user | Ready to use | Instalasi infused stand 25 unit | Marshall
08000719 | 2025-08-20 | RSUD Sunan Kalijaga Demak | Mesin Anasthesia | Draeger/Primus | ASHD-0034 | Survey | Check breathing system, check diapgram piston, check compact breathing system, running test | Ready to use |  | Amna
07000752 | 2025-08-20 | RS JIH Yogyakarta | Mesin Anasthesia | Draeger/Primus | ASMA-0144 | Service | Relace waterlock, replace sample line, cek dengan visia 2, internal kalibrasi O2 sensor, post test | Ready to use |  | Wahyu
07000753 | 2025-08-20 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASJH-0125 | Service | Relace Internal Baterai Pack M7 NIMH | Ready to use |  | Wahyu
03000397 | 2025-08-20 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | Heli DBT/129/C2 | Service | Ada USB dongle mouse optic yang nyolok di USB hub mengakibatkan keyboard touch tidak aktif | Ready to use |  | Imam
18000616 | 2025-08-19 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator Neonatal | Draeger/Babylog VN 500 | ASJD-0019 | Service | Setelah di cek, kerusakan pada pressure regulator flow sensor dan kabel flow sensor | Ready to use | Penggantian part pressure regulator flow sensor dan kabel flow sensor/cable harnes, housing flow | Irfan
08000716 | 2025-08-19 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASEL-0138 | Service | Run system test, LCD fail | Pending | Butuh penggantian LCD | Amna
07000751 | 2025-08-19 | RSUD Saras Adyatma | Patient Monitor | Draeger/Vista 120 | VGSTC0078 | Survey | Cek kabel power, cek fuse, cek baterai, cek power supply, fuse ps putus, trafo ps terbakar, unit perlu penggantian power PCBA 2607486 | Pending | Ajukan claim warranty | Wahyu
06000915 | 2025-08-19 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Babylog VN800 | ASUA-0007 | Instalasi | Unit berfungsi dengan baik | Ready to use | Accesories lengkap sesuai packing list surat jalan | Dwi
06000915 | 2025-08-19 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Babylog VN800 | ASTN-0005 | Instalasi | Unit berfungsi dengan baik | Ready to use | Accesories lengkap sesuai packing list surat jalan | Dwi
06000915 | 2025-08-19 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Babylog VN800 | ASUA-0006 | Instalasi | Unit berfungsi dengan baik | Ready to use | Accesories lengkap sesuai packing list surat jalan | Dwi
06000915 | 2025-08-19 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Babylog VN800 | ASTN-0004 | Instalasi | Unit berfungsi dengan baik | Ready to use | Accesories lengkap sesuai packing list surat jalan | Dwi
20000393 | 2025-08-19 | RSUP Makassar | MOT | SYAS | - | Service | HVAC kurang dingin, pengecekan dan setting suhu unit | Ready to use |  | Sujana
15000310 | 2025-08-15 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Fluoroscopy | Apelem/Platinum | L156F0142 | Service | Pengecekan sensor grid dan pengecekan motor grid | Ready to use | Grid tidak dapat bergerak dan sudah di perbaiki, alat sudah dapat digunakan | Sutan
06000920 | 2025-08-15 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ-0082 | Survey | General check | Ready to use | Rekomendasi dan pengantian service set 3y , O2 Cell, Breathsys Vbar 2y, baterai pack, water lock, sample line | Dwi
06000920 | 2025-08-15 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL-0039 | Survey | General check | Ready to use | Rekomendasi dan pengantian service set 3y , O2 Cell, Breathsys Vbar 2y, baterai pack, water lock, sample line | Dwi
06000920 | 2025-08-15 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL-0027 | Survey | General check | Ready to use | Rekomendasi dan pengantian service set 3y , O2 Cell, Breathsys Vbar 2y, baterai pack, water lock, sample line | Dwi
06000920 | 2025-08-15 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL-0025 | Survey | General check | Ready to use | Rekomendasi dan pengantian service set 3y , O2 Cell, Breathsys Vbar 2y, baterai pack, water lock, sample line | Dwi
06000920 | 2025-08-15 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Tiro | ASBM-0248 | Survey | General check | Ready to use | Rekomendasi dan pengantian service set 3y , O2 Cell, Breathsys Vbar 2y, baterai pack, water lock, sample line | Dwi
06000920 | 2025-08-15 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL-0047 | Survey | General check | Ready to use | Rekomendasi dan pengantian service set 3y , O2 Cell, Breathsys Vbar 2y, baterai pack, water lock, sample line | Dwi
04000514 | 2025-08-15 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Service | Uji fungsi EMH baru | Ready to use | EMH bagus, Tetapi HVG perlu perbaikan | Satria, Bahar
02000203 | 2025-08-15 | RSUD Kardinah (KSO) | CT Scan | Philips/MX 16 Slice | EP16E140017 | Service | PIU dan tube x-ray dibawa ke jakarta untuk perbaikan | Pending | Warming tube tidak bisa | Eddy
20000385 | 2025-08-15 | RSUP Tadjudin Chalid | Mamografi | Helianthus DBT | HELIDBT/133/C2 | Service | PACS error, pengecekan connection kabel; Setting IP address baru | Ready to use |  | Sujana, Nasrul
2000392 | 2025-08-14 | RSUP Makassar | C-arm | Apelem/Evo R+ | 551240241086 | Service | Camera disconnected; Set IP Address console c-arm | Ready to use |  | Sujana
11000434 | 2025-08-14 | RSUP Makassar | Instrument | Aesculap | 2511300 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000432 | 2025-08-14 | RSUP Tadjudin Chalid | Digital Radiografi Fluoroskopi | Apelem/ Platinum | J236F0882 | Service | Console tidak bisa booting; Aktifkan emergency button | Ready to use |  | Nasrul
19000442 | 2025-08-14 | RSUP dr Soeprapto Cepu | Ventilator | Draeger/Savina 300 | ASJK-02 II | Service | Cek unit, cek flow sensor, ada yang putus ganti flow sensor yang baru, running, ok | Ready to use | Disarankan Penggantian, service set 6 tahun (usia alat dari tahun 2016) | Nanda
12000243 | 2025-08-14 | RSUD Ainun Habibi Gorontalo | MOT | Syas Tech | - | Service | Ganti pengunci lemari dengan yang baru | Ready to use | Pengunci lemari rusak, sehingga diganti dengan pengunci yang baru (warranty) | Marshall
18000626 | 2025-08-14 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Fabius | ASFM-0036 | Survey | Setelah di cek ada kesumbatan | Ready to use | Sudah di bersihkan | Irfan
09000717 | 2025-08-14 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 | ASTJ-2004 | Service | Penarikan Unit demo |  |  | Ruslan
09000717 | 2025-08-14 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 | ASSM-0191 | Service | Penarikan Unit demo |  |  | Ruslan
07000749 | 2025-08-14 | RSUD Wonosari | Jaundice Meter | Draeger/JM-105 | H3512949 | Instalasi | Telah dilakukan instalasi, uji fungsi,dan training alat jaundice meter JM-105 | Ready to use | Device ready to use | Wahyu
06000914 | 2025-08-14 | RSUD Dr. Soeselo Slawi | Ventilator | Draeger/Savina 300 | ASHD-0014 | Survey | Cek semu mode, cek valve 8/9, cek blower, cek hourtime | Pending | Rekomendasi pengantian sparepart motor blower PN: 8413643 Driver motor PN: 8413620 Savina ServSet 6y PN: MX08861 O2 sensor 2pcs | Dwi
05000364 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Operating Table Electro Hydraulic | Syas/SOH 2 | 250767222000400014 | Instalasi | Instalasi unit | Ready to use | Unit dapat digunakan | Farhat
05000363 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Anestesi Trolley | Syas/Syas-SCA1-KS22 | 250736607000400013 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000363 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Anestesi Trolley | Syas/Syas-SCA1-KS22 | 250736607000400012 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000363 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Anestesi Trolley | Syas/Syas-SCA1-KS22 | 250736607000400011 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000362 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Mayo Stand | Syas/SCY 1 | 250635113000400112 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000362 | 2025-08-14 | signature Dental Care | Mayo Stand | Syas/SCY 1 | 250635113000400091 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400356 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400351 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400348 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400354 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400350 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400347 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400352 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400358 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400353 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400357 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400349 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000361 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Instrument Trolley | Syas/SCI 1 | 250769101000400355 | Instalasi | Instalasi unit | Ready to use |  | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401161 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401168 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401166 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401176 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401163 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401162 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401180 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401169 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401172 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401170 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401178 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401167 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401177 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401174 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401164 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401171 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401175 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401165 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401173 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
05000360 | 2025-08-14 | RSUD Dr Soedarso Pontianak | Syas Infused Stand | Syas/SIS 1 | 250636163000401179 | Instalasi | Instalasi Unit | Ready to use | Unit dapat digunakan | Farhat
02000204 | 2025-08-14 | RSUD R. Syamsudin (KSO) | CT Scan | Philips/MX 16 Slice | EP16E130062 | Maintenance | Tes tube dan TCU CT Scan RS Kardinah Tegal | Pending | Perlu perbaikan tube dan PIU | Eddy
04000513 | 2025-08-14 | RSUD Alimuddin Umar | Digital Radiography Flouroscopy | Apelem/Platinum | M196F0402 | Survey | Input distribution board (IDB) tidak memberikan supply ke board yang lain | Pending | Tidak bisa dilakukan lebih lanjut karena bermasalah di input distribution board (IDB) | Satria, Bahar
08000712 | 2025-08-14 | RSUD RA Kartini Jepara | Mesin Anasthesia | Draeger/Fabius Plus | ASEH-0118 | Survey | Check LCD, Cal. LCD, Running test | Pending | Needs replacement LCD | Amna
19000441 | 2025-08-13 | RSUP dr Kariadi | Dental Unit | Kavo 1058 | 2004-1002722 | Survey | Cek unit, coba bongkar, ada sistem motor pump yang mati | Ready to use | Unit bisa digunakan sementara tanpa 3 way assistan; disarankan ganti 3 way | Nanda
18000624 | 2025-08-13 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASHD-0027 | Survey | Setelah di cek baterai sudah drop | Pending | Disarankan untuk mengganti baterai baru | Irfan
11000431 | 2025-08-13 | RSUP Makassar | Mayo Stand | SYAS | 25063511300400107 | Instalasi | Unit berfungsi dengan baik | ready to use |  | Nasrul, Sujana
09000716 | 2025-08-13 | RSD Tidore | Ventilator | Draeger/Babylog VN800 HFO | ASTN-0001 | Survey | Setting batas atas dan batas bawah alarm, cek kebocoran pada ujung breathing, refresh training identifikasi penyebab alarm | Ready to use |  | Ruslan
07000748 | 2025-08-13 | RSA UGM | Vaporizer | Draeger/Vapor 2000 | ASHH-0861 | Survey | Post test, adjustmen vaporizer, post test | Ready to use |  | Wahyu
07000747 | 2025-08-13 | RSA UGM | Infant Warmer | Draeger/Resuscitaire | AC13541 | Survey | Cek pressure gauge, adjustmen pressure gauge, pegas pressure gauge rusak | Pending | Resuscitaire belum dapat digunakan, unit perlu penggantian pressure gauge | Wahyu
04000512 | 2025-08-13 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Service | Mengganti EMH dengan yang baru, Penggantian baterai CMOS | Pending | EMH tidak bisa menembak, HVG bermasalah, Menunggu keputusan rumah sakit untuk HVG dan HVPS | Satria, Bahar
08000715 | 2025-08-13 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHA-0113 | Survey | Replace PS M73 Plus, reload software, battery check, running test |  | Proses cek baterai 9 jam | Amna
20000391 | 2025-08-13 | RSUP Makassar | Meja Mayo | SYAS | 250635113000400107 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana
20000391 | 2025-08-13 | RSUP Makassar | Meja Mayo | SYAS | 250635113000400105 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana
20000391 | 2025-08-13 | RSUP Makassar | Meja Mayo | SYAS | 250635113000400110 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana
20000391 | 2025-08-13 | RSUP Makassar | Meja Mayo | SYAS | 250635113000400109 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana
20000391 | 2025-08-13 | RSUP Makassar | Meja Mayo | SYAS | 250635113000400106 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana
20000391 | 2025-08-13 | RSUP Makassar | Meja Mayo | SYAS | 250635113000400114 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana
20000391 | 2025-08-13 | RSUP Makassar | Meja Mayo | SYAS | 250635113000400113 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana
20000391 | 2025-08-13 | RSUP Makassar | Meja Mayo | SYAS | 250635113000400111 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana
20000391 | 2025-08-13 | RSUP Makassar | Meja Mayo | SYAS | 250635113000400090 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana
20000391 | 2025-08-13 | RSUP Makassar | Meja Mayo | SYAS | 250635113000400108 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Sujana
8000761 | 2025-08-13 | RSUD Kab. Temanggung | USG | MIndray/Consona N8 | LP7U56000376T | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
19000440 | 2025-08-12 | RSUD Pandan Arang Boyolali | Ventilator Transport | Draeger/Oxylog 2000 | ASKL-0036 | Survey | Cek alat, ada valve yang hilang/rusak | Pending | Disarankan penggantian valve set dan flow sensor | Nanda
19000440 | 2025-08-12 | RSUD Pandan Arang Boyolali | Ventilator Transport | Draeger/Oxylog 2000 | ASKL-0040 | Survey | Cek alat, ada valve yang hilang/rusak | Pending | Disarankan penggantian valve set dan flow sensor | Nanda
12000231 | 2025-08-12 | RSU GMIM Bethesda | Ventilator | Draeger/Savina 300 | ASTJ-0022 | Survey | Pengecekan unit | Pending | Setelah pengecekan unit, ada kerusakan pada valve mixing ( no part; 8413603), Unit masih warranty | Marshall, Irfan
12000230 | 2025-08-12 | RS GMIM Bethesda | Ventilator | Draeger/Savina 300 | ASTJ-0025 | Survey | Pengecekan unit | Pending | Setelah pengecekan unit, ada kerusakan pada valve mixing ( no part; 8413603), Unit masih warranty | Marshall, Irfan
12000229 | 2025-08-12 | RSUD Tondano | Ventilator | Draeger/Savina 300 | ASJM-0392 | Survey | Pengecekan unit | Ready to use | unit masih berfungsi dengan baik, tetapi ada kerusakan pada baterai internal, disarankan untuk mengganti baterai (masa pakai baterai sudah melebihi 2 tahun) | Marshall, Irfan
07000746 | 2025-08-12 | RS Panti Rapih | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ-0084 | Survey | Cek keypad, cek sambungan kabel keypad, servis mode, test keypad | Pending | Ajukan claim warranty, Unit perlu penggantian panel keypad | Wahyu
06000913 | 2025-08-12 | RSUD Ajibarang | CPAP | Draeger/Oxymixer | QM2025047 | Maintenace | Preventive maintenance dan pemasangan air regulator watertrap 3y do oxymixer Daeger CPAP | Ready to use | Warranty | Dwi
06000913 | 2025-08-12 | RSUD Ajibarang | CPAP | Draeger/Oxymixer | QM2025046 | Maintenace | Preventive maintenance dan pemasangan air regulator watertrap 3y do oxymixer Daeger CPAP | Ready to use | Warranty | Dwi
06000913 | 2025-08-12 | RSUD Ajibarang | CPAP | Draeger/Oxymixer | QM2025044 | Maintenace | Preventive maintenance dan pemasangan air regulator watertrap 3y do oxymixer Daeger CPAP | Ready to use | Warranty | Dwi
08000690 | 2025-08-12 | RSUD RAA Soewondo Pati | Infant Warmer | Draeger/Resuscitaire | AC 13489 | Survey | Cek 3 way failure, device not resusitasi | Pending | Needs replacement 3 way part ; breating bag, resusitasi, mask silicone | Amna
03000396 | 2025-08-12 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Fluoroscopy | Apelem/Platinum | L156F0142 | Service | Perbaikan PC WS duet reload software duet | Ready to use |  | Imam
20000387 | 2025-08-12 | RS Islam Faisal | ESWL | Medispec/EM1000 | 1114075 | Service | Penggantian EMH | Pending | EMH di retur karna pihak rumah sakit belum bisa menerima EMH baru | Sujana, Nasrul
19000435 | 2025-08-11 | RS Nasional Diponegoro Semarang | Incubator | Draeger/Isolette C2000 | AF32539 | Service | Cek alat, cek unit, ada soket timbangan dan temperature yang miring sehingga tidak muat pada soket; re connect, setting soket, OK | Ready to use |  | Nanda
18000623 | 2025-08-11 | RS Mata Manado | Mesin Anasthesia | Draeger/Perseus | ASJM-0178 | Maintenance | Visit bersama tim draeger | Ready to use | Kunjungan kontrak service ke dua | Irfan
07000745 | 2025-08-11 | RS JIH Yogyakarta | Mesin Anasthesia | Draeger/Primus | ASMA-0144 | Maintenance | Replace serv set 3y, replace breathing system set 2y, pgm function test, vgc function test, mixer function test, post test, cleaning bust | Ready to use |  | Wahyu
07000745 | 2025-08-11 | RS JIH Yogyakarta | Mesin Anasthesia | Draeger/Primus | ASMA-0128 | Maintenance | Replace serv set 3y, replace breathing system set 2y, pgm function test, vgc function test, mixer function test, post test, cleaning bust | Ready to use |  | Wahyu
19000485 | 2025-08-08 | RSUD Mijen Semarang | USG | Mindray/Z50T | HH3B2B000013 | Survey | Cek alat, cek probe, ok, konektor probe kurang pas | Ready to use | Unit sudah ok | Nanda
18000617 | 2025-08-08 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator Neonatal | Draeger/Babylog VN500 | HSJC-0048 | Service | Setelah di cek, kerusakan pada flow sensor, pressure regulator dan kabel flow sensor | Ready to use | Pergantian part, pressure regulator, flow sensor dan kabel flow sensor/cable harnes, housing flow | Irfan
07000744 | 2025-08-08 | RSUP dr Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASHA-0106 | Survey | Leak test, leak pump, cal pump, cek selang ventilator, cek safety relief valve, post test | Ready to use | Unit perlu penggantian; serv set 3y, breathing system serv set, O2 sensor | Wahyu
08000714 | 2025-08-08 | RSI Sultan Agung Semarang | Mesin Anasthesia | Draeger/Atlan A350 XL | ASRL-0101 | Survey | System check, check cozy, check absorber, check breathing system, device ready to use | Pending | Recomended to CO2 absorber ori draeger HCA only | Amna
2000388 | 2025-08-08 | RSUP makassar | MOT | SYAS | - | PraInstall | Installasi box frame | Ready to use |  | Sujana, Nasrul
11000430 | 2025-08-07 | RSUP Tadjudin Chalid | Ventilator | Draeger/Babylog VN800 | ASSJ0037 | Service | Flow tidak terbaca; Penggantian flow sensor | Ready to use |  | Nasrul
19000484 | 2025-08-07 | RS Colombia Asia Semarang | Infant warmer | Draeger/Babytherm 8010 | ARUC-1025 | Survey | Cek alat, sensor temperatur rusak, ganti sensor dengan stock dari RS, ok | Ready to use |  | Nanda
07000800 | 2025-08-07 | RSUD Purworejo | Ventilator | Draeger/Babylog VN 800 | ASPC-0048 | Instalasi | pengecekan kelengkapan alat dan fungsi alat | Ready to use | Uninstall unit demo | Wahyu
06000911 | 2025-08-07 | RSU Nirmala Purbalingga | Ventilator | Draeger/Savina 300 | ASRE-0307 | Service | Pemeriksaan flow sensor, self test unit, compare flow sensor | Pending | Flow sensor wajib penggantian | Dwi
08000713 | 2025-08-07 | RSUD KRT Setjonegoro | Mammography | Metaltronica/Helianthus DBT | HELI DBT/044/C0 | Service | Pemindahan alat mammography |  | Karena Ruangan akan di renovasi | Amna
03000395 | 2025-08-07 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELI DBT/129/C2 | Maintenance | Preventive maintenance | Ready to use | Unit berfungsi dengan baik | Imam
20000389 | 2025-08-07 | RSUP Makassar | MOT | SYAS | - | PraInstall | Installasi cabinet | Ready to use |  | Sujana
11000429 | 2025-08-06 | RSUP Dr. Wahidin Sudirohusodo | Lampu Operasi | Draeger/Polaris 600 | ASFL0280 | Service | LCD monitor kamera burem; Pengecekan LCD, LCD rusak | Pending |  | Nasrul
19000434 | 2025-08-06 | RSU Telogorejo Semarang | Mesin Anasthesia | Draeger/Primus | ASDK-0223 | Service | Cek alat, tarik error, ada indikasi keruskaan pada valve ptank/mixer PCB | Pending | Akan dicoba dengan part dari tim draeger | Nanda
15000309 | 2025-08-06 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Fluoroscopy | Apelem/Platinum | L156F0142 | Service | Cek UPS ok, cek power suplay tidak ada listrik, cek power suplay ok | Ready to use |  | Sutan
12000228 | 2025-08-06 | RSUD Zainal Umar Sidiki | Ventilator | Draeger/Savina 300 | ASTJ-2002 | Instalasi | Install unit backup dan uji fungsi | Ready to use |  | Marshall
03000394 | 2025-08-06 | RSUD Pagelaran | Mamografi | Metaltronica/Helianthus DBT | HELI DBT/136/C2 | Maintenance | Preventive maintenance | Ready to use | Unit berfungsi dengan baik | Imam
07000743 | 2025-08-06 | RSUP Dr Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD-0026 | Survey | System test, cek error log, cek waterlock dan sampling line, replace waterlock dan sampling line | Ready to use |  | Wahyu
08000711 | 2025-08-06 | RS As-Suyuthiyyah Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASUA-0003 | Maintenance | System check, check o2 sensor, check flow sensor, running test, device ready to use | Ready to use | Unit berfungsi dengan baik | Amna
04000511 | 2025-08-06 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Fluoroscopy | Apelem/Platinum | L156F0142 | Survey | Uji kesesuaian keperluan perpanjang UKES Gratia | Ready to use |  | Satria, Bahar, Sutan
19000433 | 2025-08-05 | RSUD dr Soewondo Kendal | Ventilator | Draeger/Babylog VN 500 | ASNL-0149 | Survey | Cek unit, ganti dengan flow sensor yang baru, running , ok | Ready to use | Unit sudah diganti dengan flow sensor yang baru | Nanda
15000308 | 2025-08-05 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Fluoroscopy | Apelem/Platinum | L156F0142 | Service | Install ulang software | Ready to use | Hardisk dan RAM akan diganti | Sutan
18000622 | 2025-08-05 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASFM-0036 | Survey | Setelah di cek, ada kerusakan pada flow power, minimodul | Pending | Disarankan untuk mengganti part yang bermasalah | Irfan
09000707 | 2025-08-05 | RSD Tidore | Mesin Anasthesia | Draeger/Atlan A300 | ASUF-0382 | Instalasi | Install unit, uji fungsi, dan training | Ready to use |  | Ruslan
08000710 | 2025-08-05 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHA-0113 | Survey | Hit software, tarik error, cek power supply | Pending | Device not ready to use, reschedule, identification | Amna
05000359 | 2025-08-05 | RSUD Berkah Pandeglang | Digital Radiography Flouroscopy | Apelem/Platinum | J206FO497 | Service | Perbaikan dan pembersihan kontraktor pada apanel listrik utama | Ready to use |  | Farhat,Eddy
20000400 | 2025-08-04 | RSUP Makassar | MOT | SYAS | - | PraInstall | Install dudukan pendant | Ready to use |  | Sujana, Nasrul
19000483 | 2025-08-04 | RST Bhakti Wira Tamtama Semarang | CPAP | Draeger/Seattle PAP | QM20255073 | Instalasi | Install Unit | Ready to use | User sudah di training | Nanda
19000483 | 2025-08-04 | RST Bhakti Wira Tamtama Semarang | CPAP | Draeger/Seattle PAP | QM20255076 | Instalasi | Install Unit | Ready to use | User sudah di training | Nanda
07000742 | 2025-08-04 | RSPAU Dr. S. Hardjoloekito | Ventilator Transport | Draeger/Oxylog 2000 Plus | SSAD-0006 | Survey | Cek device, cek error, cek servis mode, cek display supply voltage, VNEG dan VDA faulty | Pending | Unit perlu penggantian prog pba controler (5705018) | Wahyu
07000741 | 2025-08-04 | RSPAU Dr. S. Hardjoloekito | Ventilator Transport | Draeger/Oxylog 3000 Plus | ASSB-0125 | Survey | Device check, cek error log, v9 dan v8 faulty, cek sensor board, cek board pba controler | Pending | Unit perlu penggantian board pba controler (5790841) | Wahyu
06000910 | 2025-08-04 | RSU An Ni'mah | Ventilator | Draeger/Savina 300 | ASPJ-0377 | Survey | Cek self test, running test, cross check IMT analyzer | Ready to use | Rekomendasi perbaikan central O2/cek central O2 | Dwi
05000358 | 2025-08-04 | RSUD Cilegon | Digital Radiography Flouroscopy | Apelem/Platinum | J206F0488 | Service | Update software adam dari 1.2 sampai 1.4, update platinum view 3.0 samapi 4.2 | Ready to use |  | Farhat,Eddy
20000400 | 2025-08-04 | RSUP Makassar | Pendant MOT | SYAS | - | PraInstall | Installasi breaket dudukan | Ready to use |  | Sujana
2000390 | 2025-08-03 | RSUP Makassar | MOT | SYAS | - | PraInstall | Installasi dudukan pendant | Ready to use |  | Sujana
15000307 | 2025-08-01 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-Ray (KSO) | Apelem/Rafale EV 30 | 009/15/01095 | Service | Bersihkan board S22, cek fuse ok | Ready to use |  | Sutan
09000715 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Evita V800 | ASNK-0340 |  | Penarikan Unit demo | Ready to use | Unit sudah di packing dan siap untuk di kirim | Ruslan
07000740 | 2025-08-01 | RSUD Wonosari | C-Arm | Apelem/Evo R+ | 551-21-060-911 | Service | Cek nyala, cek kabel power, cek fuse, cek emergency stop, post test, test expose | Ready to use |  | Wahyu
08000709 | 2025-08-01 | RSUD RA Kartini Jepara | Mesin Anasthesia | Draeger/Fabius Plus | ASMK-0138 | Service | Pemasangan o2 sensor capsule, system check, running test | Ready to use | Device ready to use | Amna
08000709 | 2025-08-01 | RSUD RA Kartini Jepara | Mesin Anasthesia | Draeger/Fabius Plus | ASEH-0118 | Service | Pemasangan o2 sensor capsule, system check, running test | Ready to use | Device ready to use | Amna
08000709 | 2025-08-01 | RSUD RA Kartini Jepara | Mesin Anasthesia | Draeger/Fabius Plus | ASMK-0136 | Service | Pemasangan o2 sensor capsule, system check, running test | Ready to use | Device ready to use | Amna
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671801 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671793 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671842 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671838 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671796 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671758 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671816 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671789 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671813 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671869 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671863 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671861 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671865 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671811 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671867 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000599 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Syringe Pump | Bbraun/Perfusor Space | 671852 | Instalasi | Install , Uji fungsi, dan training syringe pump | Ready to use |  | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 964924 | Instalasi | Install, uji fungsi, dan ada kerusakan pada baterai | Pending | Sementara proses klaim garansi battery | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 814166 | Instalasi | Install, uji fungsi, dan ada kerusakan pada baterai | Pending | Sementara proses klaim garansi battery | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 964928 | Instalasi | Install, uji fungsi, dan ada kerusakan pada baterai | Pending | Sementara proses klaim garansi battery | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 964927 | Instalasi | Install, uji fungsi, dan ada kerusakan pada baterai | Pending | Sementara proses klaim garansi battery | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 964923 | Instalasi | Install, uji fungsi, dan ada kerusakan pada baterai | Pending | Sementara proses klaim garansi battery | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 964351 | Instalasi | Install, uji fungsi, dan ada kerusakan pada baterai | Pending | Sementara proses klaim garansi battery | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 814169 | Instalasi | Install, uji fungsi, dan ada kerusakan pada baterai | Pending | Sementara proses klaim garansi battery | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 964926 | Instalasi | Install, uji fungsi, dan ada kerusakan pada baterai | Pending | Sementara proses klaim garansi battery | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 814161 | Instalasi | Install dan uji fungsi | Ready to use | Unit berfungsi dengan baik | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 964928 | Instalasi | Install dan uji fungsi | Ready to use | Unit berfungsi dengan baik | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 814162 | Instalasi | Install dan uji fungsi | Ready to use | Unit berfungsi dengan baik | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 814167 | Instalasi | Install dan uji fungsi | Ready to use | Unit berfungsi dengan baik | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 814168 | Instalasi | Install dan uji fungsi | Ready to use | Unit berfungsi dengan baik | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 814170 | Instalasi | Install dan uji fungsi | Ready to use | Unit berfungsi dengan baik | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 814172 | Instalasi | Install dan uji fungsi | Ready to use | Unit berfungsi dengan baik | Ruslan
09000598 | 2025-08-01 | RSUD Dr. H. Chasan Boesoirie | Infuse Pump | Bbraun/Infusomat Space P | 964917 | Instalasi | Install dan uji fungsi | Ready to use | Unit berfungsi dengan baik | Ruslan
05000357 | 2025-08-01 | RS Dr. Suyoto | Dental Unit | KAVO/Primus 1058 | 2021-01-10207193 | Service | Dilakukan pembersihan pada penampungan dan potong selang yang bocor | Ready to use | Unit berfungsi dengan baik | Farhat,Ipan
20000384 | 2025-07-31 | RSUP Makassar | Pendant MOT | SYAS Pegasus |  | Instalasi | Instalasi pendant MOT gedung B (Anasthesia) | Pending | Unit masih belum lengkap accesories | Sujana, Nasrul
04000510 | 2025-07-31 | RSUD Al Ihsan | Bor Tulang Neuro | Bbraun/Elan4 | 008095 | Service | Pengecekan fisik, input data SN ke E-tracking |  | Input data SN ke E tracking berhasil, pengecekan selanjutnya di pegang teknisi Bbraun, Handpiece dibawa teknisi Bbraun | Satria, Bahar
19000432 | 2025-07-31 | RSUD Dr Soeratno Gemolong | Jaundice Meter | Draeger/JM105 | B3503803 | Service | Cek unit, error 05, calibration date EXP | Pending | Unit di tarik untuk di kalibrasi dan penggantian baterai, kelengkapan unit saja | Nanda
05000356 | 2025-07-31 | Signature Dental Care | Dental Unit | KAVO/Primus 1059 | 10207191/2019-09 | Survey | Setelah dilakukan pengecekan, mainboard rusak | Pending | Alat belum bisa digunakan, diperlukan penggantian mainboard | Farhat,Ipan
12000221 | 2025-07-31 | RSUP Prof. Dr. R. D. Kandou Manado | Infant Warmer | Draeger/Resuscitaire | CY03657 | Service | Pegantian Heater Assy | Ready to use |  | Marshall
16000294 | 2025-07-31 | RS Dr. Suyoto | Dental Unit | Kavo/Primus 1058 | 7021-01-10707193 | Survey | Pengecekan selang air pada dental unit dan tidak ada masalah pada dental unit | Ready to use |  | Ipan, Farhat
06000909 | 2025-07-31 | RSUD Ajibarang | USG | Mindray/PC 30 | 9P-18009329 | Maintenance | Running test, Recharge baterai | Ready to use |  | Dwi
20000383 | 2025-07-30 | RSUP Makassar | Lampu Operasi | SYAS | 0L9680/80XVZW00144 | Service | Pengecekan arm | Ready to use | Unit baik setelah dilakukan perbaikan | Sujana
08000708 | 2025-07-30 | RSUD RA Kartini Jepara | Mesin Anasthesia | Draeger/Primus | ASMA-0193 | Service | Pemasangan service set 3y, System check, running test | Ready to use |  | Amna
08000708 | 2025-07-30 | RSUD RA Kartini Jepara | Mesin Anasthesia | Draeger/Primus | ASHM-0012 | Service | Pemasangan service set 3y, System check, running test | Ready to use |  | Amna
14000356 | 2025-07-30 | RSUP Vertikal Surabaya | Meja Operasi | SYAS/EPIONE/OP850 | OP850XV2Y0003 | Service | Cek remote control | Ready to use |  | Fahim
07000738 | 2025-07-30 | RSUD Bagas Waras | Mesin Anasthesia | Draeger/Perseus A500 | ASJD-0067 | Maintenance | Replace Baterai Pack dan Cleaning Bust | Ready to use |  | Wahyu
07000738 | 2025-07-30 | RSUD Bagas Waras | Mesin Anasthesia | Draeger/Perseus A500 | ASJD-0067 | Service | Replace DLC Pump | Ready to use |  | Wahyu
06000908 | 2025-07-30 | RSV Duta Mulya | Mesin Anasthesia | Draeger/Atlan A300 XL | ASRD-0072 | Survey | cek breath system, cek valve membrane, cek baterai, cek SC Software, cek Flow Sensor, cek water blok, sample line | Ready to use | Rekomendasi service set, breathsys 2y, flow sensor, water lock, dan sample line | Dwi
03000393 | 2025-07-30 | RSUD Ulin Banjarmasin | Mamografi | Metaltronica/Heliantus BYM | HELI 10048/CB1 | Survey | Kalibrasi Detector | Pending | Disarankan reload software dan kalibrasi ulang detector untuk memastikan kerusakannya | Imam
20000376 | 2025-07-29 | RSUP Tadjudin Chalid | ESWL | Medispec EM1000 | 0922014 | Service | Pengecekan kabel monitor | Ready to use | Unit berfungsi dengan baik | Sujana
09000597 | 2025-07-29 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF 0100 | Instalasi |  | Ready to use | Unit berfungsi dengan baik | Ruslan
09000597 | 2025-07-29 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 300 | VSSTF 0099 | Instalasi |  | Ready to use | Unit berfungsi dengan baik | Ruslan
09000596 | 2025-07-29 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 ID | ASUA-3009 | Instalasi | Install dan uji fungsi | Ready to use |  | Ruslan
09000596 | 2025-07-29 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 ID | ASUA-3008 | Instalasi | Install dan uji fungsi | Ready to use |  | Ruslan
09000596 | 2025-07-29 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 ID | ASUA-3007 | Instalasi | Install dan uji fungsi | Ready to use |  | Ruslan
09000596 | 2025-07-29 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 ID | ASUA-3006 | Instalasi | Install dan uji fungsi | Ready to use |  | Ruslan
09000595 | 2025-07-29 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 120 | VGSTJ0131 | Instalasi | Install dan uji fungsi patient monitor | Ready to use |  | Ruslan
09000595 | 2025-07-29 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 120 | VGSTJ0128 | Instalasi | Install dan uji fungsi patient monitor | Ready to use |  | Ruslan
09000595 | 2025-07-29 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 120 | VGSTJ0125 | Instalasi | Install dan uji fungsi patient monitor | Ready to use |  | Ruslan
09000595 | 2025-07-29 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 120 | VGSTJ0122 | Instalasi | Install dan uji fungsi patient monitor | Ready to use |  | Ruslan
09000595 | 2025-07-29 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 120 | VGSTJ0134 | Instalasi | Install dan uji fungsi patient monitor | Ready to use |  | Ruslan
19000431 | 2025-07-29 | RSUD Budi Rahayu Magelang | Ventilator | Draeger/Evita V600 | ASTD-0042 | Maintenance | mantenance unit warranty ventilator | Ready to use |  | Nanda
15000306 | 2025-07-29 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Flouroscopy | Apelem/Platinum | L156F0142 | Service | Sesuaikan tube dan detektor ok, setting keyboard, PC Komputer ok | Ready to use |  | Sutan
08000707 | 2025-07-29 | RSUD Kab. Temanggung | C-Arm | Apellem/EVO R | 551-15-011-469 (12R008) | Survey | Cek unit, check eksposer time, check eksposer counter, check tube | Pending | rekomendasi adjusment kv. ma dan kolimator, ukes fluoroskopi c-arm, install software, penambahan hardisk dan ram | Amna
08000706 | 2025-07-29 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHA-0132 | Survey | Device on, system check, battery check, battery checkin progress 48 hours | Pending | the device is charge and battery check for up to 48 hours | Amna
03000392 | 2025-07-29 | RSUD Ratu Zalecha | Digital Radiography Flouroscopy | Apelem/Platinum DRF | D186F0271 | Service | Cek kapasitas disk system, cek kapasitas disk data, backup database dan reload software duet versi 1.13 | Ready to use |  | Imam
14000355 | 2025-07-29 | RSUP Surabaya | Scrub OK | Lokal |  | Service | Ganti pompa motor | Ready to use |  | Fahim
07000737 | 2025-07-29 | RSUP Dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASHA-0109 | Service | Cek blok gas mixer, cek sambungan selang, selang aux, flow meter bocor, ganti selang, dan adjustment flow control | Ready to use |  | Wahyu
18000641 | 2025-07-28 | RSUD Ainun Habibi Gorontalo | Bor Ortopedi | Bbraun/Aculan 4 | 3075 | Instalasi | Instalasi unit | Ready to use | SN baterai;9553 dan 8365, Salah satu baterai tidak berfungsi (rusak) yang SN nya;8365 | Irfan
07000754 | 2025-07-28 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASJH-0125 | Maintenance | Replace Particle filter, replace kit air filter, cleaning dust | Ready to use |  | Wahyu
07000754 | 2025-07-28 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASJH-0125 | Service | Replace valve drive, replace pressure regulator, replace exp vlve, replace kabel harnes flow sensor | Ready to use |  | Wahyu
12000220 | 2025-07-28 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | ASNL0152 | Service | Pergantian Pressure Regulator | Ready to use |  | Marshall
18000632 | 2025-07-28 | RSUD Otanaha Gorontalo | Ventilator | Draeger/Eviita V800 | ASPL-0095 | Survey | Setelah di cek kerusakan pada pressure regulator | Ready to use | Unit masih berfungsi hanya ada bunyi di unit, disarankan penggantian pressure regulator | Irfan
06000907 | 2025-07-28 | RS Pertamina Cilacap | Ventilator | Draeger/Savina 300 | ASNC-00725 | Survey | Cek Unit, test running, self test, dan cek leaking | Ready to use |  | Dwi
06000907 | 2025-07-28 | RS Pertamina Cilacap | Ventilator | Draeger/Savina 300 | ASNL-0723 | Survey | Cek Unit, test running, self test, dan cek leaking | Ready to use |  | Dwi
06000907 | 2025-07-28 | RS Pertamina Cilacap | Ventilator | Draeger Savina 300 | ASNL-01328 | Survey | Cek Unit, test running, self test, dan cek leaking | Ready to use |  | Dwi
06000907 | 2025-07-28 | RS Pertamina Cilacap | Ventilator | Draeger Savina 300 | ASNC-0723 | Survey | Cek Unit, test running, self test, dan cek leaking | Ready to use |  | Dwi
11000428 | 2025-07-28 | RS Mata Makassar | Mesin Anasthesia | Draeger/Atlan A350 | ASPD0142 | Survey | Pendampingan pasien | Ready to use |  | Nasrul
18000636 | 2025-07-27 | RSUD Ainun Habibi Gorontalo | Bor Tulang | Draeger/Acculan 4 | 7166 | Instalasi | Unit berfungsi dengan baik dan ada salah satu baterai yang rusak | Ready to use | baterai akan dikirim ke kantor Bbraun untuk di cek | Irfan
18000636 | 2025-07-27 | RSUD Ainun Habibi Gorontalo | Bor Tulang | Draeger/Acculan 4 | 3075 | Instalasi | Unit berfungsi dengan baik dan ada salah satu baterai yang rusak | Ready to use | baterai akan dikirim ke kantor Bbraun untuk di cek | Irfan
20000382 | 2025-07-25 | RSV Makassar | MOT | MOT SYAS |  | Instalasi |  |  | Pra instalasi unit pendant dan cabinet MOT | Sujana, Nasrul
08000705 | 2025-07-25 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASHL-0063 | Service | Pemasangan service set 3y | Ready to use |  | Amna
19000437 | 2025-07-24 | RS Telogorejo | Mesin Anasthesia | Draeger/Primus | ASDK-0223 | Survey | Cek alat, bongkar bagian mixer, tidak ditemukan ada air masuk ke dalam mixer | Pending | Unit ada indikasi kerusakan pada valve tank | Nanda
15000305 | 2025-07-24 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Flouroscopy | Apelem/Platinum | L156F0142 | Service | Cek sensor dekat detektor (bersihkan) ok, cek pu box ok | Ready to use |  | Sutan
06000906 | 2025-07-24 | RSUD Samin Surosentiko Randublatung | USG | Mindray Consona N8 | LP7U56000375T | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
03000391 | 2025-07-24 | RSUD Dr. Djasemen Saragih | Mamografi | Metaltronica/Heliantus DBT | HELIDBT/137/C2 | Survey | Cek V-RAM master LCD dan Input tegangan, master LCD | Pending | V-RAM error not booting | Imam
05000355 | 2025-07-24 | RSUP Dr. Hasan Sadikin | Acculan 4 Charging Unit | Bbraun/GA320 | 7142 | Maintenance |  | Ready to use |  | Farhat, Satria
05000354 | 2025-07-24 | RSUP Dr. Hasan Sadikin | Acculan 4 | Bbraun/GA346 | 9403/9552 | Maintenance |  | Ready to use |  | Farhat, Satria
05000353 | 2025-07-24 | RSUP Dr. Hasan Sadikin | Acculan 4 | Bbraun/GB894R | 3241 | Maintenance |  | Ready to use |  | Farhat, Satria
05000353 | 2025-07-24 | RSUP Dr. Hasan Sadikin | Acculan 4 | Bbraun/GA344 | 3044 | Maintenance |  | Ready to use |  | Farhat, Satria
05000353 | 2025-07-24 | RSUP Dr. Hasan Sadikin | Acculan 4 | Bbraun/GB884R | 2946 | Maintenance |  | Ready to use |  | Farhat, Satria
05000353 | 2025-07-24 | RSUP Dr. Hasan Sadikin | Acculan 4 | Bbraun/GB891R | 2910 | Maintenance |  | Ready to use |  | Farhat, Satria
14000354 | 2025-07-24 | RSUP Surabaya | HVAC (OK 2) | Lokal |  | Service | Cek ruangan dan cek Panel PLC | Pending | Tidak dapat di setting ulang dan di reset, membutuhkan kalibrasi PLC dari pabrik | Fahim
09000704 | 2025-07-23 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Evita V800 | ASNK-0340 |  | Uji fungsi ventilator | Ready to use | Demo unit | Ruslan
09000704 | 2025-07-23 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 | ASSM-0191 |  | Uji fungsi ventilator | Ready to use | Demo unit | Ruslan
09000704 | 2025-07-23 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Savina 300 | ASTJ-2004 |  | Uji fungsi ventilator | Ready to use | Demo unit | Ruslan
08000703 | 2025-07-23 | RSUD Batang | Seattle PAP | Draeger/Oxymixer | QM20230134 | Instalasi | Training demo unit |  |  | Amna
19000430 | 2025-07-23 | RS Bhakti Wira Tamtama | CPAP | Draeger/Seattle CPAP | QM20255073 | Instalasi | Unit sudah berfungsi dengan baik | Ready to use | Unit sudah di install dengan baik dan sudah uji fungsi dengan baik | Nanda
14000353 | 2025-07-23 | RSUP Surabaya | Scrub | Lokal |  | Service | Cek sambungan air dan cek pompa motor | Pending | Pompa motor tidak dapat digunakan, menunggu part dari pabrik | Fahim
18000635 | 2025-07-23 | RSUD Tani Nelayan | Ventilator | Evita V800/Draeger | ASPH-0001 | Survey | Mencoba part yang dipinjamkan dari draeger tapi unit tetap masih error | Pending | Tim draeger akan membaca hasil error lagi dengan tarik eror | Irfan
19000430 | 2025-07-23 | RS Bhakti Wira Tamtama | CPAP | Draeger/Seattle CPAP | QM20255076 | Instalasi | Instalation | Ready to use | Unit sudah di install dengan baik dan sudah uji fungsi dengan baik | Nanda
11000427 | 2025-07-23 | RS Mata Makassar | Mesin Anasthesia | Draeger/Atlan A350 | ASPD0142 | Survey | Pendampingan unit ke pasien | Ready to use |  | Nasrul
20000381 | 2025-07-22 | RSV Makassar | MOT | MOT SYAS |  | Instalasi | Pra instalasi pendant | Ready to use |  | Sujana
19000429 | 2025-07-22 | RSUP Dr Kariadi Semarang | Dental Unit | KAVO/1058TM | 2004-1002719 | Service | Cek unit, cek alat, ada kebocoran pada valve vacuum air, retakan di lem, ok | Ready to use | Bagian 3 way bocor, disarankan ganti 3 way | Nanda
08000702 | 2025-07-22 | RS Budi Agung Juwana | Mesin Anasthesia | Draeger/Fabius Plus | ASTK-0003 | Maintenance | General check, system check, check flow sensor, check leakage and complain test, running test | Ready to use |  | Amna
14000352 | 2025-07-22 | RSUP Surabaya | HVAC | Lokal |  | Service | Cek ruangan dan cek Panel PLC | ready to use | Humidity pada panel PLC di setting ulang | Fahim
07000735 | 2025-07-22 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASPF-0047 | Survey | System test, cek error log, PATO Faully, pressure regulator faulty, replace o2 sensor dan replace pressure regulator | Pending | Unit perlu penggantian o2 sensor dan pressure regulator | Wahyu
03000390 | 2025-07-22 | RSUP Prof. Dr. R. D. Kandou Manado (KSO) | CT Scan | Philips/Ingenuity CT | 320133 | Survey | Cek disk data (tidak bisa masuk ke windows explorer karena sistem di password) | Ready to use |  | Imam
18000643 | 2025-07-22 | RSUD Zainal Umar Sidiki | Mesin Anasthesia | Draeger/Perseus A500 | ASPH-0123 | Service | Melakukan error log dan device check | Ready to use | Disarankan untuk melakukan kontrak service, agar unit terjadwal dalam pemeliharaan | Irfan
20000372 | 2025-07-21 | RSUP Tadjudin Chalid | Digital Radiografi Fluoroskopi | Apelem/Platinum | 1236F0882 | Service | Melakukan penambalan timbal dan penggantian | Ready to use |  | Sujana
19000428 | 2025-07-21 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Primus/Draeger | ASKJ-0036 | Survey | Cek unit, cek koneksi, ada kerusakan pada layar LCD | Ready to use | Disarankan ganti layar LCD | Nanda
19000427 | 2025-07-21 | RSUD Dr. Soediran Mangun Sumarso | Pasien Monitor Central | Drager/Vista 120 CMS | 35297369247 | Service | Cek unit, ganti router pinjaman dari IPSRS, running on | Ready to use | Unit sementara di backup dengan router milik RS, Disarankan ganti router | Nanda
08000701 | 2025-07-21 | RS Mardi Rahayu Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASKL-0012 | Service | Install O2 sensor capsule, system check, running test | Ready to use |  | Amna
08000701 | 2025-07-21 | RS Mardi Rahayu Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASKL-0013 | Service | Install O2 sensor capsule, system check, running test | Ready to use |  | Amna
06000905 | 2025-07-21 | RSUD Ajibarang | USG | Mindray/DC-30 | 9P18009329 | Service | Replace baterai kit DC-30 | Ready to use |  | Dwi
20000374 | 2025-07-20 | RS Mata Kemenkes Makassar | Mesin Anasthesia | Draeger/Atlan A350 XL | ASPD-0142 | Instalasi | Unit demo |  | Unit ready to use (demo unit) | Sujana, Nasrul
14000351 | 2025-07-20 | RSUP Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0146 | Service | Cek kipas bunyi gantian pada bagian depan | Ready to use | Mengendorkan baut kipas agar tidak berbunyi | Fahim
11000426 | 2025-07-18 | RSUP Makassar | Ventilator | Draeger/Evita V300 | ASPH0033 | Service | Batterai failure; Baterai rusak | Pending | Disarankan penggantian baterai | Nasrul
20000373 | 2025-07-18 | RSUP Makassar | Phototherapy | Draeger/Billilux | 1126320019 | Instalasi | Install Unit | Ready to use | Unit dapat digunakan | Sujana, Nasrul
09000703 | 2025-07-18 | RSUD Labuha | X-Ray Digital | Apelem/Camargue DR | CPD36860K20 | Survey | Survey pemindahan Alat X-Ray Digital Camargue DR | Ready to use | Segera dibuatkan penawaran pemindahanunit dari gedung lama ke gedung baru | Ruslan
19000426 | 2025-07-18 | RS Telogorejo Semarang | Mesin Anasthesia | Draeger/Primus | ASDK-0223 | Service | Tarik error, cal. mixer, cal. VGC, Running | Ready to use |  | Nanda
19000425 | 2025-07-18 | RS Nasional Diponegoro | Incubator Transport | Draeger/Ti500 | AF10588 | Survey | Cek unit, cek tegangan baterai sudah drop | Pending | Disarankan Penggantian Baterai | Nanda
19000424 | 2025-07-18 | RS Nasional Diponegoro | Pasien Monitor | Draeger/Vista 120 | VGSNE5094 | Service | Cek alat, exchange dengan unit lain, ok | Ready to use | Unit sudah diganti part dengan NIBP yang lain | Nanda
08000700 | 2025-07-18 | RS Sebening Kasih | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ-0082 | Maintenance | Check, check O2 sensor, check flow sensor, running test | Ready to use | Rekomendasi ganti O2 sensor capsule | Amna
06000904 | 2025-07-18 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEH-0048 | Survey | Calib Flow Sensor, calib o2, test leak system | Ready to use | Rekomendasi penggantian service set 3y, cozy breathing, flow sensor o2 cell | Dwi
03000389 | 2025-07-18 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Flouroscopy | Apelem/Platinum | L156F0142 | Service | Kalibrasi Ulang meja | Ready to use |  | Imam
02000202 | 2025-07-17 | RSUD Nyi Ageng Serang Kulon Progo | UPS | Yoshiga/HEXA 33-160 KL | 110025809830204300002 | Instalasi | Pemindahan ruang unit UPS dan install ulang UPS | Ready to use |  | Eddy, Wahyu
20000350 | 2025-07-17 | RSUP Makassar | MOT | SYAS |  | Service | Pengecekan unit, cleaning door dan restart power | Ready to use | Unit dapat digunakan kembali | Sujana
19000423 | 2025-07-17 | RSUD Dr. Moewardi Surakarta | USG | Mindray/MG | CCI-04002265 | Survey | Cek unit, exchange keyboard mini, ok | Pending | Control panel sudah diganti tapi ada tambahan, Kerusakan pada keyboard mini | Nanda
14000357 | 2025-07-17 | RSUD Dr. Soetomo Surabaya | ESMR | Medispec/CS-2 | 03-10297 | Survey | Cek power supply ESMR, cek power supply USG, cek EKG, cek fungsi shock wave, cek water sirkulasi system | Pending | Perlu diganti System water sirkulasi, aplication head ESMR | Fahim
06000903 | 2025-07-17 | RSUD Ajibarang | CPAP | Draeger/Oxymixer | QM20220254 | Sevice | Penggantian oxymixer service kit, dan service certificate | Ready to use |  | Dwi
05000352 | 2025-07-17 | Signature Dental Care | Dental Unit | - | - | Service | Dilakukan penambahan HDMI spliter pada ruang OK, lantai 3 dan antara proyektor | Ready to use |  | Farhat, Satria, Ivan
09000600 | 2025-07-17 | RSD Tidore | Mesin Anasthesia With BIS | Draeger/Atlan A300 | ASUF-0382 | Instalasi | Install, Uji fungsi dan training | ready to use | Unit berfungsi dengan baik | Ruslan
07000739 | 2025-07-16 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Atlan A300 | ASTK-0061 | Service | Cek flowsensor, cek ventilator piston, cek sensor board, dan reload software | Ready to use |  | Wahyu
20000370 | 2025-07-16 | RSUP Makassar | MOT | SYAS | HVAC OK 2 Gedung C | Service | Pengecekan unit indor dan outdor ( Restart unit) setting suhu | Pending | Perlu pengacakan lebih lanjut HVAC adanya perbedaan pembacaan suhu dan RIT anatar di monitor dan panel start up | Sujana
20000361 | 2025-07-16 | RSUP Makassar | MOT HVAC | SYAS |  | Survey | Cek setting suhu MOT | Pending |  | Sujana
18000634 | 2025-07-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN 500 | ASEK0110 | Maintenance | pergantian part presure regulator | Ready to use |  | Irfan
09000593 | 2025-07-16 | RSD Tidore | Ventilator | Draeger/Babylog VN800 | ASTN-0001 | Instalasi | Install, uji fungsi, dan di training, unit ventilator babylog VN800 di RSUDKota Tidore | Ready to use |  | Ruslan
11000425 | 2025-07-16 | RSUP Dr. Wahidin Sudirohusodo | Syringe Pump | Bbraun/Perfusor Space | 725745 | Survey | Baterai tidak bisa di isi ulang; Baterai rusak | Pending |  | Nasrul
11000425 | 2025-07-16 | RSUP Dr. Wahidin Sudirohusodo | Syringe Pump | Bbraun/Perfusor Space | 716667 | Survey | Baterai tidak bisa di isi ulang; Baterai rusak | Pending |  | Nasrul
20000371 | 2025-07-15 | RSUP Makassar | MOT | Syas | OK 2 | Service | Setting suhu | Ready to use | Unit baik setelah di setting suhu | Sujana
19000436 | 2025-07-15 | RSUD Tugurejo Semarang | Mesin Anasthesia | Draeger/Primus | ASHJ-0180 | Survey | Cek unit, kalibrasi device check ulang, ada kebocoran pada selang pernafasan | Ready to use | Sudah diganti Selang pernapasan | Nanda
15000304 | 2025-07-15 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Cek travel switch ok, kalibrasi lifter ok,kalibrasi tilitng ok | Ready to use |  | Sutan
08000699 | 2025-07-15 | RSI Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASLK-0114 | Service | Install seal, gasket, diaphgram membran, ventillteler, O2 sensor capsule, system check, running test | Ready to use |  | Amna
12000218 | 2025-07-15 | RSUD Ainun Habibie | Meja Operasi | SYAS Tech | 250552511140400027 | Instalasi | Install, uji fungsi dan training user | Ready to use |  | Marshall
07000734 | 2025-07-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE-0234 | Survey | Replace Conv. kit display chimei, post test | Ready to use | Unit perlu penggantian conv. kit display chimei | Wahyu
04000509 | 2025-07-15 | RSUD Malingping | Mamografi | Metaltronica/Heliantus DBT | HELIDBT/131/C2 | Service | Adjust mirror, collimator, kompresi, dan log check | Pending | Masih ada kendala tegangan dan suhu ruangan | Satria, Imam, Bahar
18000633 | 2025-07-15 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina 300 | ASTK-0074 | Survey | Setelah dicek unit tidak bermasalah unit hanya pemakaian terlalu lama dan butuh kalibrasi manual | Ready to use |  | Irfan
12000217 | 2025-07-15 | RSUD Ainun Habibie | MOT | SYAS Thec |  | Instalasi | Install, uji fungsi dan training user | Ready to use |  | Marshall
19000421 | 2025-07-14 | RSUD Dr. Soehadi Prijonegoro | Ventilator | Draeger/Babylog VN800 | ASTH-0035 | Maintenance | Maintenance unit warranty. pre use check,cleaning, running, ok | Ready to use |  | Nanda
12000219 | 2025-07-14 | RSUD Ainun Habibie | MOT | SYAS Tech |  | Service | Penggantian gas medical air dan vacum | Ready to use | Sudah diganti dengan yang baru | Marshall
07000733 | 2025-07-14 | RSUD Bagas Waras | Pasien Monitor | Draeger/Vista 120 | VGSNF1838 | Survey | test on, cek error, cek software, cek setting modul, reload software, dan post test | Ready to use |  | Wahyu
11000424 | 2025-07-14 | RSUP Tadjudin Chalid | Mesin Anasthesia | Draeger/Primus | ASJF0037 | Service | Sensor check fail, ventilator fail; Penggantian port distrubuting valve dan electronic valve | Ready to use |  | Nasrul
14000360 | 2025-07-13 | RSUP Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0146 | Servive | Membongkar dan cek bagian tube | Ready to use | mengendorkan baut agar kipas dapat berputar | Fahim
20000380 | 2025-07-12 | RSV Makassar | MOT | SYAS MOT |  |  | Standby operasi perdana jantung gedung D | Ready to use | Unit dapat digunakan | Sujana
19000420 | 2025-07-11 | RSUP Dr Kariadi Semarang | Infant Warmer Resusitasi | Draeger/Resuscitaire | AF10367 | Service | Cek unit, ganti seal di valve control mini 5 way, running, ok | Ready to use |  | Nanda
15000303 | 2025-07-11 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Flouroscopy | Apelem/Platinum | L156F0142 | Service | Cek travel switch ok, bersihkan CAR RAM ok | Ready to use |  | Sutan
08000698 | 2025-07-11 | RSUD Dr. Loekmono Hadi Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASKE-0198 | Survey | Check cozy, check piston, check absorber, check breathing circuit, system check, running test | Ready to use |  | Amna
06000902 | 2025-07-11 | RSUI Harapan Anda Tegal | Ventilator | Draeger VN500 | ASJC-0050 | Survey | Device Check, Baterai test running | Pending | Rekomendasi Service set ventilator indue baterai | Dwi
06000902 | 2025-07-11 | RSUI Harapan Anda Tegal | Ventilator | Draeger VN500 | ASMK-0055 | Survey | Device Check, Baterai test running | Pending | Rekomendasi Service set ventilator indue baterai | Dwi
06000902 | 2025-07-11 | RSUI Harapan Anda Tegal | Ventilator | Draeger VN600 | ASTE-0058 | Survey | Device Check, Baterai test running | Pending | Rekomendasi Service set ventilator indue baterai | Dwi
06000902 | 2025-07-11 | RSUI Harapan Anda Tegal | Ventilator | Draeger VN600 | ASTE-0057 | Survey | Device Check, Baterai test running | Pending | Rekomendasi Service set ventilator indue baterai | Dwi
06000902 | 2025-07-11 | RSUI Harapan Anda Tegal | Ventilator | Draeger VN800 | ASRL-0029 | Survey | Device Check, Baterai test running | Pending | Rekomendasi Service set ventilator indue baterai | Dwi
03000388 | 2025-07-11 | RS Immanuel | Digital Radiography Flouroscopy | Apelem/Platinum | H226F0F45 | Service | Kalibrasi Table ulang | Ready to use |  | Imam
18000637 | 2025-07-11 | RS Mata Manado | Mesin Anasthesia | Fabius/Draeger | ASNE-0025 | Survey | Cek sensor o2 | Pending | Disarankan untuk mengganti o2 sensor | Irfan
18000639 | 2025-07-11 | RSUP Prof. Dr. R. D. Kandou Manado | Bor Tulang | BBraun/ELAN4 | 003477 |  | Penarikan unit backup bersama tim Bbraun | Ready to use |  | Irfan
20000375 | 2025-07-11 | RS Mata Kemenkes Makassar | Anasthesia | Draeger/Atlan A350 | ASNH-0125 |  | Penarikan Unit demo | Pending |  | Sujana, Nasrul
19000418 | 2025-07-10 | RSUD Dr Moewardi Surakarta | USG | Mindray/Mg | CCI-04002265 | Survey | Cek alat, cek keyboard mini | Pending | Ada indikasi kerusakan pada keyboard mini, akan dicarikan dulu untuk compare | Nanda
08000697 | 2025-07-10 | RSUD RAA Soewondo Pati | Ventilator | Draeger/Evita V300 | ASHA-0098 | Survey | System check, check hit software battery internal and external | Pending | Needs replacement internal battery and external battery | Amna
06000901 | 2025-07-10 | RSUI Harapan Anda Tegal | Baby Incubator | Draeger Isolatte | DSRJ-0061 | Survey | Cek temperatur sensor, compare skin temperatur, dan running test | Pending | Rekomendasi penggantian skin temperatur/rusak | Dwi
06000901 | 2025-07-10 | RSUI Harapan Anda Tegal | Baby Incubator | Draeger Isolatte | DSPL-0022 | Survey | Cek temperatur sensor, compare skin temperatur, dan running test | Pending | Rekomendasi penggantian skin temperatur/rusak | Dwi
06000901 | 2025-07-10 | RSUI Harapan Anda Tegal | Baby Incubator | Draeger Isolatte | DSRJ-0067 | Survey | Cek temperatur sensor, compare skin temperatur, dan running test | Pending | Rekomendasi penggantian skin temperatur/rusak | Dwi
14000359 | 2025-07-10 | RSUD Dr. Moh. Saleh | UPS | Yoshiga/GPH 33-120 | 501212020260 P3800003 | Maintenance | Cek kondisi baterai ,cek tegangan, cek fisik | Ready to use |  | Fahim
20000349 | 2025-07-09 | RSUP Tadjudin Chalid | Digital Radiografi Fluoroskopi | Apelem/Platinum | J236F0882 | Service | Pengukuran uji paparan pada kaca tambal di ruangan operator | Ready to use | Unit baik, perlu perbaikan pada kaca PB dan area operator | Sujana
19000417 | 2025-07-09 | RSUD Dr. Soediran Mangun Sumarso | Central Monitor | Draeger/Vista 120CMS | 35297369247 | Survey | Cek alat, set up connection wifi | Pending | Masih belum bisa kemungkinan kerusakan pada router,akan di coba dengan router lain | Nanda
03000387 | 2025-07-09 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Flouroscopy | Apelem/Platinum | L156F0142 | Service | Pasang PC Console dan Identifikasi error meja | Ready to use |  | Imam
07000750 | 2025-07-09 | RSUD Purworejo | Jaundice Meter | Draeger/JM105 | H3500310 | Instalasi | Installl, Uji fungsi, dan training unit demo jaundice meter JM-105 | Ready to use |  | Wahyu
19000422 | 2025-07-08 | RSUD Dr Adhyatma MPH Semarang | Incubator Transport | Draeger/TI500 | MR11799 | Survey | Cek alat, cek baterai, drop, sensor temp. tidak ada | Pending | Unit baterai drop dan sensor temp hilang, disarankan ganti baterai dan sensor temperatur | Nanda
11000423 | 2025-07-09 | RSUD Latemmamala | Ventilator | Draeger/Babylog VN500 | ASML0095 | Survey | Kalibrasi gas supply fail, air supply fail, O2 sensor fail; Pressure regulator rusak, kit maintenance valve | Pending |  | Nasrul
20000379 | 2025-07-08 | RSV Makassar | MOT | SYAS MOT |  | Instalasi | Pengecekan ruangan MOT untuk digunakan operasi perdana | Ready to use | Unit siap digunakan | Sujana
11000421 | 2025-07-08 | RSUP Makassar | Mesin Anasthesia | Draeger/Primus | ASRH0049 | Service | O2 sensor fail, gas sensor fail; Flow power rusak | Pending |  | Nasrul
19000414 | 2025-07-07 | RS Syubbannul Waton Magelang | Incubator Transport | Draeger/TI500 | DH12427 | Service | Cek alat, charger di luar, running, ok | Ready to use |  | Nanda
08000696 | 2025-07-07 | RSUD RA Kartini Jepara | Ventilator | Draeger/Babylog VN 500 | ASMK-0029 | Survey | System check, remote device check | Pending | Compressor not pressure | Amna
18000638 | 2025-07-07 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASHD0032 | Survey | Tarik error | Pending | Hasil tarik eror masih menunggu konfirmasi dari tim draeger karena mereka yang membaca hasilnya | Irfan
011000419 | 2025-07-07 | RSUP Wahidin Sudirohusodo | Infant Warmer | Draeger/Babyroo TN300 | DSTJ-0113 | Survey | Refresh training user kamar operasi | Ready to use |  | Nasrul
20000378 | 2025-07-07 | RSV Makassar | MOT | SYAS MOT |  | Instalasi | Pra instalasi pendant | Ready to use |  | Sujana
06000850 | 2025-07-05 | RSU An Ni'mah | Ventilator | Draeger/Savina 300 | ASPJ-0377 | Service | Oxytrace VE 2 pcs,Savina service set 2y replacemen, kalibrasi o2 and cek reset baterai | Ready to use |  | Dwi
06000849 | 2025-07-04 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASKJ-0050 | Survey | Cek service connect, test compare valve distributing MV2 | Pending | Kerusakan pada valve distributing MV2 akan dibuatkan SPH | Dwi
17000217 | 2025-07-04 | RSUP Jayapura | Digital Radiography Flouroscopy | Apelem/Platinum Neo | H248L0157 | Service | Replace detector dan kalibrasi detector | Ready to use |  | Bahar, Eddy
08000695 | 2025-07-04 | RS Mitra Bangsa Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASNI-0041 | Survey | General check, system check, check flow sensor, check leakage and complain test, running test | Ready to use |  | Amna
19000416 | 2025-07-03 | RSUD Dr Moewardi Surakarta | USG | Mindray/Mg | CCI-04002265 | Service | Cek unit, control panel assy yang baru sudah terpasang, upgrade software ke latest version os 3.3.14 dan droppler 4.9.0 tapi masih belum bisa | Pending | Unit masih belum bisa, sudah laporan ke tim mindray | Nanda
04000508 | 2025-07-03 | RSUP H. Adam Malik Medan | Mamografi | Metaltronica/Heliantus DBT | HELIDBT/083/C1 | Maintenance | Refresh training teknisi terkait error dan trouble, setting kolimasi mamo dan pendampingan ukes mamography | Ready to use |  | Satria, Imam
20000369 | 2025-07-03 | RSUP Makassar | C-Arm | Apelem/EVO RT | 51240241086 | Survey | Dilakukan refresh training digantikan radiologi rumah sakit | Ready to use |  | Sujana, Nasrul
08000694 | 2025-07-03 | RSUD Batang | Seatle PAP | Draeger/Oxymixer | QM20230134 | Instalasi | Install Demo Unit |  |  | Amna
19000413 | 2025-07-02 | RSUP Dr Kariadi | Ventilator | Savina 300/Draeger | ASMK-0305 | Service | Cek alat, coba charger baterai di luar, pasang, ok | Ready to use |  | Nanda
12000216 | 2025-07-02 | RSUD Zainal Umar Sidiki | Mesin Anasthesia | Draeger/Perseus A500 | ASPH-0123 | Survey | Identifikasi setelah tarik error terjadi permasalahan pada sensor PCA dan sensor pressure baging | Pending | Disarankan untuk mengganti part, untuk saat ini lagi menunggu part backup dari draeger untuk di uji coba | Marshall
07000732 | 2025-07-02 | RS Pratama | Mesin Anasthesia | Draeger/Fabius Plus | ASHL-0048 | Survey | Cek kabel flow sensor, cek flow sensor, replace flow sensor, cek kabel housing o2, cek o2 sensor, dan post test | Ready to use | Unit disarankan penggantian o2 sensor dan o2 sensor housing, service set 3y | Wahyu, Wening
06000847 | 2025-07-02 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM-0072 | Survey | Compare MV2cek, Cek Post test service mode | Pending | Leak in Sytem VGC, Cek lebih lanjut | Dwi
06000846 | 2025-07-02 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASTF-0005 | Survey | Cek HIT Software SC, Cek Status O2 PIA2 | Pending | PIA 2 Problem Claim Warrranty | Dwi
20000368 | 2025-07-02 | RSU Faisal Makassar | ESWL | Medispec/EM 1000 | 1114075 | Service | Penggantian selang sirkulasi air | Pending | Perlu penggantian EMH baru | Sujana
02000196 | 2025-07-02 | RSI Faisal | ESWL | Medispec/EM 1000 | 1114075 | Service | Penggantian selang sirkulasi air | Ready to use | Perlu penggantian EMH yang sudah EXP | Eddy
15000302 | 2025-07-01 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiography Flouroscopy | Apelem/Platinum | L156F0142 | Service | Cek tegangan ke PC, tegangan ada 12 VDC, PC konsul dikirim ke mas Imam (pusat) | Pending | PC Konsul tidak hidup | Sutan
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400293 | Instalasi |  | Ready to use |  | Amna
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400287 | Instalasi |  | Ready to use |  | Amna
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400286 | Instalasi |  | Ready to use |  | Amna
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400290 | Instalasi |  | Ready to use |  | Amna
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400284 | Instalasi |  | Ready to use |  | Amna
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400294 | Instalasi |  | Ready to use |  | Amna
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400282 | Instalasi |  | Ready to use |  | Amna
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400291 | Instalasi |  | Ready to use |  | Amna
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400289 | Instalasi |  | Ready to use |  | Amna
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400283 | Instalasi |  | Ready to use |  | Amna
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400285 | Instalasi |  | Ready to use |  | Amna
18000640 | 2025-07-01 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger /TN 500 | ASNF0004 | Survey | Melakukan tarik error dan cek heater | Pending | Heater rusak | Irfan
07000731 | 2025-07-01 | RSUD Panembahan Senopati | Ventilator | Draeger/Evita V300 | ASJH-0121 | Maintenance | Replace kit air filter | Ready to use |  | Wahyu
20000367 | 2025-07-01 | RSUP Makassar | C-Arm | Apelem/Evo R+ | 551-24-024-1088 | Service | Cek connection IP address software | Ready to use |  | Sujana
02000195 | 2025-07-01 | RSUV Makassar | C-Arm | Apelem/Evo R+ | 551-24-024-1086 | Service | Cek connection IP address | Ready to use |  | Eddy
06000845 | 2025-07-01 | RSU Muhammadiyah Siti Aminah Bumiayu | Mesin Anasthesia | Draeger/Fabius Plus | ASND-0060 | Survey | Cek cozy, breathing, sell test unit service mode | Ready to use |  | Dwi
08000693 | 2025-07-01 | RSUD RA Kartini Jepara | Foot Step | SYAS Tech SFS 1 | 250661301000400288 | Instalasi |  | Ready to use |  | Amna
19000419 | 2025-06-30 | RS JIH Surakarta | Mesin Anasthesia | Draeger/Primus | ASMM-0125 | Maintenance | Selftest, cleaning, running, ok | Ready to use | Maintenance unit service kontrak bulan juli 2025 | Nanda
19000415 | 2025-06-30 | RSUD Dr Moewardi | USG | Mindray/Mg | CCI-04002265 | Service | Cek unit, selftest, keyboard rusak, pasang control panel yang baru masih belum bisa | Pending | Unit masih belum bisa, akan di coba untuk upgrade software dulu | Nanda
20000366 | 2025-06-30 | RSU Makassar | MOT HVAC | SYAS |  | Survey | Settingan HVAC MOT | Pending | Perlu perbaikan HVAC | Sujana
06000844 | 2025-06-30 | RSUD Panti Nugroho | Ventilator | Draeger/Evita V300 | ASCD-0168 | Survey | Test unit software | Ready to use |  | Dwi
19000412 | 2025-06-30 | RS JIH Solo | Mesin Anasthesia | Draeger/Primus | ASMA-0146 | Maintenance | Maintenance unit garansi | Ready to use |  | Nanda
19000411 | 2025-06-30 | RSUD Dr. Moewardi | USG | Mindray/Mg | CCI-04002265 | Service | Cek alat, penggantian control panel | Pending | Dijadwalkan lagi untuk update software | Nanda
07000729 | 2025-06-30 | RSUD Merah Putih | Jaundice Meter | Draeger/JM105 | H3510377 | Survey | Cek kondisi docking, cek kabel USB, dan cek software JM-105 | Ready to use | Disarankan penggantian kabel USB | Wahyu
08000692 | 2025-06-30 | RSUD Dr. Loekmono Hadi Kudus | Mesin Anasthesia | Draeger/Perseus A500 | ASUC-0046 | Instalasi | Test Function | Ready to use |  | Amna
20000365 | 2025-06-27 | RSU Makassar | C-Arm | Apelem/Evo R+ | 551-24-024-1086 | Survey | Pengecekan unit camera dan surcei kerusakan | Pending | Perlu perbaikan lebih lanjut | Sujana
15000301 | 2025-06-26 | RSUP Dr. Mohammad Hoesin | X-ray Digital | Apelem/Camargue DR | IMA 704028 | Service | Cek suplay kolimator | Ready to use | Lampu kolimator tidak hidup dan sudah diperbaiki | Sutan
09000650 | 2025-06-26 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0192 | Survey | Tombol on/off, setelah di cek kerusakan pada switch on/off | Pending | Disarankan untuk mengganti switch on/off | Irfan
09000592 | 2025-06-26 | RSUD Sanana | Premium Bed 3 Crank |  | 250562404000400106 | Instalasi | Install premium bed 3 crank Jumlah 1 unit | Ready to use |  | Ruslan
09000591 | 2025-06-26 | RSUD Sanana | Pediatrick Bed 2 Crank |  | 250326302000400026 | Instalasi | Install paket pediatrick bed 2 crank, cabinet 76 dan overbed table sudah di distribusikan ke ruangan | Ready to use |  | Ruslan
09000589 | 2025-06-26 | RSUD Sanana | Premium strecher Hydraulic | Syastech | 250331214000400012 | Instalasi | Install uji fungsi dan di training ke user | Ready to use |  | Ruslan
09000588 | 2025-06-26 | RSUD Sanana | Premium Ginekologi table | Syastech | 250534919000410022 | Instalasi | Unit telah di install dan di training ke user | Ready to use |  | Ruslan
09000587 | 2025-06-26 | RSUD Sanana | Wheel Chair | Syastech | 250531313000400263 | Instalasi | Unit telah di install dan di distribusikan ke ruangan | Ready to use |  | Ruslan
07000730 | 2025-06-26 | RSUP dr. Soeradji | Ventilator | Draeger/Baby Log-VN800 | ASKL-0092 | Service | Replace Pressure Regulator 2 pcs | Ready to use |  | Wahyu
08000691 | 2025-06-26 | RSUD Dr. Loekmono Hadi Kudus | Mesin Anasthesia | Draeger/Perseus A500 | ASUC-0046 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
08000689 | 2025-06-25 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | AINH-0240 | Survey | Cek sistem, cek mpgm, cek minimodul pump | Pending | Disarankan penggantian minimodul pump dan board mixer | Amna
05000351 | 2025-06-25 | Signature Dental Care | Dental Unit | - | - | Survey | Pengecekan jalur HDMI dan LAN untuk koneksi DU lantai 2 ke ruang seminar lantai 3 | Pending | Jalur sudah ok tinggal menunggu klinik beli kebutuhan untuk koneksi | Farhat, Ivan
20000364 | 2025-06-25 | RSU Makassar | Lampu operasi MOT | SYAS |  | Survey | Setting arm | Ready to use |  | Sujana
07000728 | 2025-06-25 | RSPAU Dr. S. Hardjolukito | Ventilator | Draeger/Evita V600 | ASRL-0039 | Service | Replace O-ring | Ready to use |  | Wahyu
15000300 | 2025-06-24 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Ganti sensor grid ok, cek lifter ok, ganti travel switch ok | Ready to use |  | Sutan
09000649 | 2025-06-24 | RS Mata Manado | Mesin Anasthesia | Draeger/Perseus A500 | ASJM-0178 | Survey | Setelah di cek kerusakan pada baterai | Pending | Unit masih dalam garansi kontrak service total care, baterai baru akan dikirimoleh tim draeger untuk diganti | Irfan
07000727 | 2025-06-24 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Instalasi | Install dan post test | ready to use |  | Wahyu, Handoko
06000843 | 2025-06-24 | RSUD Prembun Kebumen | Incubator Transport | Draeger/TI 500 | HH 12657 | Service | Pergantian batery 12v | Ready to use |  | Dwi
19000410 | 2025-06-24 | RSUD Tidar Magelang | BMD | Medilink/Medix DR | 15016M258 | Service | Cek alat ,PSU tidak ada, diambilkan di PSU yang lama, running dan ok | Ready to use |  | Nanda, Syamsudin, Satria
09000648 | 2025-06-23 | RSUP Prof. Dr. R. D. Kandou Manado | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J190100103 | Survey | Setelah di cek unit tidak ada masalah hanya RS perlu order fiber optik | Pending | Karena stock fiber sudah habis | Irfan
09000586 | 2025-06-23 | RSUD Sanana | OT Elektrohydrolic | Syastech | 250367221000400007 | Instalasi | Unit telah di install uji fungsi dan di training | Ready to use |  | Ruslan
09000585 | 2025-06-23 | RSUD Sanana | Standar Ginekologi, table dan food step | Syastech | 250364112000400104 | Instalasi | Install sebanyak 3 unit sudah di bagikan ke ruangan dan food step sebanyak 6 unit | Ready to use |  | Ruslan
09000584 | 2025-06-23 | RSUD Sanana | Food Trolly | Syastech | 250435213000900014 | Instalasi | Install sebanyak 2 unit dan sudah dibagikan ke ruangan gizi | Ready to use |  | Ruslan
09000584 | 2025-06-23 | RSUD Sanana | Food Trolly | Syastech | 250435213000400013 | Instalasi | Install sebanyak 2 unit dan sudah dibagikan ke ruangan gizi | Ready to use |  | Ruslan
09000583 | 2025-06-23 | RSUD Sanana | Instrumen Trolly | Syastech | 250469101000400143 X 4 unit | Instalasi | Install 4 unit dan sudah dibagikan ke ruangan | Ready to use |  | Ruslan
09000582 | 2025-06-23 | RSUD Sanana | Premium Baby Bed | Syastech | 250526211000400058 | Instalasi |  | Ready to use |  | Ruslan
09000582 | 2025-06-23 | RSUD Sanana | Premium Baby Bed | Syastech | 250526211000400057 | Instalasi |  | Ready to use |  | Ruslan
09000581 | 2025-06-23 | RSUD Sanana | Infused Stand | Syastech | 250336163000400645 X 25 unit | Instalasi | Install sebanyak 25 unit dan sudah di distribusikan ke ruangan sesuai permintaan rumah sakit | ready to use |  | Ruslan
07000721 | 2025-06-23 | RSUD Purworejo | Ventilator | Draeger/Baby Log-VN800 | ASPC-0037 | Instalasi | Demo unit dan cek kelengkapan aksesoris | Ready to use |  | Wahyu
08000688 | 2025-06-23 | RS PKU Muhammadiyah Mayong | Anaesthesia | Draeger/Fabius Plus | ASRJ-0081 | Service | Replace o2 sensor capsule | Ready to use |  | Amna
19000409 | 2025-06-23 | RS Ken Saras | Mesin Anasthesia | Draeger/Fabius Plus G5 Premium | AS BC-0055 | Service | Coba penggantian baterai namun masih belum bisa | Pending | Disarankan Kontrak service 2 tahun, free mainboard | Nanda
16000292 | 2025-06-23 | Signature Dental Care | Dental Unit | Kavo/1058TM | 10207194 | Service | Install ulang aplikasi conexio | Ready to use |  | Ipan
20000348 | 2025-06-20 | RSUD Anuntaloko | Ventilator | Draeger/Evita V300 | ASHA 0097 | Service | Tarik error log | Ready to use | Unit perlu penggantian part | Sujana
011000416 | 2025-06-20 | RSUP Wahidin Sudirohusodo | Ventilator | Draeger/Savina 300 | ASHD-0009 | Survey | Cek alat | Pending | Disarankan penggantian service set 3y, batterai PS500, flow sensor, power supply | Nasrul
011000416 | 2025-06-20 | RSUP Wahidin Sudirohusodo | Ventilator | Draeger/Savina 300 | ASEF-0135 | Survey | Cek alat | Pending | Disarankan penggantian service set 3y, batterai PS500, flow sensor, power supply | Nasrul
011000417 | 2025-06-20 | RSUP Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASMF-0003 | Survey | Cek PS500, cek baterai internal, system test | Pending | Disarankan lakukan penggantian maintenance kit, baterai internal dan baterai PS500 | Nasrul
07000726 | 2025-06-20 | RSUD Kota Yogyakarta | Mesin Anasthesia | Draeger/Primus | ASLD-0030 | Maintenance | Tes fungsi mixer, tes fungsi VGC, dan tes fungsi PGM | Ready to use |  | Wahyu
07000725 | 2025-06-20 | RSUD Kota Yogyakarta | Mesin Anasthesia | Draeger/Primus | ASMF-0113 | Maintenance | Tes fungsi mixer, tes fungsi VGC, dan tes fungsi PGM | Ready to use |  | Wahyu
03000386 | 2025-06-20 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Kalibrasi meja dan ganti limit switch collum | Ready to use |  | Imam
20000347 | 2025-06-19 | RSD Madani | Ventilator | Draeger/Savina 300 | ASSL-0013 | Maintenance | Pengecekan Warranty | Ready to use |  | Sujana
20000346 | 2025-06-19 | RSUD Anutapura | Ventilator | Draeger/Babylog VN500 | ASML-0098 | Maintenance |  | Ready to use |  | Sujana
4000500 | 2025-06-19 | RSUP Dr. M. Djamil Padang | Pass Box | SYAS | - | Instalasi | Telah diinstall 2 unit passbox dan 1 lemari yang tersambung listrik | Ready to use |  | Satria, Bahar, Sutan, Farhat
08000687 | 2025-06-19 | RSUD Dr. Loekmono Hadi Kudus | Wheel Chair | Syastech/SW 01-SS | 250631313000400482 | Instalasi | Installation | Ready to use |  | Amna
08000687 | 2025-06-19 | RSUD Dr. Loekmono Hadi Kudus | Wheel Chair | Syastech/SW 01-SS | 250631313000400481 | Instalasi | Installation | Ready to use |  | Amna
08000687 | 2025-06-19 | RSUD Dr. Loekmono Hadi Kudus | Wheel Chair | Syastech/SW 01-SS | 250631313000400480 | Instalasi | Installation | Ready to use |  | Amna
08000687 | 2025-06-19 | RSUD Dr. Loekmono Hadi Kudus | Wheel Chair | Syastech/SW 01-SS | 250631313000400479 | Instalasi | Installation | Ready to use |  | Amna
08000687 | 2025-06-19 | RSUD Dr. Loekmono Hadi Kudus | Wheel Chair | Syastech/SW 01-SS | 250631313000400478 | Instalasi | Installation | Ready to use |  | Amna
08000687 | 2025-06-19 | RSUD Dr. Loekmono Hadi Kudus | Wheel Chair | Syastech/SW 01-SS | 250631313000400483 | Instalasi | Installation | Ready to use |  | Amna
08000687 | 2025-06-19 | RSUD Dr. Loekmono Hadi Kudus | Wheel Chair | Syastech/SW 01-SS | 250631313000400476 | Instalasi | Installation | Ready to use |  | Amna
08000687 | 2025-06-19 | RSUD Dr. Loekmono Hadi Kudus | Wheel Chair | Syastech/SW 01-SS | 250631313000400475 | Instalasi | Installation | Ready to use |  | Amna
19000408 | 2025-06-19 | RS Nasional Diponegoro | Infant Warmer | Draeger/Resuscitaire | AF10306 | Service | Cek alat, setting zero pressure , set pressure regulator dan saklar on/off rusak | Ready to use | Unit sementara bisa digunakan tapi berdengung | Nanda
08000687 | 2025-06-19 | RSUD Dr. Loekmono Hadi Kudus | Wheel Chair | Syastech/SW 01-SS | 250631313000400477 | Instalasi | Installation | Ready to use |  | Amna
07000724 | 2025-06-19 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM-0073 | Maintenance | Replace Service set 3 year, replace breathing assy 2 year | Ready to use |  | Wahyu
07000723 | 2025-06-19 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM-0072 | Survey | Pengecekan device | Pending |  | Wahyu
4000499 | 2025-06-18 | RSUP Dr. M. Djamil Padang | Pendant Anasthesi | SYAS/1422A | SYPA2501001 | Instalasi | Unit sudah terpasangan namun belum tersambung dengan listrik | Ready to use |  | Satria, Bahar, Sutan, Farhat
4000499 | 2025-06-18 | RSUP Dr. M. Djamil Padang | Pendant Surgery | SYAS/1422S | SYPS2501002 | Instalasi | Unit sudah terpasangan namun belum tersambung dengan listrik | Ready to use |  | Satria, Bahar, Sutan, Farhat
07000722 | 2025-06-18 | RSUD Ajibarang | Seattle PAP | Draeger/Oxymixer | QM20252046 | Instalasi | Uji fungsi dan cek instalasi alat | Ready to use |  | Wahyu
19000407 | 2025-06-18 | RSUD Dr. Gunawan Mangunkusomo Ambarawa | Incubator Transport | Draeger/Ti 500 | KY 11399 | Service | Cek Unit dan cek fan | Pending | Unit sudah lebih dari 3 tahun, baterai sudah drop, dan disarankan penggantian baterai internal dan transport incubator | Nanda
19000406 | 2025-06-18 | RSUD Dr. Gunawan Mangunkusomo Ambarawa | Incubator Transport | Draeger/Ti 500 | KY 11400 | Service | Cek unit, cek alat, dan baterai internal drop | Pending | Disarankan penggantian baterai internal dan transport | Nanda
07000722 | 2025-06-18 | RSUD Ajibarang | Seattle PAP | Draeger/Oxymixer | QM20255047 | Instalasi | Uji fungsi dan cek instalasi alat | Ready to use |  | Wahyu
07000722 | 2025-06-18 | RSUD Ajibarang | Seattle PAP | Draeger/Oxymixer | QM20255044 | Instalasi | Uji fungsi dan cek instalasi alat | Ready to use |  | Wahyu
03000385 | 2025-06-18 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | J166F0184 | Survey | Pengecekan Seluruh fungsi alat dan fungsi pergerakan meja | Pending | Disarankan penggantian part PC Detector | Imam, Irfan
16000291 | 2025-06-17 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Service | Penggantian applicator | Ready to use |  | Ipan
09000647 | 2025-06-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0180 | Survey | Setelah dicek kerusakan mungkin ke kabel flow sensor | Pending |  | Irfan
20000362 | 2025-06-17 | RSU Makassar | Aesculab Transfonder | Aesculab |  |  |  |  | Serah terima kekurangan accesories transfonder | Sujana
02000199 | 2025-06-16 | RSUD Dr. Drajat Prawiranegara (KSO) | Cathlab | Siemens/Artis one | 82377 | Service | Perbaikan board generator dan cek soket kabel yang terkoneksidi board interface | Ready to use |  | Eddy
08000685 | 2025-06-16 | RSUD Dr. Loekmono Hadi Kudus | Infant Warmer | Draeger/Resuscitaire | AY-11083 | Survey | Pengembalian board infant warmer resuscitaire | Pending | Disarankan Penggantian kit upgrade 4th ED 230V | Amna
19000405 | 2025-06-16 | RSUD Dr. Moewardi Surakarta | BMD | Medilink/Medix DR | A180 16 M 301 | Survey | Cek unit, cek alat, QC, coba scanning, dan ok | Ready to use |  | Nanda
16000290 | 2025-06-16 | Signature Dental Care | Dental Unit | Kavo/1058TM | 10207194 | Maintenance | Perapihan kabel HDMI dan CPU | Ready to use |  | Ipan
03000384 | 2025-06-16 | RSUD Al-Ihsan | Bor Bedah Syaraf | B-Braum/Elan 4 1 GA 800 | 008095 | Instalasi |  | Ready to use |  | Imam
20000360 | 2025-06-13 | RSU Makassar | MOT | SYAS |  | Survey | Perlu perbaikan HVAC | Pending | Perlu perbaikan HVAC | Sujana
07000719 | 2025-06-13 | RSUD Nyi Ageng Serang (KSO) | CT Scan | Siemens/Somatom 90 Top | 120196 | Survey | Replace Baterai, cek tegangan baterai 430V, cek tegangan UPS 435V | Ready to use | Unit dapat digunakan dengan keterbatasan | Wahyu
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0351 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
4000498 | 2025-06-12 | RSUP Jayapura | Digital Radiografi Fluoroskopi | Apelem/Platinum | H248L0157 | Survey | Training user dan pendampingan pasien | Ready to use |  | Satria, Bahar
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0304 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0308 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0287 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0213 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0286 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0319 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0312 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0307 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0281 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0310 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0284 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0283 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0316 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0309 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0295 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0300 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
09000577 | 2025-06-12 | RSUD Sanana | Patient Monitor | Draeger/Vista 300 | VSSTD 0313 | Instalasi | Install uji fungsi dan di training | Ready to use |  | Ruslan
18000604 | 2025-06-12 | RSUD Iwan Bokings | Mesin Anasthesia | Draeger/Atlan A100 XL | ASUB-0212 | Instalasi |  | Ready to use | Telah Selesai Training User Mesin Anasthesia Atlan | Irfan
20000245 | 2025-06-12 | RSUP Makassar | Fluruoscopy | Apelem/Platinum NEO | G248L0145 | Service | Penggantian kabel PACS | Ready to use |  | Sujana
011000414 | 2025-06-12 | RSUP Makassar | Ventilator | Draeger/Babylog VN800 | ASTM-0102 | Survey | Cek baterai, konfigurasi baterai internal | ready to use |  | Nasrul
02000198 | 2025-06-12 | RSUD Dr. Drajat Prawiranegara (KSO) | Cathlab | Siemens/Artis one | 82377 | Sevice | Pelumasan dan pembersihan break | Ready to use | Disarankan pasang unit dehumidifier untuk mengurangi kelembapan ruangan | Eddy
08000684 | 2025-06-12 | RS PKU Fastabiq Sehat Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASTH-0060 | Service | Pemasangan cozy upgrade (a-cone switch TSB310) | Pending | disarankan penggantian o2 sensor capsule | Amna
19000404 | 2025-06-12 | RSUD Tidar Magelang | BMD | Medilink/Medix DR | 15016M258 | Survey | Cek unit, cek koneksi kabel LAN dan ok | pending | Scanner not connect, QC not ready dan akan dijadwalkan lagi | Nanda
15000299 | 2025-06-11 | RSUP Dr. Mohammad Hoesin (KSO) | Printer Fuji | Fuji/Drypix | 66431571 | Service | Atur ulang settingan komputer dengan jumlah film | Ready to use |  | Sutan
09000646 | 2025-06-11 | RSUD Otanaha Gorontalo | Ventilator | Evita V800/Draeger | ASPL-0095 | Survey | Setelah dicek flow sensor putus | pending | Harus diganti yang baru agar alat berfungsi dengan baik | Irfan
19000403 | 2025-06-11 | RSUD Dr. Soeratno Gemolong | Jaundice Meter | Draeger/JM105 | B3503803 | Survey | Error 05. Cek alat dan test unit | pending | Ada kerusakan pada baterai internet dan kalibrasi sudah on dute | Nanda
16000289 | 2025-06-11 | RSU Haji Medan | Dental Unit | Kavo/1058TM | 10207192 | Survey | Muncul error, pengecekan mainboard dental unit | Pending | Disarankan penggantian mainboard | Ipan
08000683 | 2025-06-11 | RSJD Amino Gondohutomo | Jaundice Meter | Draeger/JM-105 | H3510377 | Service | Replace Baterai JM-105 | Ready to use |  | Amna
09000580 | 2025-06-11 | RSUD Sohan kep. Sula | Patient Monitor Invasne with BIS | Draeger/Vista 120 | VGSTJ0117 | Instalasi | Install uji fungsi dan di training sepaket dengan mesin anasthesia ATLAN A300 | Ready to use |  | Ruslan
09000576 | 2025-06-11 | RSUD Sanana | Mesin Anasthesia | Draeger/Atlan A300 with BIS | ASSD-0233 | Instalasi | Install ujifungsi dan training Mesin Anasthesia ATLAN A300 with BIS di ruangan OK RSUD Sanana | Ready to use |  | Ruslan
12000227 | 2025-06-10 | RSUP Prof. Dr. R. D. Kandou Manado | Bor Tulang | BBraun/ELAN4 | GA-800 | Survey | Uji fungsi unit backup | Ready to use |  | Marshall
07000717 | 2025-06-10 | RSUD Bagas Waras | Mesin Anasthesia | Draeger/Perseus A500 | ASJD-0067 | Survey | Self test, cek error log, dan pump faulty | Ready to use | Perlu penggantian oli pump 200 | Wahyu
08000686 | 2025-06-10 | RSUD Kayen Pati | Infant Warmer | Draeger/Resuscitaire | BV 05963 | Survey | Service mode, cek 07, cal SSR, device running | Ready to use |  | Amna
19000402 | 2025-06-10 | RS UNS Surakarta | Ventilator | Evita V800 | ASRH-0010 | Survey | Demo unit ventilator | Ready to use |  | Nanda
17000216 | 2025-06-10 | RSUP Jayapura | Digital Radiografi Fluoroskopi | Apelem/Platinum | H248L0157 | Service | Mini PC rusak. Penggantian Mini PC (Warranty) | Ready to use |  | Bahar, Satria
3000383 | 2025-06-10 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Maintenance | Preventive Maintenance Contract Service | Ready to use |  | Imam
11000422 | 2025-06-10 | RS Regional dr. Hasri Ainun Habibie | Bubble CPAP | Draeger/Oxymixer | QM20246035 | Service | E87 (humidifier) ; Penggantian electronic PCB Assembly | Ready to use |  | Nasrul
3000377 | 2025-06-06 | RSUD dr. H. Moh. Anwar | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/084/C1 | Service | PC workstation error. Setting bootable HDD bios OC | Ready to use |  | Imam
18000602 | 2025-06-06 | RSUD Tani Nelayan | Kompresor Medis Air | PN 300 | 2410103039 | Instalasi |  | Ready to use | Uji Fungsi Satu unit compresor air medis | Irfan
19000438 | 2025-06-05 | Puskesmas Setabelan Surakarta | Bed Gynecolog | Syastech | 250 564 112000400107 | Instalasi | Unit bed sudah ter install dengan baik | Ready to use |  | Nanda
15000298 | 2025-06-05 | RSUP Dr. Mohammad Hoesin | X-ray Digital | Apelem/Camargue DR | IMA704028 | Service | Install ulang aplikasi DROC OK | Ready to use |  | Sutan
02000197 | 2025-06-05 | RSUD Kab. Temanggung | C-Arm | Apelem/EVO R | 551-15-011-469 | Service | Penggantian baterai CMOS | Ready to use |  | Eddy
19000439 | 2025-06-04 | RS UNS Surakarta | Ventilator | Draeger/Evita V800 | ASPH-0120 | Instalasi | Unit demo sudah terinstall dengan baik | Ready to use |  | Nanda
18000603 | 2025-06-04 | RSUD Zainal Umar Sidiki Gorontalo Utara | Ventilator | Babylog VN800/Draeger | ASSD-0008 | Service | Update software | Ready to use |  | Irfan
08000682 | 2025-06-04 | RSUD Dr. Loekmono Hadi Kudus | Infant Warmer | Draeger/Resuscitaire | AY 11083 | Survey | Service mode, cek skep dan cek voltage | Pending | Error 04 | Amna
06000842 | 2025-06-04 | RSUD Bumi ayu | Infant Warmer | Draeger/Resuscitaire | CY03650 | Survey | Cek heater assembly, board heater, compare board | Pending | Rekomendasi penggantian heater assembly | Dwi
07000718 | 2025-06-04 | RSUD Merah Putih | X-ray General | Apelem/Camargue | 703979 | Service | Cek board inverter, cek generator, cek kabel HTT anoda, cek pelumas kabel HTT, dan menambahkan pelumas pada kabel HTT | Ready to use |  | Wahyu, Imam
15000297 | 2025-06-03 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Bongkar table top, cek motor dekat detektor (bersihkan) | Ready to use |  | Sutan
12000225 | 2025-06-03 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Helianthus DBT | HELI DBT/130 | Survey | Keluarkan film yang nyangkut dari dalam tray | Ready to use |  | Marshall
09000575 | 2025-06-03 | RSUD Jailolo Halmahera Barat | Patient Monitor | Draeger/Vista 300 | VSSTE 0117 | Instalasi | Install uji fungsi dan di training dengan jumlah unit | Ready to use |  | Ruslan
06000841 | 2025-06-03 | RSUD. Brebes | Mesin Anasthesia | Draeger/Primus | ASHH-0153 | Survey | Cek dan compare sparepart, peep value, pump assembly | Pending | Ada kerusakan Sparepart peep value dan pump assembly | Dwi
09000575 | 2025-06-03 | RSUD Jailolo Halmahera Barat | Patient Monitor | Draeger/Vista 300 | VSSTE 0115 | Instalasi | Install uji fungsi dan di training dengan jumlah unit | Ready to use |  | Ruslan
18000601 | 2025-06-03 | RSUD Iwan Boking Gorontalo | Mesin Anasthesia | Draeger/Atlan A100 | ASUB-0212 | Instalasi | Install dan uji fungsi satu unit Mesin Anasthesia Atlan A100 Draeger | Ready to use |  | Irfan
08000681 | 2025-06-03 | RSUD Ki Ageng Selo | Mesin Anasthesia | Draeger/Fabius Plus | ASMK-0137 | Service | Replace o2 sensor capsule | Ready to use |  | Amna
3000382 | 2025-06-03 | RSUD Dr. Soediran Mangun Sumarso | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/134/C2 | Survey | Workstation tidak connect ke printer, error. Setting tray printer, software workstation | Ready to use |  | Imam, Nanda
07000716 | 2025-06-03 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Atlan A300 | ASTK-0060/ASTK-0061 | Survey | Cek konektor air, cek filter, cek pressure regulator, cek blok mixer, dan cek breathing system | Ready to use |  | Wahyu
20000359 | 2025-06-02 | RSP Makassar | MOT | SYAS |  | Survey | Pengecekan HVAC MOT suhu naik turun | Pending | perlu perbaikan HVAC | Sujana
07000715 | 2025-06-02 | RSUD Merah Putih | X-ray General | Apelem/Camargue | 703979 | Survey | Cek ceiling, bersikan rel ceiling, dan cek roda ceiling | Ready to use | Disarankan ganti roda ceiling | Wahyu
19000481 | 2025-06-02 | RSUP Dr. Kariadi Semarang | ESWL | Medispec/EM1000 | 0922164 | Survey | Ruangan sudah di cek yang diperlukan pembuatan plumbing untuk kabel saja | Ready to use |  | Nanda
19000401 | 2025-06-02 | RSUD Dr. Soediran Mangun Sumarso | Mamografi | Helianthus DBT/Metaltronica | HELIDBT/134/C2 | Service | Cek alat, restore WS apps, setting and adjustment | Ready to use |  | Nanda
15000296 | 2025-05-28 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Service | Bersihkan board S22 pakai contact cleaner ok, cek fuse ok | Ready to use |  | Sutan
12000214 | 2025-05-28 | RSUD Ainun Habibi | Draeger | Resucitaire (Infant Warmer) | DY03435 | Service | Ganti part Valve, control mini S-way dan uji fungsi | Ready to use |  | Marshall
12000212 | 2025-05-28 | RSUD Ainun Habibi | Mesin Anasthesia | Draeger/Fabius | ASLL-016 | Service | Pergantian part dan uji fungsi unit (O2 Sensor) | Ready to use |  | Marshall
12000211 | 2025-05-28 | RSUD Hasri Ainun Habibie | Mesin Anasthesia | Draeger Atlan A300 | ASMM-0118 | Service | Pergantian part dan uji fungsi unit ( Oxxy cell O2 sensor) | Ready to use |  | Marshall
19000480 | 2025-05-28 | RSUD Dr. Soediran Mangun Sumarso Wonogiri | Pasien Monitor Central | Draeger/Vista 120 CMS | 35297369247 | Service | Cek computer server, setting ulang IP, reconnecting dan connect all | Ready to use |  | Nanda
7000713 | 2025-05-28 | RSUD Bagas Waras | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASHM0014 | Survey | Mve dan Vte tidak muncul. Device check, cek dan kalibrasi flow sensor | Pending | Unit perlu penggantian Diff. Pressure flow sensor | Wahyu
7000714 | 2025-05-28 | RSUD Bagas Waras | Infant Warmer | Draeger/Resuscitaire | JY03385 | Service | Error 05. Relay heater rusak | Pending | Unit perlu penggantian PCB power dan control | Wahyu
8000679 | 2025-05-28 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Service | Pemasangan service set Kit Cozy part | Ready to use |  | Amna
9000568 | 2025-05-28 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Evita V300 | ASMD0029 | Service | Gas supply failed. Penggantian pressure regulator dan medical compressor | Ready to use |  | Ruslan
9000570 | 2025-05-28 | RSUD Dr. H. Chasan Boesoirie | ESWL | Medispec/E3000 | 0922007 | Service | Unit tidak bisa menyala. Penggantian baterai CMOS dan setting bios | Ready to use |  | Ruslan
9000569 | 2025-05-28 | RSUD Dr. H. Chasan Boesoirie | Medical Air Compressor | Draeger/PN3000 | 2409103119 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Ruslan
011000412 | 2025-05-27 | RSUD Makassar | Ventilator | Draeger/Evita V300 | ASDH-0031 | Survey | Cek out-input baterai external | Ready to use | Unit dapat digunakan dengan keterbatasan atau baterai drop | Nasrul, Sujana
18000646 | 2025-05-27 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | HSEK0109 | Service | Melakukan tarik error menggunakan service kontak draeger dan bermasalah pada pressure regulator | Pending | Disarankan untuk mengganti Pressure regulator | Irfan R
19000479 | 2025-05-27 | RSUD Dr. Moewardi | BMD | Medilink/Medix DR | A18016M301 | Sevice | Penggantian Monoblock, setting dan adjusment | Ready to use | Unit sementara dapat digunakan, kurang remote dari principal | Nanda
8000678 | 2025-05-27 | RSI Aisyiyah Kudus | Ventilator | Draeger/Babylog VN800 | ASPF0017 | Survey | Refresh training | Ready to use |  | Amna
20000358 | 2025-05-27 | RSP Makassar | MOT | Syas | - | Service | Perbaikan Scrub sink MOT | Ready to use |  | Sujana
3000381 | 2025-05-27 | RSU Haji Medan | Bone Mineral Densitometer | Medilink/Medix DR | D197H0560 | Survey | Generator fault. Cek tegangan supply inverter, cek koneksi cpu board. Kalibrasi filament gagal | Pending | Monoblock kemungkinan rusak | Imam
011000411 | 2025-05-26 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Babyleo TN500 | ASTH-0130 | Service | Cek engsel pintu access port | Pending | Disarankan ganti engsel pintu access port | Nasrul
18000647 | 2025-05-26 | RSUD Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | VGSSF0022 | Survey | Pengecekan dan setting EKG | Ready to use |  | Irfan R
8000677 | 2025-05-26 | RSI Aisyiyah Kudus | Meja Operasi | SYAS | 250552512100400010 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
7000709 | 2025-05-26 | RSUP dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Service | Penggantian flow power | Ready to use |  | Wahyu
7000712 | 2025-05-26 | RSUD Wonosari | Ventilator | Draeger/Babylog VN800 | ASPF0004 | Service | Gas supply calibration failure. Cek gas central, kompresor | Ready to use | Unit bisa digunakan namun disarankan penggantian air inlet filter dan baterai internal | Wahyu
20000363 | 2025-05-24 | RSUP Makassar | HVAC MOT | SYAS |  | Survey | Pengecekan suhu MOT | Pending | Perlu perbaikan HVAC | sujana
8000660 | 2025-05-23 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKK0117 | Survey | Leak/compliance test fail, ventilator fail. Service mode, calibration motor pumo | Ready to use |  | Amna
7000711 | 2025-05-23 | RSPAU dr. S. Hardjolukito | Ventilator | Draeger/Babylog VN600 | ASRM0091 | Service | Unit tidak menyala. Cek kabel monitor, cek power supply, cek board power, cek M48 | Pending | Unit perlu penggantian M48 set | Wahyu
16000284 | 2025-05-23 | GiGi Dental House | Dental Unit | Kavo/1058TM | 10207196 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Ipan, Farhat
16000285 | 2025-05-23 | GiGi Dental House | Dental Unit | Kavo/E50 | 30211887 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Ipan, Farhat
19000478 | 2025-05-23 | RSUP Dr. Kariadi Semarang | Ventilator | Draeger/Savina 300 | ASMK-0301 | Maintenance | Pasang maintenance kit | Ready to use |  | Nanda
16000286 | 2025-05-23 | GiGi Dental House | Dental Unit | Kavo/E50 | 30211888 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Ipan, Farhat
12000213 | 2025-05-23 | RSUD Ainun Habibie | MOT SYAS | HVAC SYAS |  | Survey | Bersihkan filter medium filter | Ready to use |  | Marshall
20000344 | 2025-05-23 | RSUD Trikora Salakan | Mesin Anasthesia | Draeger/Primus | ASHD-0036 | Sevice | Pemasangan part service set, PBA VGC Analog, dan presure regulator | Ready to use |  | Sujana
3000380 | 2025-05-23 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Meja error tidak bisa posisi tidur. Cek koneksi DRV 1 dan reconnect | Ready to use |  | Imam
19000477 | 2025-05-22 | RSUD Ki Ageng Selo Grobogan | Infant Warmer | Draeger/Resuscitaire | AC13515 | Survey | Cek alat, cek service mode, test 10, fail. exchange other board, ok | Ready to use | Board display rusak namun masih bisa ketika dipasang | Nanda
011000413 | 2025-05-22 | RSUD Kota Makassar | ESST | Medispec/ED1000 | 0316014 | Survey | Cek ECD, cek baterai epron | Ready to use |  | Nasrul
15000295 | 2025-05-22 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Tube tidak bisa bergerak; Bersihkan sensor grid | Ready to use |  | Sutan
16000283 | 2025-05-22 | RSPPM Soedirman | Dental Unit | Kavo/1058TM | 10207193 | Service | Threeway syringe ada bocor angin. Perbaikan threeway syringe | Ready to use |  | Ipan, Farhat
8000676 | 2025-05-22 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASNH0240 | Survey | Gas mixer fault. Reload software | Ready to use |  | Amna
4000496 | 2025-05-21 | RSUD Dr. H. Abdul Moeloek | Mobile X-ray | Apelem/Rafale | 1769292100001 | Survey | Pengecekan mobile Xray | Ready to use |  | Satria, Bahar
7000710 | 2025-05-22 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Babyleo TN500 | ASTL0117 | Maintenance | Maintenance Warranty | Ready to use |  | Wahyu
14000357 | 2025-05-21 | RSUD Dr. Soetomo Surabaya | ESMR | Medispec/CS-2 | 03-10297 | Survey | Cek power supply ESMR, Cek Power Supply USG, Cek EKG, Cek shock wave, cek water sirkulasi, system | Pending | perlu penggantian System water circulation rusak, aplication head ESMR rusak | Fahim
011000409 | 2025-05-21 | RSUD Makassar | C-arm | Apelem/Evo R | 551-24-024-1086 | Service | Maintenance Checklist Prosedure | Ready to use |  | Nasrul
011000408 | 2025-05-21 | RSUD Makassar | Mobile X-ray | Apelem/Rafale BDR | 971-24-084-483 | Service | Maintenance chechlist prosedure | Ready to use |  | Nasrul, Imam
011000407 | 2025-05-21 | RSUD Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0145 | Service | Maintenance checklist, x-ray room configurasi checks | Ready to use |  | Nasrul, Imam
18000648 | 2025-05-21 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0165 | Survey | pengecekan dan ada kerusakan pada baterai | Pending | Disarankan penggantian baterai | Irfan R
8000675 | 2025-05-21 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHA0113 | Service | Baterai fail & system check nebulizer fail. Penggantian baterai dan pressure regulator | Ready to use |  | Amna
8000675 | 2025-05-21 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHN 0038 | Service | Baterai fail & system check nebulizer fail. Penggantian baterai dan pressure regulator | Ready to use |  | Amna
6000837 | 2025-05-21 | RSKBD Panti Nugroho | Ventilator | Draeger/Evita V300 | ASLD0168 | Survey | O2 measurement fail, flow sensor fail, air supply low, baterai fail. Selft test. | Pending | Disarankan penggantian O2 cell dan service set 2 tahun | Dwi
16000287 | 2025-05-20 | Signature Dental Care | Dental Unit | KAVO 1058 |  | Service | Setting pada air bowl, pembersihan saluran air pada 3 jalur,pemindahan tray dari kiri ke kanan, setting timer dan perapihan kabel mini pc | Ready to use |  | Farhat, Ipan
16000288 | 2025-05-20 | Signature Dental Care | Dental Unit | KAVO 1058 |  | Service | Instalasi ulang pada software | Ready to use |  | Farhat, Ipan
18000649 | 2025-05-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | ASJD0020 | Survey | pengecekan dan ada kerusakan pada kabel flow sensor | Pending | Disarankan penggantian cable harnes dan baby log | Irfan
17000215 | 2025-05-20 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Survey | HVG tidak bisa menghasilkan output energi, EMH rusak. Cek input HVG, cek bagian dalam EMH, cek leveling EMH dan thermal switch | Pending | EMH rusak | Bahar, Satria
3000379 | 2025-05-20 | RSUP Makassar | C-arm | Apelem/Evo R+ | 551240241086 | Service | Penggantian inverter board dan reload software | Ready to use |  | Imam, Nasrul
6000835 | 2025-05-20 | RSUD Dr. M. Ashari | Mesin Anasthesia | Draeger/Primus | ASMF0114 | Survey | Pengecekan unit, self test, running test | Ready to use | Rekomendasi penggantian waterlock, sample line, flow sensor | Dwi
19000476 | 2025-05-20 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Fabius Plus/Draeger | ARZM-0051 | Maintenance | Pasang service set 3 year dan cozy 2 year | Ready to use | Unit sudah dipasang partnya | Nanda
6000835 | 2025-05-20 | RSUD Dr. M. Ashari | Mesin Anasthesia | Draeger/Primus | ASLD0025 | Survey | Pengecekan unit, self test, running test | Ready to use | Rekomendasi penggantian waterlock, sample line, flow sensor | Dwi
6000835 | 2025-05-20 | RSUD Dr. M. Ashari | Mesin Anasthesia | Draeger/Primus | ASMJ0131 | Survey | Pengecekan unit, self test, running test | Ready to use | Rekomendasi penggantian waterlock, sample line, flow sensor | Dwi
7000692 | 2025-05-19 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASJH0121 | Maintenance | Maintenance Warranty | Ready to use | Penggantian O2 sensor, pressure regulator, internal baterai, particle filter, kit air filter | Wahyu
15000294 | 2025-05-19 | RSUP Dr. Mohammad Hoesin | X-ray Digital | Apelem/Camargue DR | IMA704028 | Service | Kolimator mati; Bersihkan board kolimator dengan contact cleaner | Ready to use |  | Sutan
8000674 | 2025-05-19 | RS Panti Wilasa Dr. Cipto | Ventilator | Draeger/Savina300 | ASMM0026 | Survey | VT high, paw high alarm. General check, system check, cek membrane valve dan flow sensor. Unit sudah berfungsi kembali | Ready to use |  | Amna
6000834 | 2025-05-19 | RS PKU Muhammadiyah Gombong | Mesin Anasthesia | Draeger/Fabius Tiro | ARZM0024 | Survey | Display tidak ada. Display unit copot | Pending | Disarankan penggantian display fabius PN: 607102 | Dwi
19000475 | 2025-05-19 | RSUD Mijen Semarang | USG | Mindray/Z50T | HH3B2B000013 | Survey | Cek alat, cek detail 4D board not available | Ready to use | Perlu penambahan board dan probe | Nanda
20000357 | 2025-05-19 | RSUP Makassar | MOT HVAC | Syas | - | Service | Pendampingan penggantian kompresor AC MOT | Pending | Disarankan penggantian Kompresor HVAC | Sujana
12000224 | 2025-05-16 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Perseus A500 | ASHE0004 | Survey | Cek part baterai | Ready to use | Baterai sudah ditemukan | Marshall
20000343 | 2025-05-16 | RS UNHAS | Mesin Anasthesia | Draeger/Primus | ASRF-0060 | Survey | Pengecekan power suplay | Pending | Perlu penggantian power suplay | Sujana, Nasrul
7000691 | 2025-05-16 | RS JIH | Mesin Anasthesia | Draeger/Atlan A350 | ASRJ0460 | Maintenance | Maintenance Warranty | Ready to use |  | Wahyu
8000673 | 2025-05-16 | RSUD Dr. Loekmono Hadi | Infant Warmer | Draeger/Resuscitaire | AY11083 | Service | Error 05, heater rusak. Penggantian heater assy | Ready to use |  | Amna
6000833 | 2025-05-16 | RSUD Dr. Soeselo Slawi | Ventilator | Draeger/Savina300 | ASHD0014 | Service | Penggantian baterai external | Ready to use |  | Dwi
7000688 | 2025-05-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Atlan A350 | ASTK0060 | Service | Sampling line occluded dan monitoring gas fail. Penggantian mini module pump | Ready to use |  | Wahyu
19000474 | 2025-05-15 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Fabius Plus/Draeger | AR2M-0051 | Service | Pasang sparepart | Ready to use | Sparepart sebagian sudah dipasang tetapi ruangan mau digunakan untuk operasi, dilanjutkan lain hari pasang sparepart nya | Nanda
8000672 | 2025-05-15 | RSUD Kraton Pekalongan | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASHM0009 | Survey | Error sensor S3 pressure too high. Service mode, test 33, pressure regulator fail | Ready to use | Unit masih bisa digunakan | Amna
20000356 | 2025-05-15 | RSUP Makassar | MOT HVAC OK3/Cathlab | Syas | - | Survey | Pengecekan HVAC | Pending | Perlu perbaikan HVAC | Sujana
011000410 | 2025-05-14 | RS Unismuh Makassar | Mesin Anasthesia | Draeger/ATLAN | ASPE-0199 | Survey | System test, cek val baterai external | Pending | Segera ganti baterai external dan maintenance kit | Nasrul
19000473 | 2025-05-14 | RSUD Dr H Soewondo Kendal | Ventilator | Draeger/Baby Log VN 500 | ASNL-0149 | Survey | Cek alat dan cek device | Ready to use | Disarankan penggantian baterai internal, pressure regulator, particle filter, dan kit air filter | Nanda
8000663 | 2025-05-14 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina300 | ASTJ0006 | Maintenance | Maintenance Warranty | Ready to use |  | Amna
8000663 | 2025-05-14 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina300 | ASTJ0007 | Maintenance | Maintenance Warranty | Ready to use |  | Amna
8000663 | 2025-05-14 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina300 | ASNJ0041 | Maintenance | Maintenance Warranty | Ready to use |  | Amna
6000839 | 2025-05-14 | RSUD Ajibarang | Mesin Anasthesia | Draeger/Primus | ASPM0019 | Service | Pemasangan primus water lock, sample line dan flow sensor | Ready to use |  | Dwi
011000406 | 2025-05-09 | RSUD Makassar | Ventilator | Draeger/Evita V300 | ASPH-0032 | Instalasi |  | Ready to use |  | Nasrul
011000406 | 2025-05-09 | RSUD Makassar | Ventilator | Draeger/Evita V300 | ASPH-0033 | Instalasi |  | Ready to use |  | Nasrul
011000406 | 2025-05-09 | RSUD Makassar | Ventilator | Draeger/Evita V300 | ASPH-0031 | Instalasi |  | Ready to use |  | Nasrul
2000193 | 2025-05-09 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Survey | Shockwave tidak bisa. Pengecekan HVG dan HVPS. HVG rusak | Pending |  | Eddy
8000671 | 2025-05-09 | RS PKU Muhammadiyah Mayong | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ0081 | Maintenance | Perlu penggantian O2 sensor | Ready to use |  | Amna
8000671 | 2025-05-09 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Babylog VN600 | ASSK0028 | Maintenance | Perlu penggantian O2 sensor | Ready to use |  | Amna
15000293 | 2025-05-09 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Data base penuh. Hapus data | Ready to use |  | Sutan
15000292 | 2025-05-09 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Service | Filament fault. Pembersihan board S22 | Ready to use |  | Sutan
7000687 | 2025-05-09 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Atlan A300 | ASTK0060 | Survey | Sampling line occluded . Cek unit dan pump error | Pending | Unit disarankan penggantian mini module pump | wahyu
7000686 | 2025-05-09 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Service | Penggantian flow power | Ready to use |  | Wahyu
9000573 | 2025-05-09 | RSD Tidore | Ventilator | Draeger/Babylog VN500 | ASJK0042 | Survey | O2 supply dan air supply tidak terbaca. General check | Pending | Unit perlu penggantian M12.4 (PATO) dan M1.3 (Kit proportional valve) | Ruslan
6000840 | 2025-05-09 | RSUD dr. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN800 | ASEL0081 | Survey | Regular visit JV dengan draeger pusat | Ready to use |  | Dwi
18000650 | 2025-05-09 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASSM0159 | Survey | Pengecekan dan ada kerusakan pada cable harnes | Pending | Disarankan penggantian kabel harnes | Irfan R, Marshall
9000574 | 2025-05-09 | RSUD Sanana | X-ray Digital | Apelem/Camargue DR | 251023012 | Survey | DAP tidak aktif. Pengecekan unit DAP. Masih tidak sesuai ukuran | Pending |  | Ruslan
18000608 | 2025-05-08 | RS Mata Manado | Mesin Anasthesia | A500/Draeger | ASJM-0178 | Instalasi | Pemasangan aksesoris Mesin Anasthesia kontrak service | Ready to use |  | Irfan R, Marshall
19000472 | 2025-05-08 | RSUD Simo Boyolali | Mesin Anasthesia | Draeger/Fabius Plus XL Workstation | ASTC-0004 | Survey | Setting , adjustment dan cek device | Ready to use |  | Nanda
8000670 | 2025-05-08 | RSUD dr. R. Soetijono Blora | Ventilator | Draeger/Babylog VN800 | ASKJ0004 | Survey | Refresh training. Cover kit corner gas inlet fail, part masih kurang | Pending | Alat belum bisa digunakan | Amna
7000685 | 2025-05-08 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | MF31848 | Survey | Nilai kelembaban tidak sesuai. Cek evaporator, heater evaporator | Pending | Unit disarankan penggantian Evaporator Reservoir | Wahyu
2000192 | 2025-05-07 | RSUD Kab. Temanggung | Bone Mineral Densitometer | Apelem/Medix DR | H237H1091 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Eddy
8000669 | 2025-05-07 | RS Assuyuthiyyah | Mesin Anasthesia | Draeger/Fabius Plus | ASUA0003 | Survey | Install conversation kit heater | Ready to use |  | Amna
7000684 | 2025-05-07 | RSUD Merah Putih | Bubble CPAP | Draeger/Oxymixer | QM20230134 | Survey | Pengecekan kelengkapan unit demo | Ready to use |  | Wahyu
6000832 | 2025-05-07 | RSUD Prembun | Ventilator | Draeger/Babyog VN500 | ASLF0056 | Maintenance | Preventive Maintenance Warranty & Refresh Training | Ready to use |  | Dwi
6000832 | 2025-05-07 | RSUD Prembun | Ventilator | Draeger/Babyog VN500 | ASLF0057 | Maintenance | Preventive Maintenance Warranty & Refresh Training | Ready to use |  | Dwi
18000607 | 2025-05-07 | RSUP Prof. Dr. R. D. Kandou Manado | Bor Tulang | Bbraun/ELAN4 | 03946 | Survey | Pengecekan Unit | Pending | Ada kerusakan pada salah satu handpiece bor, sudah dibuat penawaran pergantian handpiece bor | Irfan
12000222 | 2025-05-06 | RSUD Provinsi Sulawesi Utara | Mamografi | Helianthus DBT | Heli DBT/130 | Survey | Restart gambar yang hilang | Ready to use |  | Marshall
20000355 | 2025-05-06 | RSUP Makassar | MOT HVAC | Syas | - | Service | Pengecekan HVAC, setting suhu HVAC |  |  | Sujana
18000645 | 2025-05-06 | RSUD Provinsi Sulawesi Utara | Mamografi | Helianthus | Heli DBT/130/C2 | Survey | Pengecekan Unit | Ready to use |  | Irfan R, Marshall
19000471 | 2025-05-06 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 Workstation | ASPH-0124 | Service | Cek alat dan cek device | Ready to use |  | Nanda
19000470 | 2025-05-06 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL-0089 | Service | Cek device dan setting parameter | Ready to use |  | Nanda
19000470 | 2025-05-06 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL-0075 | Service | Cek device dan setting parameter | Ready to use |  | Nanda
19000470 | 2025-05-06 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRF-0023 | Service | Cek device dan setting parameter | Ready to use |  | Nanda
19000470 | 2025-05-06 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRH-0031 | Service | Cek device dan setting parameter | Ready to use |  | Nanda
19000470 | 2025-05-06 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL-0092 | Service | Cek device dan setting parameter | Ready to use |  | Nanda
19000470 | 2025-05-06 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL-0091 | Service | Cek device dan setting parameter | Ready to use |  | Nanda
19000470 | 2025-05-06 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL-0088 | Service | Cek device dan setting parameter | Ready to use |  | Nanda
19000469 | 2025-05-06 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL-0090 | Service | Cek Unit dan Tarik error | Pending | Disarankan penggantian blower unit GS 500 | Nanda
8000668 | 2025-05-06 | RS Assuyuthiyyah | Mesin Anasthesia | Draeger/Fabius Plus | ASUA0003 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Amna
6000831 | 2025-05-06 | RS JIH Purwokerto | Mesin Anasthesia | Draeger/Primus | ASPC0085 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Dwi
6000831 | 2025-05-06 | RS JIH Purwokerto | Mesin Anasthesia | Draeger/Primus | ASNJ0040 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Dwi
011000405 | 2025-05-06 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Syastech/4 Motor | 240779101260400154 | Survey | Cek remote, cek motor, cek kontrol box | Ready to use |  | Nasrul
011000404 | 2025-05-05 | RSUP Dr. Wahidin Sudirohusodo | Infant Warmer | Draeger/Babyleo TN500 | ASTH-0128 | Survey | adjust settingan arm | Ready to use |  | Nasrul
6000840 | 2025-05-05 | RSUD Ajibarang | Ventilator | Draeger/Babylog VN800 | ASMK0054 | Service | Pemasangan particle filter dan kit air filter | Ready to use |  | Dwi
3000378 | 2025-05-03 | RSUD Tidar Magelang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0496 | Survey | Motor lifter tidak berfungsi untuk turun, ada bunyi pada motor. Resistor break rusak | Pending |  | Imam
6000829 | 2025-05-02 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Babylog VN800 | ASTE0047 | Maintenance | Maintenance Warranty | Ready to use |  | Dwi
11000403 | 2025-04-30 | RSUD Andi Makassau | Ventilator | Draeger/Babylog VN800 | ASSF0003 | Service | Update software versi terbaru | Ready to use |  | Nasrul
7000681 | 2025-04-30 | RSUD Sleman | Ventilator | Draeger/Babylog VN800 | ASRD0053 | Maintenance | Penggantin kit air filter, particle filter, pembersihan dari debu | Ready to use |  | Wahyu
6000826 | 2025-04-30 | RSUD Dr. Soeselo Slawi | Central Monitor | Draeger/Vista 120 CMS | VL5FF0018 | Service | Error windows. Install windows 10 64bit, install cms dan konfigurasi cms | Ready to use |  | Dwi
11000402 | 2025-04-29 | RS Regional dr. Hasri Ainun Habibie | Ventilator | Draeger/Babylog VN800 | ASSM0042 | Service | Update software versi terbaru | Ready to use |  | Nasrul
11000401 | 2025-04-29 | RS Regional dr. Hasri Ainun Habibie | Bubble CPAP | Draeger/Oxymixer | QM20246035 | Survey | Error E87. PCB rusak | Pending | Perlu penggantian PCB | Nasrul
19000468 | 2025-04-29 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH-0124 | Survey | Survey kerusakan touch screen | Pending | Unit ada kerusakan pada modul touch screen | Nanda
12000204 | 2025-04-29 | RSUD Noongan | X-ray Digital | Apelem/Camargue DR | IMA703798 | Survey | Hasil foto hitam, komputer sudah lemot. Kemungkinan rusak ada pada detektor dan perlu upgrade processor dan RAM | Pending |  | Marshall
8000666 | 2025-04-29 | RSI Sunan Kudus | Ventilator | Draeger/Savina300 | ARUM0308 | Survey | Alarm flow measurement inop. Cek flow sensor, valve expiratory | Pending | Perlu penggantian service set 6 thn dan O2 sensor | Amna
8000666 | 2025-04-29 | RSI Sunan Kudus | Ventilator | Draeger/Savina300 | ARUE0022 | Survey | Alarm flow measurement inop. Cek flow sensor, valve expiratory | Pending | Perlu penggantian service set 6 thn dan O2 sensor | Amna
11000400 | 2025-04-28 | RSUD Kota Makassar | Bone Mineral Densitometer | Medilink/Medix DR | 460624 | Survey | Generator fault. Cek trafo, cek inverter. Tegangan terlalu tinggi | Pending | Menunggu pengecekan lebih lanjut setelah UPS baru ada | Nasrul
6000825 | 2025-04-28 | RS JIH Purwokerto | Bubble CPAP | Draeger/Oxymixer | Q1420230185 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Dwi
6000824 | 2025-04-28 | RS JIH Purwokerto | Mesin Anasthesia | Draeger/Primus | ASPL0085 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Dwi
6000824 | 2025-04-28 | RS JIH Purwokerto | Mesin Anasthesia | Draeger/Primus | ASPL0254 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Dwi
6000824 | 2025-04-28 | RS JIH Purwokerto | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0040 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Dwi
20000342 | 2025-04-28 | RSUP Tadjudin Chalid | ESWL | Medispec/E1000 | 0922014 | Service | Pengecekan sirculation selang dan penambahan air | Ready to use |  | Sujana
7000683 | 2025-04-25 | RSUD Sleman | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0128 | Service | Indikator O2 supply fail. Penggantian PCB optimix display | Ready to use |  | Wahyu
8000665 | 2025-04-25 | RS Aisyiyah Kudus | Bubble CPAP | Draeger/Oxymixer | QM20215174 | Service | Alarm humidifier chamber. Setting HC, setting AA, setting OO | Ready to use |  | Amna
11000399 | 2025-04-25 | RSUD Kota Makassar | ESWL | Medispec/E3000 | 814134 | Survey | Pengecekan kerusakan alat. Console unit rusak | Pending |  | Nasrul
3000375 | 2025-04-25 | RSUD Dr. Saiful Anwar | Bone Mineral Densitometer | Medilink/Medix DR | J207H0735 | Service | Update software versi terbaru | Ready to use |  | Imam
18000589 | 2025-04-24 | RSUD Dr. Iwan Boking | Ventilator | Draeger/Savina 300 | ASTM0169 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000589 | 2025-04-24 | RSUD Dr. Iwan Boking | Ventilator | Draeger/Savina 300 | ASTM0168 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
18000589 | 2025-04-24 | RSUD Dr. Iwan Boking | Ventilator | Draeger/Savina 300 | ASTM0170 | Test Function | Unit berfungsi dengan baik | Ready to use |  | Irfan R
3000376 | 2025-04-24 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Error pergerakan meja; Perbaikan driver motor | Ready to use |  | Imam, Satria
17000214 | 2025-04-24 | RSUP dr. Ben Mboi | ESWL | Medispec/EM1000 | 0922162 | Service | EMH expired, C-arm tidak menyala; Setting tanggal mundur generator, penggantian baterai cmos dan pengisian air | Ready to use |  | Bahar
11000398 | 2025-04-23 | RSUP Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Cek EMH, cek koneksi EMH, service koneksi EMH | Ready to use |  | Nasrul
12000208 | 2025-04-23 | RSUD Ainun Habibie | MOT | SYAS | - | Survey | Pengecekan unit. Unit berfungsi dengan baik | Ready to use |  | Marshall, Irfan
4000495 | 2025-04-23 | RSUP Jayapura | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | H248L0157 | Service | Console tidak menyala. PSU dan CPU dibawa untuk di cek | Pending |  | Satria, Imam
15000290 | 2025-04-23 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Service | HVPS dan HVG dibawa ke workshop untuk dicek lebih lanjut | Pending |  | Sutan
12000202 | 2025-04-22 | RSUD Toto Kabila | X-ray Digital | Apelem/Camargue DR | IMA703943 | Survey | Window dan software unit crash. Disarankan untuk reload software | Pending |  | Marshall, Irfan
5000301 | 2025-04-22 | Signature Dental Care | Dental Unit | Kavo/1058TM | 10207194 | Service | Suction tersumbat, sehingga air tidak lancar keluar. Pembersihan jalur suction dan penambahan air oxygenal | Ready to use |  | Farhat, Ipan
15000289 | 2025-04-22 | RSUP Dr. Mohammad Hoesin | X-ray Digital | Apelem/Camargue DR | IMA704028 | Service | Tube ada yang turun, perbaikan soket vert ke soker ang break. Tensator spring ada yang kendor | Ready to use | Sementara unit bisa digunakan dengan perbaikan sementara soket vert ke ang | Sutan
8000664 | 2025-04-22 | RS Islam Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0114 | Survey | Ventilator fail. Pengecekan motor pump, kalibrasi motor pump. O2 sensor, absorber dan kit maintenance Co2y perlu diganti | Pending |  | Amna
11000397 | 2025-04-21 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Primus | ASHD-0026 | Survey | Cek VGC, system test, test running | Ready to use | Disarankan ganti baterai internal dan service set 3y | Nasrul
19000466 | 2025-04-21 | RSJD Dr. Amino | Mesin Anasthesia | Draeger/Fabius Plus | ASLL-0145 | Service | Cek aset dan cek leak test | Ready to use | Disarankan ganti o2 sensor | Nanda
6000821 | 2025-04-21 | RS Mata Purwokerto | Mesin Anasthesia | Draeger/Fabius Plus | ASTH0001 | Service | Serah terima barang quality issue A-cone switch | Ready to use |  | Dwi
8000661 | 2025-04-21 | RS Mitra Bangsa Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0041 | Service | Pemasangan O2 sensor capsule | Ready to use |  | Amna
15000288 | 2025-04-21 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Service | Frekuensi tembakan lemah. Kalibrasi dengan quick seasoning | Pending | Ada kerusakan pada HVG | Sutan
7000682 | 2025-04-17 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Evita V300 | ASPC0205 | Survey | Gas supply calibration failed. System check, cek error log, cek proportional valve; Penggantian proportional valve | Ready to use | Unit berfungsi dengan baik | Wahyu
16000282 | 2025-04-17 | Signature Dental Care | Dental Unit | Kavo/1058TM | 10207191 | Service | Selenoid rusak dan air bowl tidak keluar; Penggantian selenoid valve | Ready to use |  | Ipan
4000507 | 2025-04-17 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Survey | Pengambilan applicator yang bermasalah. Qty: 2pcs | Pending |  | Farhat
8000662 | 2025-04-17 | RSUD Dr. R. Soedjati Soemodiarjo | UPS | OAXACA/AP Series | - | Survey | Survey ruangan | Ready to use |  | Amna
19000464 | 2025-04-17 | RS JIH Surakarta | Mesin Anasthesia | Draeger/Primus | ASMA-0146 | Maintenance | Maintenance kunjungan kontrak service | Ready to use |  | Nanda
19000463 | 2025-04-17 | RS JIH Surakarta | Mesin Anasthesia | Draeger/ATLAN A350 | ASMM0125 | Maintenance | Kunjungan maintenance kontrak service | Ready to use |  | Nanda
6000823 | 2025-04-17 | RSUD Brebes | Ventilator | Draeger/Evita V800 | ASPF0164 | Survey | Neonatal flow sensor fail | Pending | Disarankan penggantian board neonatal | Dwi
2000341 | 2025-04-16 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Penggantian EMH | Ready to use |  | Nasrul
4000506 | 2025-04-16 | FKG Universitas Trisakti | Dental Unit | Kavo/E30TM | 10111149 | Service | Perbaikan foot control, penggantian rubber gasket dan dampening device | Ready to use |  | Farhat, Ipan
4000505 | 2025-04-16 | FKG Universitas Trisakti | Dental Unit | Kavo/E30TM | 10111155 | Survey | Kerusakan pada foot control, selang angin | Ready to use |  | Farhat, Ipan
17000213 | 2025-04-16 | RSUD Dr. Drajat Prawiranegara | ESWL | Medispec/EM1000 | 0617048 | Service | Membran emh bocor, penggantian membran head baru | Ready to use |  | Bahar, Satria
15000287 | 2025-04-16 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Monitor bergaris; Bersihkan card BNC | Ready to use |  | Sutan
8000667 | 2025-04-16 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHA0113 | Survey | System failure 0769.32927. Reload software dan hit software | Ready to use |  | Amna
6000822 | 2025-04-16 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASHA0095 | Service | Leak test failure. Penggantian basis board, board m.13, pvt monitoring, selenoid valve, valve 5.1 | Ready to use |  | Dwi
9000565 | 2025-04-16 | RSUD Weda | UPS | Yoshiga/GPH-33 | 562015005780L5200031 | Survey | Unit diaktifkan untuk digunakan | Ready to use |  | Ruslan
19000465 | 2025-04-15 | RSUD Budi Rahayu Magelang | Ventilator | Draeger/Evita V600 | ASTD-0042 | Maintenance | Maintenance unit warranty | Ready to use |  | Nanda
9000567 | 2025-04-15 | RSD Tidore | Baby Incubator | Draeger/Isolette 8000 | KY07192 | Survey | motor failed, hinge access rusak. Cek motor kipas. Perlu penggantian service set 3 tahun dan beberapa hinge access | Pending |  | Ruslan
9000567 | 2025-04-15 | RSD Tidore | Baby Incubator | Draeger/Isolette 8000 | KY07193 | Survey | motor failed, hinge access rusak. Cek motor kipas. Perlu penggantian service set 3 tahun dan beberapa hinge access | Pending |  | Ruslan
18000594 | 2025-04-15 | RSUD Toto Kabila | Ventilator | Draeger/Babylog VN500 | ASMK0031 | Survey | Unit restart sendiri; Kerusakan pada board M16 | Pending | Disarankan penggantian board M16 | Irfan R, Marshall
4000504 | 2025-04-15 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Survey | Applikator kurang cairan dan roda habis. Sudah diisi air applicatornya | Ready to use |  | Farhat, Ipan
17000212 | 2025-04-15 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Survey | Tidak bisa shock dan tidak ada muncul error; HVG rusak dan selang air rapuh | Pending |  | Bahar, Satria
8000651 | 2025-04-15 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina300 | ASMN0162 | Survey | Breathing circuit check fail;. Pengecekan expiratory block, expiratory valve, inspiration block; Inspiration block rusak | Pending | Perlu penggantian inspiration block | Amna
12000206 | 2025-04-15 | RSUD Toto Kabila | X-ray Digital | Apelem/Camargue DR | IMA703943 | Survey | Blue screen pada console dan software sudah lemot. Pengecekan baterai bios dan penghapusan data pasien | Pending | Baterai bios sudah harus diganti | Marshall, Irfan
19000462 | 2025-04-14 | RSUD Dr. Moewardi Surakarta | USG | Mindray/Mg | CC-1040022 | Survey | Cek alat, selftest, keyboard error, dan cleaning | Pending | Disarankan penggantian control panel assy | Nanda
18000588 | 2025-04-14 | RSUD Dr. Iwan Boking | Ventilator | Draeger/Savina 300 | ASTM0169 | Instalasi | Unit dapat dinyalakan dengan baik | Ready to use |  | Irfan R, Marshall
18000588 | 2025-04-14 | RSUD Dr. Iwan Boking | Ventilator | Draeger/Savina 300 | ASTM0168 | Instalasi | Unit dapat dinyalakan dengan baik |  |  | Irfan R, Marshall
18000588 | 2025-04-14 | RSUD Dr. Iwan Boking | Ventilator | Draeger/Savina 300 | ASTM0170 | Instalasi | Unit dapat dinyalakan dengan baik |  |  | Irfan R, Marshall
8000654 | 2025-04-14 | RSUD Sunan Kalijaga | Ventilator | Draeger/Evita V600 | ASNK0368 | Survey | Packing demo unit |  |  | Amna
7000680 | 2025-04-14 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN800 | ASTM0051 | Service | Leakage patient 100%. System check, cek exp valve, cek insp valve, cek flow sensor. Penggantian flow sensor | Ready to use |  | Wahyu
8000654 | 2025-04-14 | RSUD Sunan Kalijaga | Ventilator | Draeger/Savina300 | ASMJ0025 | Survey | Packing demo unit |  |  | Amna
8000654 | 2025-04-14 | RSUD Sunan Kalijaga | Ventilator | Draeger/Babylog VN800 | ASNM0024 | Survey | Packing demo unit |  |  | Amna
6000828 | 2025-04-14 | RS Raffa Medika | Kompressor | Xentonic | XIMC0001 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
6000828 | 2025-04-14 | RS Raffa Medika | Kompressor | Xentonic | XIMC0002 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Dwi
12000203 | 2025-04-12 | RSUD Toto Kabila | Dental Unit | Kavo/Primus 1058 | 10205320 | Survey | Waterblock, board unit, ext way board dan board dompis rusak | Pending |  | Marshall, Irfan
19000461 | 2025-04-11 | RS St Elizabeth Semarang | Mesin Anasthesia | Draeger/Fabius Tiro | ASHA-0114 | Survey | Cek unit dan cek fisik o2 sensor | Ready to use | Unit bisa digunakan tanpa monitoring o2, disarankan penggantian o2 sensor dan housing o2 sensor | Nanda
19000460 | 2025-04-11 | RS St Elizabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHC-0076 | Survey | Cek unit dan cek fisik o2 sensor | Ready to use | Unit bisa digunakan tanpa monitoring o2, disarankan penggantian o2 sensor dan housing o2 sensor | Nanda
19000459 | 2025-04-11 | RS St Elizabeth Semarang | Mesin Anasthesia | Draeger/Primus | ASKE-0213 | Maintenance | Maintenance unit Anasthesia primus | Ready to use |  | Nanda
4000502 | 2025-04-11 | FKG Universitas Trisakti | Dental Unit | Kavo/E30TM | 10111155 | Survey | Pengecekan kerusakan foot control | Pending |  | Farhat, Ipan
10000230 | 2025-04-11 | RSUD H. Boejasin | MOT | SYAS | - | Service | HVAC tidak menyala dan ruangan tidak dingin; Setting HVAC menjadi positive | Ready to use |  | Kreshna
8000659 | 2025-04-11 | RSUD Kraton Pekalongan | Lampu Operasi | Draeger/Polaris 600 | ASDK0401 | Service | Lampu mati; Pemasangan power supply | Ready to use |  | Amna
12000205 | 2025-04-11 | RSUD Ainun Habibie | MOT | SYAS | - | Survey | Alarm gas module menyala, tekanan low. Pengecekan kompresor dari RS; Kompresor gas medik salah satu rusak. | Pending | Disarankan perbaiki atau melakukan penggantian kompresor | Marshall, Irfan
6000820 | 2025-04-11 | RS Orthopaedi Purwokerto | Ventilator | Draeger/Savina300 | ASTC0258 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Dwi
19000457 | 2025-04-10 | RSUP Dr Kariadi Semarang | Ventilator | Draeger/Savina 300 | ASMK-0304 | Maintenance | Pasang maintenance kit 2 year | Ready to use |  | Nanda
9000566 | 2025-04-10 | RSUD Labuha | UPS | Yoshiga/GPI-33 | E2211CO12260621353 | Service | Bypass control board rusak. Penggantian part bypass control board | Ready to use |  | Ruslan
8000658 | 2025-04-10 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASHL0063 | Survey | Gas dan ventilation failed alarm. Check software, cek blok gas modul | Pending | Unit masih belum berfungsi. Akan dijadwalkan pengecekan lebih lanjut | Amna
19000456 | 2025-04-10 | RSUP Dr. Kariadi Semarang | Ventilator | Draeges/Savina 300 | ASMK-0300 | Maintenance | Pasang maintenance kit 2 year | Ready to use |  | Nanda
6000815 | 2025-04-10 | RS Nirmala | Ventilator | Draeger/Savina300 | ASRE0307 | Survey | Flow sensor fail; Perlu penggantian O2 cell dan flow sensor | Ready to use |  | Dwi
3000374 | 2025-04-10 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 0720015 | Survey | Hardware not connected; Kerusakan pada power supply 24V | Pending |  | Imam
19000458 | 2025-04-09 | RSUD Dr. Moewardi Surakarta | USG | Mindray/MG | CCI-04002265 | Survey | Cek unit dan restart unit | Ready to use |  | Nanda
12000207 | 2025-04-09 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Survey | Hasil foto hilang. Restart hasil foto. Foto sudah kembali | Ready to use |  | Marshall
4000492 | 2025-04-09 | RSUP Dr. M. Djamil Padang | Laser Surgical | Jena Surgical/Multipulse TM1470 | J190100147 | Maintenance | Preventive Maintenance | Ready to use |  | Satria
7000678 | 2025-04-09 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASJN0023 | Maintenance | Total care Maintenance | Ready to use |  | Wahyu
1000228 | 2025-04-09 | RSUD Mas Amsyar | MOT | SYAS | - | Service | Kontaktor panel HVAC dengan PLN terbakar; Penggantian kontraktor | Ready to use |  | Kreshna
15000286 | 2025-04-09 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Data base penuh. Hapus data | Ready to use |  | Sutan
8000657 | 2025-04-09 | RSUD RA Kartini | Lampu Operasi | Draeger/Polaris 600 | ASLL0237 | Survey | Survey untuk pemindahan unit ke ruangan baru | Ready to use |  | Amna
6000814 | 2025-04-09 | RSUD Cilacap | Bubble CPAP | Draeger/Oxymixer | QM20230175 | Service | Flow meter low dan high pecah dan tidak bisa diatur; Penggantian flow meter high and low | Ready to use | Claim Warranty | Dwi
11000395 | 2025-04-09 | RSUD Haji Makassar | Ventilator | Draeger/Evita V300 | ASPH0056 | Survey | Internal baterai drop | Pending | Perlu dilakukan penggantian baterai internal | Nasrul
3000373 | 2025-04-09 | RSUP Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0146 | Service | Worklist tidak jalan, koneksi PACS sering lost connect, disable given name dan cke koneksi kabel LAN | Ready to use |  | Imam
19000467 | 2025-04-09 | RSUP Dr Kariadi | Ventilator | Draeger/Savina 300 | ASMK-0300 | Maintenance | Pasang maintenance kit 2 year | Ready to use |  | Nanda
19000467 | 2025-04-09 | RSUP Dr Kariadi | Ventilator | Draeger/Savina 300 | ASMK-0304 | Maintenance | Pasang maintenance kit 2 year | Ready to use |  | Nanda
20000338 | 2025-04-08 | RSUP Tadjudin Chalid | UPS | Yoshiga/10 KVA |  | Service | Pengecekan unit baterai 16 pcs | Pending | Perlu penggantian baterai UPS | Sujana
20000337 | 2025-04-08 | RSUP Tadjudin Chalid | ESWL | Medispec/EM1000 | 0922014 | Survey | Pengecekan unit EMH, cleaning water dan circulation | ready to use | Perlu penggantian selang circulation | Sujana
18000593 | 2025-04-08 | RSD Tidore | Ventilator | Draeger/Babylog VN500 | ASJK0042 | Survey | Unit tidak bisa dikalibrasi. Device check; Valve O2 rusak | Pending |  | Irfan
6000827 | 2025-04-08 | RSU An Ni'mah | Bubble CPAP | Draeger/Oxymixer | QM20215459 | Service | Alarm over pressure oxygen; Adjustment pressure O2 central regulator | Ready to use |  | Dwi
12000201 | 2025-03-27 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Helianthus DBT | HELI DBT/130/C2 | Survey | Restart hasil foto | Ready to use |  | Marsall
6000813 | 2025-03-27 | RSUD Ajibarang | CPAP | Draeger/Oxymixer | QM20215544 | Survey | Installasi unit demo | Ready to use |  | Dwi
8000648 | 2025-03-27 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0033 | Survey | Power fail. General check; Baterai rusak | Pending |  | Amna
2000336 | 2025-03-27 | RSUP Makassar | C-arm | Apelem/Evo R+ | 551240241086 | Service | Error missing pulse. Pengecekan board memory dan board inverter control dan inverter board rusak | Pending |  | Sujana, Nasrul
6000813 | 2025-03-27 | RSUD Ajibarang | CPAP | Draeger/Oxymixer | QM20215544 | Survey | Installasi unit demo | Ready to use |  | Dwi
6000819 | 2025-03-26 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASKJ0050 | Service | Pemindahan unit ke ruangan baru | Ready to use |  | Dwi
19000455 | 2025-03-26 | RSUD Bung Karno | USG | Mindray/Resona 7 | 7V-OB002253/MFG.2020-11-07 | Maintenance | Kunjungan dan survey populasi USG Mindray | Ready to use | Disarankan kontrak service | Nanda
16000281 | 2025-03-26 | RSUD Kab. Bekasi | Bed Pediatric | SYAS/2cranck | 240526302000400043 | Service | Penggantian handle bed | Ready to use |  | Ipan, Farhat
8000655 | 2025-03-26 | RS Mitra Bangsa Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0041 | Survey | Alarm O2 sensor fail dan inspiration O2 flow alarm. System check; O2 sensor kadaluarsa | Pending |  | Amna
12000250 | 2025-03-25 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0004 | Survey | Cek sensor PDMIX dan board optimix (rusak) | Ready to use | Unit sementara sudah bisa digunakan tapi dengan menggunakan part pinjaman dari draeger | Marshall, Ruslan
18000587 | 2025-03-25 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | ASNL0151 | Survey | Tidak bisa kalibrasi unit karna tidak ada flow sensor | Pending |  | Irfan R
18000586 | 2025-03-25 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | 8417801 | Survey | Nebulizer tidak berfungsi, pressure regulator rusak | Pending |  | Irfan R
6000816 | 2025-03-25 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASKJ0050 | Service | Pemindahan unit ke ruangan baru | Ready to use |  | Dwi
4000493 | 2025-03-25 | RS Otak Moh. Hatta | ESWL | Medispec/EM1000 | 0922163 | Survey | Refresh training | Ready to use |  | Satria
7000677 | 2025-03-25 | RS PKU Muhammadiyah Gamping | Ventilator | Draefer/Evita V300 | ASPC0205 | Survey | Gas supply sensor calibration failed. Cek error log, pembacaan air proportional valve tidak stabil; Proportional valve rusak | Pending |  | Wahyu
19000454 | 2025-03-25 | RSUP Dr. Kariadi | Patient Monitor | Draeger/Vista 120 S | S4SNH0664 | Survey | Cek unit, cek settingan, setting and adjusment | Ready to use |  | Nanda
6000812 | 2025-03-25 | RSUD dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Atlan A350 | ASRL0046 | Survey | Breathing system leakage. Pembersihan co2y breath, ganti heating system, bersihkan valve, penggantian soda lime | Ready to use |  | Dwi
6000811 | 2025-03-24 | RSUD Bumiayu | Infant Warmer | Draeger/Resuscitaire | CY03650 | Preventive | Preventive Maintenance | Ready to use |  | Dwi
6000811 | 2025-03-24 | RSUD Bumiayu | Incubator Transport | Draeger/Isolette TI500 | AY11232 | Preventive | Preventive Maintenance | Ready to use |  | Dwi
6000811 | 2025-03-24 | RSUD Bumiayu | Infant Warmer | Draeger/Babyleo TN500 | ASLJ0005 | Preventive | Preventive Maintenance | Ready to use |  | Dwi
6000810 | 2025-03-24 | RSUD Bumiayu | Infant Warmer | Draeger/Resuscitaire | DSPF0026 | Survey | Unit error 21. Pengecekan keseluruhan unit, temp prob. Control board bermasalah | Pending | Disarankan penggantian control board PN:Mu27145 | Dwi
8000653 | 2025-03-24 | RSUD Sunan Kalijaga | Ventilator | Draeger/Savina300 | ASTJ0006 | Survey | Training user demo alat | Ready to use |  | Amna
3000372 | 2025-03-24 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Lifter Et130, unplug and plug DRV 1 | Ready to use |  | Imam
8000656 | 2025-03-21 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Primus | ASNE0161 | Service | Pemasangan O2 sensor dan flow sensor | Ready to use |  | Amna
8000656 | 2025-03-21 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Primus | ASKL0011 | Service | Pemasangan O2 sensor dan flow sensor | Ready to use |  | Amna
8000656 | 2025-03-21 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Primus | ASKL0120 | Service | Pemasangan O2 sensor dan flow sensor | Ready to use |  | Amna
8000656 | 2025-03-21 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Primus | ASEL0138 | Service | Pemasangan O2 sensor dan flow sensor | Ready to use |  | Amna
19000453 | 2025-03-21 | RS JIH Surakarta | Ventilator | Draeger/Baby log VN 600 | ASSK-0027 | Maintenance | Maintenance Unit warranty | Ready to use |  | Nanda
4000491 | 2025-03-21 | RS Otak Moh. Hatta | ESWL | Medispec/EM1000 | 922163 | Service | Aluminium emh pecah; Penggantian EMH dan pendampingan pasien | Ready to use |  | Satria
6000809 | 2025-03-20 | RS Mata Purwokerto | Mesin Anasthesia | Draeger/Fabius Plus | ASTH0001 | Maintenance | Preventive Maintenance | Ready to use |  | Dwi
11000392 | 2025-03-20 | RSUP Tadjudin Chalid | Mesin Anasthesia | Draeger/Primus | ASJF0637 | Survey | Ventilation valve fail. Electrovalve MV3 dan MV2 rusak | Pending |  | Nasrul
11000393 | 2025-03-20 | RSUP Tadjudin Chalid | Mesin Anasthesia | Draeger/Primus | ASJL0023 | Survey | O2 sensor fail, gas sensor fail; O2 sensor, ILCA3 rusak, baterai internal rusak dan flow power rusak | Pending |  | Nasrul
11000394 | 2025-03-20 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina300 | ASJL0396 | Survey | Baterai internal drop, flow sensor tidak terbaca. | Pending | Disarankan penggantian service set 2 tahun dan 3 tahun | Nasrul
8000647 | 2025-03-20 | RSUD Samin Surosentiko | USG | Mindray/Consona N9 | LQ6-32000376 | Survey | Installasi unit demo | Ready to use |  | Amna
6000809 | 2025-03-20 | RS Mata Purwokerto | Mesin Anasthesia | Draeger/Fabius Plus | ASTH0001 | Maintenance | Total care Maintenance | Ready to use |  | Dwi
10000225 | 2025-03-20 | RSUD Mas Amsyar | Pendant | SYAS | SY-PS2402004 | Service | Pemasangan gas kompresor | Ready to use |  | Kreshna
19000452 | 2025-03-19 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL-0090 | Service | Cek alat dan tarik error | pending | Ada kerusakan pada blower, Disarankan Penggantian blower | Nanda
8000646 | 2025-03-19 | RSUD Sunan Kalijaga | Ventilator | Draeger/Evita V600 | ASNK0368 | Survey | Install demo dan training | Ready to use |  | Amna
3000371 | 2025-03-18 | RSUD Dr. Drajat Prawiranegara | ESWL | Medispec/EM1000 | 1118085 | Survey | Pengetesan board c-arm eswl | Ready to use |  | Imam
2000340 | 2025-03-17 | RSUP Makassar | Mobile X-ray | Apelem/Rafale BDR | 97124084483 | Service | Koneksikan PACS | Ready to use |  | Sujana, Nasrul
12000199 | 2025-03-14 | RSUD Hasri Ainun Habibie | MOT | Syastech | - |  | Training MOT Syastech, pencucian filter kondensor, outdoor, indoor, dan setting PLC |  | Training Sudah selesai | Marshall
19000451 | 2025-03-14 | RST Bhakti Wira Tamtama | Bed Stretcher | Syastech/bed stretcher | 250368302000400021 | Instalasi | Unit berfungsi dengan baik | Ready to use
6000808 | 2025-03-14 | RSU Muhammadiyah Siti Aminah | Mesin Anasthesia | Draeger/Fabius Plus | ASND0060 | Survey | O2 sensor fail, flow sensor calibration. Self test dan kalibrasi O2 sensor | Ready to use | Disarankan penggantian O2 sensor, flow sensor dan service set 3 tahun | Dwi
8000644 | 2025-03-14 | RS Budi Agung | Mesin Anasthesia | Draeger/Fabius Plus | ASTK0003 | Maintenance | General check, system check, running test | Ready to use |  | Amna
6000808 | 2025-03-14 | RSU Muhammadiyah Siti Aminah | Mesin Anasthesia | Draeger/Fabius Plus | ASND0060 | Survey | O2 sensor fail; O2 sensor expired, flow sensor dan service set 3 tahun perlu diganti | Ready to use |  | Dwi
7000676 | 2025-03-14 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Maintenance | Total care Maintenance | Ready to use |  | Wahyu
7000675 | 2025-03-14 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Maintenance | Total care Maintenance | Ready to use |  | Wahyu
9000561 | 2025-03-14 | RSUD Dr. H. Chasan Boesoirie | Mesin Anasthesia | Draeger/Perseus A500 | ASSK0059 | Survey | Masalah teknis yang menyebabkan mode ventilasi low oksigen. Pengembalian saklar O2 ke posisi O2 | Ready to use |  | Ruslan
15000285 | 2025-03-14 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Survey | Alat mati total. Cek generator, alat menyala. Touchscreen LCD rusak | Pending | Sudah diberikan penawaran harga | Sutan
7000674 | 2025-03-13 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Service | Installasi unit untuk backup | Ready to use |  | Wahyu
11000391 | 2025-03-13 | RSI Faisal Makassar | ESWL | Medispec/EM1000 | 1214012 | Survey | EMH expired, sirkulasi bocor | Pending | Disarankan penggantian EMH | Nasrul
8000645 | 2025-03-13 | RSUD Sunan Kalijaga | Ventilator | Draeger/Evita V600 | ASNK0368 | Survey | Installasi demo unit | Ready to use |  | Amna
4000490 | 2025-03-13 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 0720015 | Survey | Motor bunyi keras lama | Pending | Perlu penggantian power supply | Imam, Satria
10000229 | 2025-03-13 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D18680271 | Service | Kolimator tidak terbuka; Restart unit dan perbaiki koneksi kabel | Ready to use |  | Kreshna
9000560 | 2025-03-13 | RSUD Jailolo | Ventilator | Draeger/Evita V300 | ASND0085 | Survey | humidifier mati total | Ready to use | Unit dalam kondisi normal namun humidifier disarankan diganti | Ruslan
9000560 | 2025-03-13 | RSUD Jailolo | Ventilator | Draeger/Evita V300 | ASND0079 | Survey | humidifier mati total | Ready to use | Unit dalam kondisi normal namun humidifier disarankan diganti | Ruslan
9000560 | 2025-03-13 | RSUD Jailolo | Ventilator | Draeger/Evita V300 | ASND0080 | Survey | humidifier mati total | Ready to use | Unit dalam kondisi normal namun humidifier disarankan diganti | Ruslan
6000807 | 2025-03-12 | RSU An Ni'mah | Ventilator | Draeger/Babylog VN600 | ASSK0026 | Service | Air supply error. Cek input supply air, cek proportional valve, cek system ventilator; Air kompressor masuk ke system | Ready to use |  | Dwi
9000564 | 2025-03-12 | RSUD Tobelo | X-ray Digital | Apelem/Camargue DR | IMA703799 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Ruslan
9000563 | 2025-03-12 | RSUD Tobelo | Ventilator | Draeger/Babylog VN800 | ASTF0020 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Ruslan
4000489 | 2025-03-12 | RSUD Sidoarjo Barat | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0755 | Service | SID tidak terbaca, adjusting potensio sid | Ready to use |  | Imam, Satria
7000673 | 2025-03-12 | RSUD Panembahan Senopati Bantul | Patient Monitor | Draeger/Vista 120 | V2SHK0117 | Survey | Layar blackout; Mainboard rusak | Pending |  | Wahyu
16000280 | 2025-03-12 | RSUD Kab. Bekasi | Bed Pediatric | SYAS | 240526302000400043 | Survey | Handle rel bed patah | Pending | Perlu penggantian handle plastik dan penyangga rel | Farhat, Ipan
9000558 | 2025-03-12 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0011 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Ruslan
9000558 | 2025-03-12 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0012 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Ruslan
8000643 | 2025-03-12 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASJE0014 | Survey | System check, gas sensor fail; Kerusakan pada pressure regulator dan mix air | Pending |  | Amna
19000900 | 2025-03-12 | RS Ken Saras | Mesin Anasthesia | Draeger/Fabius GS Premium | ASCL0014 | Survey | Layar LCD mati; Kabel soket rusak | Pending | Perlu penggantian kabel soket mainboard to LCD | Nanda
12000200 | 2025-03-11 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Helianthus DBT/Drypix Smart | Heli DBT/130/C2 | Service | Daily check, tes expose, dan mengembalikan posisi film | Ready to use |  | Marshall
19000898 | 2025-03-11 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Isolette8000 | FC09049 | Survey | Cek water supply, cek humidity cek heater; Evaporizer humidity rusak | Pending | Perlu diganti evaporizer humidity | Nanda
19000899 | 2025-03-11 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Isolette8000 | MC09024 | Service | Cek water supply, heater kotor. Bersihkan heater | Ready to use |  | Nanda
4000501 | 2025-03-11 | Signature Dental Care | Dental Unit | Kavo/Primus 1058 | 10207191 | Service | Installasi software intraoral kavo | Ready to use |  | Farhat, Ipan
6000806 | 2025-03-11 | RSU Wiradadi Husada | Ventilator | Draeger/Evita V300 | ASNL0238 | Service | Device failure (gas supply sensor, O2 sensor fail; Penggantian proportional valve | Ready to use |  | Dwi
3000370 | 2025-03-11 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Imam, Bahar
7000672 | 2025-03-11 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASJH0126 | Survey | VT tidak stabil, nebulizer fail; Valve drive, pressure regulator dan internal baterai rusak | Pending |  | Wahyu
6000806 | 2025-03-11 | RSU Wiradadi Husada | Ventilator | Draeger/Evita V300 | ASNL0238 | Service | Device failure (gas supply sensor, O2 sensor fail; Penggantian proportional valve | Ready to use |  | Dwi
2000335 | 2025-03-10 | RSKD Gigi dan Mulut | Dental Unit | Kavo/Primus 1058 | 10207187 | Survey | Pengecekan software | Pending | Disarankan penggantian board dentist | Sujana
8000642 | 2025-03-10 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V800 | ASPC0426 | Survey | Installasi dan training demo unit | Ready to use |  | Amna
19000894 | 2025-03-10 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Primus | ASHM0012 | Service | Absolute pressure sensor fail; Penggantian pressure sensor | Ready to use |  | Nanda
6000805 | 2025-03-10 | RSUD Ajibarang | USG | Mindray/DC30 | GP18009329 | Survey | Baterai communication error | Pending | Perlu penggantian baterai kit | Dwi
6000900 | 2025-03-08 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASLD0042 | Service | Breathing system leakage. Penggantian 2/2 distribution valve | Ready to use |  | Dwi
7000671 | 2025-03-10 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0092 | Maintenance | Total care Maintenance | Ready to use |  | Wahyu
18000591 | 2025-03-07 | RSUD Samratulangi Tondano | Ventilator | Draeger/Savina300 | ASJM0392 | Survey | O2 rendah, pressure regulator rusak | Pending |  | Irfan R
2000334 | 2025-03-07 | RSUP Tadjudin Chalid | ESWL | Medispec/EM1000 | 0922014 | Service | Penggantian EMH | Ready to use |  | Sujana, Nasrul
19000893 | 2025-03-07 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0150 | Survey | Alarm sytem failure. Alarm karna loudspeaker kanan rusak | Pending | Disarankan penggantian loudspeaker 1 set | Nanda
10000226 | 2025-03-06 | RSUD H. Boejasin | MOT | SYAS | - | Service | Pintu sliding door tidak tertutup; Pemasangan baut pada penyangga pintu | Ready to use |  | Kreshna
7000670 | 2025-03-06 | RSUD Nyi Ageng Serang (KSO) | CT Scan | Siemens/Somatom | - | Service | Penggantian baterai UPS | Pending | Baterai tidak terbaca | Wahyu
15000284 | 2025-03-06 | RS Otak Moh. Hatta | ESWL | Medispec/EM1000 | 0922163 | Survey | Cek unit; Membran rusak | Pending | Akan dijadwalkan kembali | Sutan
19000892 | 2025-03-06 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Survey | Monoblock error, sudah dicoba perbaiki per komponen namun masih belum bisa | Pending | Disarankan penggantian monoblock | Nanda
3000369 | 2025-03-06 | RSUD Pagelaran | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/136/C2 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Imam
15000283 | 2025-03-05 | RSUP Prof. Dr. M. Ali Hanafiah | X-ray Digital | Apelem/Camargue | - | Survey | Survey ruangan | Ready to use |  | Sutan
4000485 | 2025-03-05 | RSUP Surabaya | C-arm | Apelem/Evo R+ | 551240231085 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Satria, Bahar
4000485 | 2025-03-05 | RSUP Surabaya | C-arm | Apelem/Evo R+ | 551240251087 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Satria, Bahar
8000641 | 2025-03-05 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHL0168 | Service | Baterai fail setelah penggantian baterai; Reset battery | Ready to use |  | Amna
4000487 | 2025-03-05 | RSUP Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0146 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Satria, Farhat
6000803 | 2025-03-05 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | AF32541 | Survey | Keypad error; Membran keypad rusak | Pending |  | Dwi
17000211 | 2025-03-05 | Siloam Hospital Lippo Village | C-arm | Ziehm 800 | 81092 | Service | Uji Kesesuaian | Ready to use |  | Bahar
6000804 | 2025-03-05 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASHL0166 | Survey | Neonatal flow sensor fail | Pending | Rekomendasi penggantian baterai internal dan external serta flow sensor | Dwi
3000368 | 2025-03-05 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Apelem | H226F0745 | Maintenance | Preventive Maintenance. Penggantian kabel LAN DRV | Ready to use |  | Imam
4000486 | 2025-03-04 | RSUP Surabaya | Lampu Operasi | SYAS/Lumina | OL9680/802401000045 | Service | Lengan lemah dan stepper patah; Penggantian lengan lampu | Ready to use |  | Satria, Farhat
4000486 | 2025-03-04 | RSUP Surabaya | Lampu Operasi | SYAS/Lumina | OL9680/802401000044 | Service | Lengan lemah dan stepper patah; Penggantian lengan lampu | Ready to use |  | Satria, Farhat
7000669 | 2025-03-04 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Survey | Layar blank, penggantian control unit yang baru namun HV enable dan terdapat kabel yang mengelupas | Pending | Unit dikirim ke HO | Wahyu
12000196 | 2025-03-03 | RSUD Liun Kendage Tahuna | Laparascopy | Bbraun | 6582 | Survey | Pemeriksaan BPK untuk unit Laparascopy | Ready to use |  | Marshall
6000818 | 2025-03-03 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0013 | Survey | Tidak mendeteksi vaporizer. Cek koneksi vapo dan supply PLN. Recalibration | Ready to use |  | Dwi
8000634 | 2025-03-03 | RSUD RAA Soewondo Pati | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/054/C1 | Survey | Connection printer fail; Setting ulang IP Address | Ready to use |  | Amna
7000668 | 2025-03-03 | RSUD Bagas Waras | Incubator Transport | Draeger/TI500 | DY11070 | Service | Baterai tidak bisa di charge; Penggantian baterai | Ready to use |  | Wahyu
19000890 | 2025-03-03 | RSUP dr. Kariadi | Infant Warmer | Draeger/Resuscitaire | MC13338 | Survey | Error 05, heater rusak | Pending | Disarankan penggantian heater | Nanda
9000557 | 2025-03-03 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | 2510220561 | Service | Wall stand miring sehingga hasil gambar terpotong; Setting posisi wall stand | Ready to use |  | Ruslan
19000889 | 2025-03-01 | RSUP dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0235 | Survey | Gas measurement failure. Flow power MPGM rusak | Pending |  | Nanda
11000396 | 2025-02-28 | RSUP Dr. Wahidin Sudirohusodo | Syringe Pump | Bbraun/Perfusor Space | 716608 | Instalasi | Unit berfungsi dengan baik | Ready to use | Terinstall 102 unit | Nasrul
8000638 | 2025-02-28 | RSJD Dr. Amino Gondohutomo | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0145 | Survey | LCD fail; Plug and play cable flexible | Pending | Rekomendasi pemasangan UPS external | Amna
19000888 | 2025-02-28 | RS Indriati Solo | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0063 | Survey | O2 sensor fail; O2 sensor kadaluarsa | Ready to use |  | Nanda
8000639 | 2025-02-28 | RSUD K.R.M.T Wongsonegoro | USG | Mindray/DC-30 | 9P-18009331 | Survey | Unit berfungsi dengan baik | Ready to use |  | Amna
7000698 | 2025-02-28 | RSUP dr. Soeradji Tirtonegoro | Infant Warmer | Draeger/Babytherm 8010 | ASMM0045 | Service | Error 35, error heater; Diganti dengan heater milik RS | Ready to use |  | Wahyu
8000635 | 2025-02-27 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0148 | Survey | LCD fail. Cek koneksi LCD cable | Ready to use |  | Amna
6000817 | 2025-02-27 | RSUD Cilacap | Mesin Anasthesia | Draeger/Perseus A500 | ASNB01234 | Survey | Installasi unit demo | Ready to use |  | Dwi
19000897 | 2025-02-27 | RSIA Umi Barokah Boyolali | Baby Incubator | Draeger/Isolette C2000 | DSRL0029 | Survey | Fan failure; Pembersihan kipas | Ready to use |  | Nanda
4000484 | 2025-02-27 | RSUD Kanjuruhan | Lampu Operasi | Draeger/Polaris 600 | ASTL1105 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bahar
9000555 | 2025-02-26 | RSUD Sanana | X-ray Digital | Apelem/Camargue DR | 251023012 | Preventive | Preventive maintenance warranty | Ready to use |  | Ruslan
7000695 | 2025-02-27 | RSUD KOTA YOGYAKARTA | Ventilator | Draeger/Babylog VN800 | ASRK0005 | Survey | Pembacaan O2 error; Kalibrasi sensor manual | Ready to use |  | Wahyu
8000631 | 2025-02-26 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASHL0063 | Service | Gas mixer failure, general check dan penggantian water trap | Ready to use |  | Amna
18000578 | 2025-02-26 | RSUP Prof. Dr. R. D. Kandou Manado | Infant Warmer | Draeger/Resuscitaire | CY03657 | Survey | Pemanas tidak berfungsi, heater rusak | Pending |  | Irfan R
19000886 | 2025-02-26 | RST dr. Soedjono Magelang | Baby Incubator | Draeger/Isolette C2000 | DSPM0035 | Survey | Pengecekan unit | Ready to use |  | Nanda
19000887 | 2025-02-26 | RST dr. Soedjono Magelang | Infant Warmer | Draeger/Resuscitaire | DY10960 | Survey | Pengecekan SSR board | Ready to use |  | Nanda
15000282 | 2025-02-25 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Service | Unit tidak hidup; Pembersihan board dengan contact cleaner | Ready to use |  | Sutan
12000191 | 2025-02-25 | RSD Liun Kendage Tahuna | Laparascopy | B-Braun | 6582 | Service | Pemasangan trolley laparascopy selesai | Ready to use |  | Marshall
12000192 | 2025-02-25 | RSD Liun Kendage Tahuna | Laparascopy | B-Braun | 1008CE189 | Survey | Terjadi permasalahan pada selenoid valve | Pending | Unit akan dibawa ke workshop Bbraun untuk dilakukan pengecekan | Marshall
8000633 | 2025-02-25 | RSUD RA Kartini | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/006/C0 | Survey | time out high speed starter. HSS board rusak | Pending |  | Amna
3000367 | 2025-02-25 | RSUD RA Kartini | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/006/C0 | Survey | Timeout high speed starter; HSS board rusak | Pending |  | Imam, Amna
19000884 | 2025-02-25 | RSUP dr. Kariadi | Baby Incubator | Draeger/Isolette C2000 | CR35569 | Service | Pemasangan front panel | Ready to use |  | Nanda
19000885 | 2025-02-25 | RSUP dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0235 | Service | Sample line tidak terbaca; Bongkar selang sample line | Ready to use | Disarankan penggantian service set 6 tahun | Nanda
20000333 | 2025-02-25 | RSUP Jayapura | Instrument | Aesculap/Amputation set | - | Instalasi | Instrument set lengkap | Ready to use |  | Sujana
7000666 | 2025-02-25 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASJN0023 | Servoce | Pressure sensor calibration failed; Bersihkan inspiration dan expiration valve | Ready to use |  | Wahyu
6000816 | 2025-02-25 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASHL0010 | Survey | Pengecekan after service (flow sensor, baterai, breathing system). Unit OK | Ready to use |  | Dwi
4000488 | 2025-02-24 | RSUP Surabaya | Lampu Operasi | SYAS/Lumina | OL9680/80XVWZ00032 | Service | Unit terlalu rendah; Adjust ketinggian lengan | Ready to use |  | Satria, Bahar
19000883 | 2025-02-24 | RSUP dr. Kariadi | Baby Incubator | Draeger/Isolette 8000 | MR07739 | Survey | check water supply, humidity low; Evaporizer dan sensor modul error | Pending |  | Nanda
15000281 | 2025-02-24 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Tube tidak bisa bergerak; Bersihkan sensor grid | Ready to use |  | Nanda
8000637 | 2025-02-21 | RS Aisyiyah Kudus | Ventilator | Draeger/Babylog VN500 | ASPF0017 | Survey | Refresh training | Ready to use |  | Amna
20000332 | 2025-02-21 | RSUP Dr. Wahidin Sudirohusodo | Dental Unit | Kavo/1058TM | 10205313 | Survey | Alarm alat berbunyi, pengecekan PCBA dentist | Pending |  | Sujana
11000389 | 2025-02-21 | RSUD Sawerigading | Mesin Anasthesia | Draeger/Perseus A500 | ASHJ0071 | Survey | Gas measurement failure, ETCo2 tidak terbaca, O2 tidak terbaca. Pump dan sensor ILCA2 rusak | Pending |  | Nasrul
9000553 | 2025-02-20 | RSD Tidore | Mesin Anasthesia | Draeger/Perseus A500 | ASJK0123 | General Check | Pengecekan kondisi unit | Pending | Pergantian part baterai akan melakukan service kontrak diperlukan ketika unit sudah beroperasi selama 4 tahun | Ruslan
7000665 | 2025-02-21 | RSPAU dr. S. Hardjolukito | Infant Warmer | Draeger/Babytherm 8004 | ASDL0042 | Survey | Error 35, heater error | Pending | Perlu penggantian heater | Wahyu
9000554 | 2025-02-20 | RSD Tidore | Ventilator | Draeger/Babylog VN500 | ASJK0042 | Survey | Pengecekean kondisi dan baterai unit drop | Pending | Disarakan pergantian part regulator batterai, kit air filter dan partical filter | Ruslan
8000636 | 2025-02-20 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHM0013 | Service | Power fail dan mixer fail. Penggantian power pack, sensor pdif low, sensor pdif high | Ready to use |  | Amna
8000636 | 2025-02-20 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASJH0095 | Service | Penggantian basis board dan proportional valve | Ready to use |  | Amna
11000390 | 2025-02-20 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2165 | Survey | Pengecekan kelengkapan unit | Ready to use | Disarankan penggantian NIBP Cliff | Nasrul
3000366 | 2025-02-20 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | mA low tube; Kalibrasi filament tube | Ready to use |  | Imam
12000198 | 2025-02-19 | RSUD Toto Kabila | Laser surgical | Jena Surgical/MultiPulse HoPlus | J1911573 | Survey | Pengambilan part module, PSI dan Laser Head | Pending | Part akan dibawa ke workshop pusat untuk dilakukan perbaikan | Marshall
17000210 | 2025-02-19 | RSUP Jayapura | Digital Radiografi Fluoroskopi | Apelem/Platinum | H248L0157 | Survey | Uji kesesuaian dan Uji paparan | Ready to use |  | Bahar, Satria
12000189 | 2025-02-18 | RSUD Toto Kabila | Dental Unit | Kavo/1058TM | 20181010205320 | Survey | Software failed terload memori corrupt | Pending | unit tetap bisa digunakan, tapi memori tidak terinstall di internal memori. Perlu dilakukan pergantian board | Marshall
12000188 | 2025-02-18 | RSUD Toto Kabila | Laser surgical | Jena Surgical/MultiPulse HoPlus | JI9115173 | Survey | Pergantian part, rembes pada laser, ganti kapasitor. Tapi masih terdapat masalah pada unit | Pending | Unit sudah selesai diganti part dan rembesan sudah diperbaiki, tapi unit masih belum bisa menembak | Marshall
7000694 | 2025-02-18 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Survey | Layar blank, control unit rusak | Pending | Unit sudah dicoba dengan control unit yang lain namun tetap rusak. | Wahyu
8000630 | 2025-02-18 | RSUD RA Kartini | C-arm | Apelem/Evo R+ | 55120027813 | Service | Expose alarm, motif missing pulse; Pengecekan console | Ready to use |  | Amna
20000329 | 2025-02-18 | RSUD Morowali | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/126/C2 | Survey | UPS fault, pengecekan baterai; | Pending | Perlu penggantian baterai 16 pcs | Sujana
11000388 | 2025-02-18 | RSUP Dr. Wahidin Sudirohusodo | Instrument | Aesculap/Basic Vascular | - | Instalasi | Serah terima basic vascular | Ready to use |  | Nasrul
18000585 | 2025-02-18 | RS Mata Sulawesi Utara | Mesin Anasthesia | Draeger/Perseus A500 | ASJM0178 | Maintenance | Kunjungan join visit contract service | Ready to use |  | Irfan R
3000364 | 2025-02-18 | RSUD Pringsewu | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0756 | Service | Detector artefak; Penggantian detector | Ready to use |  | Imam, Sutan
18000579 | 2025-02-18 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN600 | ASJC0048 | Survey | Nebulizer tidak berfungsi, pressure regulator rusak | Pending |  | Irfan R
12000187 | 2025-02-17 | RSUD Ainun Habibie | MOT | Syas Tech | - | Survey | AC kurang dingin, pengecekan outdoor unit dan set point remote kontrol | Pending | Harus dilakukan pembersihan outdoor AC, pencucian unit outdoor | Marshall
9000547 | 2025-02-15 | RSUD Weda | Dental Unit | Kavo/1058TM | 2018-11-10205459 | Survey | Pengecekan kondisi unit, motor backrest axis 2 tidak berfungsi | Ready to use | Disarankan untuk pergantian board PCBA Chair | Ruslan
9000549 | 2025-02-14 | RSUD Weda | Mesin Anasthesia | Draeger/Primus | ASNE0150 | Survey | Pengecekan Kondisi unit | Ready to use | Direkomendasikan ganti part ( service set 3 tahun) | Ruslan
6000801 | 2025-02-14 | RSUD Ajibarang | Ventilator | Draeger/Babylog VN800 | ASMN0054 | Survey | Baterai internal error | Pending | Rekomendasi penggantian baterai internal dan service set | Dwi
19000882 | 2025-02-14 | RSUP dr. Kariadi | USG | Mindray/DC32 | 9Q8B000124 | Survey | Unit berfungsi dengan baik | Ready to use |  | Nanda
19000881 | 2025-02-14 | RSUP dr. Kariadi | USG | Mindray/Nuewa i9 Exp | KV328000015 | Survey | Unit berfungsi dengan baik | Ready to use |  | Nanda
6000802 | 2025-02-14 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0010 | Service | Penggantian baterai pack | Ready to use |  | Dwi
6000802 | 2025-02-14 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0011 | Service | Penggantian baterai pack | Ready to use |  | Dwi
20000339 | 2025-02-14 | RSUD La Palaloi | Baby Incubator | Draeger/Isolette C2000 | MC37362 | Service | Error humidity, cek evaproizer. Evaporizer perlu diganti | Pending |  | Sujana
19000880 | 2025-02-13 | RSUD Kota Salatiga | ESWL | Medispec/E3000 | 10131176 | Survey | Pengecekan unit | Ready to use |  | Nanda
18000581 | 2025-02-13 | RSUD Talaud | Ventilator | Draeger/Babylog VN500 | ASND0072 | Maintenance | Unit berfungsi dengan baik | Ready to use |  | Irfan R
17000209 | 2025-02-13 | RSUP Makassar | Mobile X-ray | Apelem/Rafale BDR | 97124084483 | Survey | Uji Kesesuaian | Ready to use |  | Bahar, Nasrul
4000481 | 2025-02-13 | RSUP Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0146 | Survey | Training user | Ready to use |  | Satria, Purwono
4000483 | 2025-02-13 | RSUP Surabaya | Lampu Operasi | SYAS/Lumina | OL9680XVZW00049 | Service | Setting ketinggian lampu operasi | Ready to use |  | Satria
15000280 | 2025-02-12 | RSUD Harapan dan Doa | X-ray Digital | Apelem/Camargue DR | IMA703958 | Service | Kolimator goyang; Service kolimator | Ready to use |  | Sutan
9000548 | 2025-02-12 | RSUD Labuha | UPS | Yoshiga/GPI-33 | E2211CO12260621353 | Survey | Bypass control board unit rusak / hangus | Pending | Unit part bypass dibawa ke kantor pusat untuk diidentifikasi lebih lanjut | Ruslan
19000878 | 2025-02-12 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Survey | Monoblock dibawa untuk diperbaiki | Pending |  | Nanda
19000877 | 2025-02-12 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Survey | QC fail; Trafo sekunder diganti namun monoblock masih rusak | Pending |  | Nanda
12000179 | 2025-02-11 | RSUD Hasri Ainun Habibie | Mesin Anasthesia | Atlan Draeger | ASMM-0118 | Survey | Device check | Ready to use | Disarankan untuk mengganti O2 sensor, unit bisa digunakan tapi kalibrasi | Marshall
15000279 | 2025-02-11 | RSUD Dr. M. Yunus | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/056/C1 | Survey | Pengecekan unit; Unit berfungsi dengan baik | Ready to use |  | Sutan
12000180 | 2025-02-11 | RSUD Ainun Habibie | Baby Incubator | Dreager/Caleo | ASJE0060 | Survey | Humidifier tidak berfungsi disarankan untuk mengganti unit water heater | Ready to use | Bisa digunakan tetapi untuk kelembapan tidak bisa digunakan | Marshall
12000180 | 2025-02-11 | RSUD Ainun Habibie | Baby Incubator | Dreager/Caleo | ASJH0031 | Survey | Humidifier tidak berfungsi disarankan untuk mengganti unit water heater | Ready to use | Bisa digunakan tetapi untuk kelembapan tidak bisa digunakan | Marshall
12000180 | 2025-02-11 | RSUD Ainun Habibie | Baby Incubator | Dreager/Caleo | ASJE0059 | Survey | Humidifier tidak berfungsi disarankan untuk mengganti unit water heater | Ready to use | Bisa digunakan tetapi untuk kelembapan tidak bisa digunakan | Marshall
12000180 | 2025-02-11 | RSUD Ainun Habibie | Baby Incubator | Dreager/Caleo | ASJH0030 | Survey | Humidifier tidak berfungsi disarankan untuk mengganti unit water heater | Ready to use | Bisa digunakan tetapi untuk kelembapan tidak bisa digunakan | Marshall
12000180 | 2025-02-11 | RSUD Ainun Habibie | Baby Incubator | Dreager/Caleo | ASJE0061 | Survey | Humidifier tidak berfungsi disarankan untuk mengganti unit water heater | Ready to use | Bisa digunakan tetapi untuk kelembapan tidak bisa digunakan | Marshall
12000180 | 2025-02-11 | RSUD Ainun Habibie | Baby Incubator | Dreager/Caleo | ASJE0053 | Survey | Humidifier tidak berfungsi disarankan untuk mengganti unit water heater | Ready to use | Bisa digunakan tetapi untuk kelembapan tidak bisa digunakan | Marshall
12000177 | 2025-02-11 | RSUD Ainun Habibie | Ventilator | Draeger/Savina300 | ASLD0402 | Survey | Unit berfungsi normal | Ready to use | Unit bisa digunakan dengan baik | Marshall
12000178 | 2025-02-11 | RSUD Ainun Habibie | Patient Monitor | Draeger/Delta XL | U9SNF1501 | Survey | Pengecekan unit, berfungsi baik | Ready to use | Unit bisa digunakan dengan baik | Marshall
12000178 | 2025-02-11 | RSUD Ainun Habibie | Patient Monitor | Draeger/Delta XL | U9SNF1501 | Survey | Pengecekan unit, berfungsi baik | Ready to use | Unit bisa digunakan dengan baik | Marshall
12000178 | 2025-02-11 | RSUD Ainun Habibie | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0118 | Survey | 02 sensor fail | Ready to use | Disarankan untuk mengganti 02 Sensor | Marshall
12000174 | 2025-02-11 | RSUD Ainun Habibie | Mesin Anasthesia | Draeger/Fabius Plus | ASL10160 | Survey | O2 Sensor failed, cek O2 sensor sudah expired | Ready to use | Disarankan untuk mengganti 02 Sensor | Marshall
20000328 | 2025-02-11 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASLA0082 | Service | Reinstall software | Ready to use |  | Nasrul
11000387 | 2025-02-11 | RSUD Lakipadada | Ventilator | Draeger/Babylog VN800 | ASTE0042 | Survey | Fio2 failure. System check, Hit software, zeroing O2 sensor | Ready to use |  | Nasrul
18000577 | 2025-02-11 | RSUP Ratatotok Buyat | X-ray Analog | Apelem/Camargue Analog | CPD36816J20 | Service | Pengecekan unit | Ready to use |  | Irfan R
6000899 | 2025-02-11 | RSUD dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Atlan A350 | ASRK0304 | Maintenance | Total care Maintenance | Ready to use |  | Dwi
6000899 | 2025-02-11 | RSUD dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Atlan A350 | ASRL0046 | Maintenance | Total care Maintenance | Ready to use |  | Dwi
6000898 | 2025-02-11 | RSUD dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Fabius Tiro | ASLH0077 | Survey | O2 sensor fail | Pending | Disarankan penggantian O2 sensor capsule dan service set 3 tahun | Dwi
6000897 | 2025-02-11 | RSUD dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Primus | ASKJ0050 | Survey | Monitoring gas fail, ventilation valve fail | Pending | Disarankan penggantian ILCA3, flow power, 2/2 port distributing valve dan service set 3 tahun | Dwi
9000551 | 2025-02-10 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | CPD4122OH22 | Service | Pemeriksaan rel ceiling dan pemberian pelumas anti karat pada roda dan rel | Ready to use | Unit bisa digunakan dengan baik | Ruslan
12000190 | 2025-02-10 | RSUD Toto Kabila | Dental Unit | Kavo/1058TM | 20181010205320 | Survey | Lisensi failed, update software tidak berhasil | Ready to use | Unit bisa digunakan tapi perlu pengecekan lebih lanjut | Marshall
12000184 | 2025-02-10 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0081 | Survey | Unit berfungsi normal | Ready to use | Unit bisa digunakan dengan baik | Marshall
12000185 | 2025-02-10 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | ULSNF0251 | Survey | Unit tidak bisa connect ke monitor central | Ready to use | Unit tbisa digunakan tetapi tidak bisa konek dengan monitor central | Marshall
12000181 | 2025-02-10 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Perseus A500 | ASKK0126 | Survey | Pengecekan unit, terjadi kerusakan pada flow power | Pending | Disarankan untuk mengganti flow power unit part | Marshall
12000183 | 2025-02-10 | RSUD Toto Kabila | Ventilator | Draefer/Evita V300 | ASNC0903 | Survey | Pengecekan unit, dan terjadi kerusakan pada touchscreen monitor | Pending | Unit belum bisa digunakan perlu mengganti part | Marshall
20000324 | 2025-02-10 | RSUP Tadjudin Chalid | Laser Surgical | Jena Surgical/Multipulse HoPlus | J191150292 | Survey | Laser tidak menambah, pengecekan fiber optic | Pending | Fiber optic perlu diganti | Nasrul
11000386 | 2025-02-10 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Babylog TN500 | ASTH0132 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
18000576 | 2025-02-10 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Service | Pemasangan stabilizer | Ready to use |  | Irfan R
20000325 | 2025-02-10 | RSUP Dr. Wahidin Sudirohusodo | Dental Unit | Kavo/1058TM | 10205313 | Survey | Alarm unit berbunyi. Pengecekan PCBA unit, dentist board dan resinstall software | Pending | Unit masih belum berfungsi | Sujana
9000545 | 2025-02-08 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Babylog VN500 | ASNL0122 | Survey | Pengecekean touchscreen error dan internal battery drop | Ready to use | Unit disarankan untuk kotrak service (unit sudah beroprasional selama 4 tahun) | Ruslan
9000546 | 2025-02-08 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Evita V300 | ASMD0029 | Survey | Kompressor internal drop, gas supply unit fail | Ready to use | Unit disarankan untuk melakukan pergantian pressure regulator dan kompressor internal | Ruslan
9000544 | 2025-02-07 | RS Bhayangkara Ternate | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0450 | Survey | Pengecekan kondisi unit | Ready to use | Unit disarankan melakukan pergantian unit ketika sudah 2 sampai 4 tahun beroprasional | Ruslan
12000186 | 2025-02-07 | RSUD Toto Kabila | X-ray Digital | Apelem/Camargue DR | IMA703943 | Survey | Terdapat bulatan hitam pada monitor (dead pixel) perlu melakukan pergantian monitor | Ready to use | Unit tetap berfungsi tetapi perlu dilakukan pergantian monitor | Marshall
12000182 | 2025-02-07 | RSUD Toto Kabila | Dental Unit | Kavo/1058TM | 20181010205320 | Survey | Lisensi failed, update software tidak berhasil | Ready to use | Reinstall Software dental unit | Marshall
12000197 | 2025-02-07 | RSUD Bumi Panua | Ventilator | Draeger/Savina300 | ASNL0606 | Survey | Pengecekan unit | Ready to use | Unit berfungsi dengan baik | Marshall
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905392 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905394 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905395 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905396 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905397 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905398 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905399 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905400 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905402 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905752 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905754 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905756 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana |  |  | SO PT Esa Medika Mandiri
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905759 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905761 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905762 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905763 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 974771 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 974774 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 974777 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000326 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 974779 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
11000380 | 2025-02-07 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0108 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
18000582 | 2025-02-07 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Babyleo TN500 | ASNF0004 | Survey | Humiduty tidak berfungsi. Heater rusak | Pending |  | Irfan R
7000663 | 2025-02-07 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista120 | VGSSF0105 | Survey | Training user | Ready to use |  | Wahyu
6000896 | 2025-02-06 | RSUD Ajibarang | Mesin Anasthesia | Draeger/Primus | ASPM0019 | Survey | Leak circuit, waterlock dan sample line, air low pressure; Self test running test | Ready to use | Rekomendasi service set 3 tahun | Dwi
7000697 | 2025-02-06 | RSUP Dr. Sardjito | Infant Warmer | Draeger/Babylog TN800 | DSTK0047 | Survey | Training user | Ready to use |  | Wahyu
8000629 | 2025-02-06 | RSUD RA Kartini | Ventilator | Draeger/Savina 300 | ASFL0163 | Service | Alarm 41.0010, cek motor driver | Ready to use |  | Amna
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905582 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905584 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905590 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905589 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905594 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905595 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905596 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905598 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905599 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905583 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905386 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905387 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905388 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905391 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
11000385 | 2025-02-06 | RSUP Dr. Wahidin Sudirohusodo | Infus Pump | Bbraun/Infusomat | 905392 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
18000575 | 2025-02-06 | RS GMM Siloam Sender | Ventilator | Draeger/Savina300 | ASTJ0069 | Service | Pemasangan selang oksigen | Ready to use |  | Irfan R
19000871 | 2025-02-06 | RS Colombia Asia Semarang | USG | Mindray/ DC-70 | CJ4-15001549 | Survey | Unit berfungsi dengan baik | Ready to use |  | Nanda
7000662 | 2025-02-05 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASKM0402 | Survey | Power supply failure, ventilator failure. Pengecekan power supply dan value pneumatic, cek sensor board | Pending | Unit perlu penggantian power pack, MV2 valve, UGC PCB analog, flow power, ILCA3, service set 3 tahun | Wahyu
7000661 | 2025-02-05 | RSUD Prambanan | Jaudience Meter | Draeger/JM105 | H3500432 | Survey | Unit dibawa untuk di kalibrasi | Pending |  | Wahyu
15000277 | 2025-02-05 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Tube tidak bergerak, dan x-ray tidak bisa exposing. Matikan PU BOX 10 menit | Ready to use |  | Sutan
9000543 | 2025-02-05 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASJH0179 | Survey | Pengecekan unit baterai | Ready to use | Unit disarankan untuk melakukan pergantian servis set 2 tahun | Ruslan
9000543 | 2025-02-05 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASJH0170 | Survey | Pengecekan unit baterai | Ready to use | Unit disarankan untuk melakukan pergantian servis set 2 tahun | Ruslan
9000543 | 2025-02-05 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASJH0180 | Survey | Pengecekan unit baterai | Ready to use | Unit disarankan untuk melakukan pergantian servis set 2 tahun | Ruslan |  |  | SO PT Esa Medika Mandiri
18000583 | 2025-02-05 | RSUP Prof. Dr. R. D. Kandou Manado | Bor Tulang | Aesculap/Elan4 | 003976 | Service | Pemasangan spare part backup | Ready to use |  | Irfan R |  |  | SO PT Esa Medika Mandiri
19000870 | 2025-02-05 | RSUD Kota Salatiga | Mesin Anasthesia | Draeger/Primus | ASHL0060 | Survey | Expiration flow sensor failed. Penggantian flow sensor; Terdapat dot dilayar, LCD rusak | Ready to use | Disarankan penggantian layar lcd | Nanda |  |  | SO PT Esa Medika Mandiri
9000550 | 2025-02-04 | RSUD Dr. H. Chasan Boesoirie | Printer Mammografi | FujiFilm / Drypix | P6770193 | Service | Tray 2 Error, dilakukan setting ukuran printer | Ready to use | Unit sudah berfungsi dengan baik | Ruslan |  |  | SO PT Esa Medika Mandiri
12000170 | 2025-02-04 | RSUD Kab. Bolaang Mongondow | X-ray Digital | Apelem/Camargue DR | 84TD95 | Survey | Alat berfungsi dengan baik | Ready to use | Unit berfungsi dengan baik | Marshall |  |  | SO PT Esa Medika Mandiri
12000168 | 2025-02-04 | RSUD Datoe Binangkang | X-ray Digital | Apelem/Camargue DR | 84W100 | Survey | Terjadi kerusakan pada baterai detektor (baterai bocor) | Ready to use | Perlu dilakukan pergantian baterai detektor | Marshall |  |  | SO PT Esa Medika Mandiri
8000627 | 2025-02-04 | RSUD Brebes | Ventilator | Draeger/Babylog VN500 | ASJH0095 | Service | Alarm FIO2 failure. Cek purity gas, gas central bermasalah | Pending |  | Amna |  |  | SO PT Esa Medika Mandiri
18000584 | 2025-02-04 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | ASRK0009 | Survey | LCD blank. Konektor LCD rusak | Pending |  | Irfan R |  |  | SO PT Esa Medika Mandiri
19000896 | 2025-02-04 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASJC0044 | Survey | Air supply fail; Penggantian pressure regulator milik RS | Ready to use |  | Nanda |  |  | SO PT Esa Medika Mandiri
6000895 | 2025-02-04 | RS Duta Mulya | Ventilator | Draeger/Savina300 | ASSK0023 | Survey | Pengecekan unit; O2 cell expired | Ready to use | Disarankan penggantian O2 cell dan service set | Dwi |  |  | SO PT Esa Medika Mandiri
6000895 | 2025-02-04 | RS Duta Mulya | Ventilator | Draeger/Savina300 | ASSK0024 | Survey | Pengecekan unit; O2 cell expired | Ready to use | Disarankan penggantian O2 cell dan service set | Dwi |  |  | SO PT Esa Medika Mandiri
6000895 | 2025-02-04 | RS Duta Mulya | Ventilator | Draeger/Savina300 | ASRJ0121 | Survey | Pengecekan unit; O2 cell expired | Ready to use | Disarankan penggantian O2 cell dan service set | Dwi |  |  | SO PT Esa Medika Mandiri
4000480 | 2025-02-04 | RSUD Dr. Drajat Prawiranegara | ESWL | Medispec/EM1000 | 0617048 | Service | HVPS rusak, pemasangan HVPS 4pcs | Pending | HVPS masih rusak | Bahar, Satria |  |  | SO PT Esa Medika Mandiri
16000276 | 2025-02-04 | Signature Dental Care | Dental Unit | Kavo/1058TM | 10207194 | Service | Water IO tidak bisa diperasikan dari foot step, pengecekan kabel LAN; Setting IP Address | Ready to use |  | Ipan, Farhat |  |  | SO PT Esa Medika Mandiri
11000384 | 2025-02-04 | RSUP Dr. Wahidin Sudirohusodo | Infant Warmer | Draeger/Resuscitaire | DSTJ0114 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000384 | 2025-02-04 | RSUP Dr. Wahidin Sudirohusodo | Infant Warmer | Draeger/Resuscitaire | DSTJ0108 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000384 | 2025-02-04 | RSUP Dr. Wahidin Sudirohusodo | Infant Warmer | Draeger/Resuscitaire | DSTJ0109 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000384 | 2025-02-04 | RSUP Dr. Wahidin Sudirohusodo | Infant Warmer | Draeger/Resuscitaire | DSTJ0110 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000384 | 2025-02-04 | RSUP Dr. Wahidin Sudirohusodo | Infant Warmer | Draeger/Resuscitaire | DSTJ0111 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000384 | 2025-02-04 | RSUP Dr. Wahidin Sudirohusodo | Infant Warmer | Draeger/Resuscitaire | DSTJ0112 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000384 | 2025-02-04 | RSUP Dr. Wahidin Sudirohusodo | Infant Warmer | Draeger/Resuscitaire | DSTJ0113 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000383 | 2025-02-03 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Babylog TN500 | ASTH0129 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000383 | 2025-02-03 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Babylog TN500 | ASTH0130 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000383 | 2025-02-03 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Babylog TN500 | ASTH0131 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000383 | 2025-02-03 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Babylog TN500 | ASTH0139 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000383 | 2025-02-03 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Babylog TN500 | ASTH0140 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000383 | 2025-02-03 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Babylog TN500 | ASTJ0174 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
19000895 | 2025-02-03 | RS Telogorejo | Ventilator | Draeger/Savina300 | ASLL0025 | Survey | Alarm baterai menyala, tekanan drop; Baterai rusak | Pending | Disarankan penggantian baterai set | Nanda
2200006 | 2025-02-03 | Siloam Hospital manado | Hemodialisa | Bbraun | 227271 | Service | Test temperature not OK; Pembersihan sensor temperature dan kalibrasi | Ready to use |  | Ari
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0108 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0096 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0097 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0046 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0033 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0035 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0109 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0095 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0101 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0105 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000382 | 2025-02-01 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0100 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0099 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0048 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0098 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0102 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0104 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0106 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0107 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTH0110 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0034 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
11000381 | 2025-01-31 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette 8000 | DSTJ0041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
6000894 | 2025-01-31 | RST Wijayakusuma | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0181 | Survey | O2 sensor fail. Baterai dan o2 cell tidak berfungsi | Pending | Disarankan penggantian service set 3 tahun dan O2 cell | Dwi |  |  | SO PT Esa Medika Mandiri
6000894 | 2025-01-31 | RST Wijayakusuma | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0176 | Survey | O2 sensor fail. Baterai dan o2 cell tidak berfungsi | Pending | Disarankan penggantian service set 3 tahun dan O2 cell | Dwi |  |  | SO PT Esa Medika Mandiri
8000628 | 2025-01-31 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASKL0012 | Service | Baterai fail | Pending | Ajukan claim warranty battery pack | Amna |  |  | SO PT Esa Medika Mandiri
7000660 | 2025-01-31 | RSUD Merah Putih | X-ray Analog | Apelem/Camargue Analog | IMA703979 | Service | Cover handle kendor; Kecangkan dan pasang baut bara untuk cover tube dan cover handle | Ready to use |  | Wahyu |  |  | SO PT Esa Medika Mandiri
16000275 | 2025-01-31 | RSUD Kab. Bekasi | Bed Pediatric | SYAS | 240526302000900000 | Service | Penggantian handle bed | Ready to use | Claim warranty | Ipan |  |  | SO PT Esa Medika Mandiri
19000869 | 2025-01-30 | RS Nasional Diponegoro | Incubator Transport | Draeger/Isolette TI500 | AF10588 | Service | Baterai internal drop; Baterai sudah harus diganti | Pending |  | Nanda |  |  | SO PT Esa Medika Mandiri
6000893 | 2025-01-30 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Babylog VN800 | ASTE0047 | Survey | Training user | Ready to use |  | Dwi |  |  | SO PT Esa Medika Mandiri
19000876 | 2025-01-25 | RS Paru Ario Wirawan | Ventilator | Draeger/Savina300 | ASTL0146 | Survey | Installasi unit demo | Ready to use |  | Nanda |  |  | SO PT Esa Medika Mandiri
19000688 | 2025-01-25 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRH0033 | Service | Pressure sensor calibration failed; Pemasangan membrane | Ready to use |  | Nanda |  |  | SO PT Esa Medika Mandiri
2000191 | 2025-01-24 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Survey | Error generator; Monoblock rusak | Pending |  | Eddy |  |  | SO PT Esa Medika Mandiri
9000542 | 2025-01-24 | RSUD Prof. Dr. H. Aloei Saboe | Ventilator | Draeger/Evita V300 | ASME0153 | Survey | O2 supply tidak terbaca, FIO2 error, baterai fail; Baterai rusak | Pending | Disarankn penggantian baterai pack dan kompressor portable | Ruslan |  |  | SO PT Esa Medika Mandiri
9000541 | 2025-01-24 | RSUD Prof. Dr. H. Aloei Saboe | Mesin Anasthesia | Draeger/Fabius Plus | ASJH0080 | Survey | System check, daily check, mode vintilasi gagal, baterai fail | Pending | Disarankan penggantian baterai dan service set 3tahun | Ruslan |  |  | SO PT Esa Medika Mandiri
18000573 | 2025-01-24 | RSD Tidore | Ventilator | Draeger/Babylog | ASJL0036 | Survey | Flow tidak terbaca; Flow sensor rusak | Pending |  | Irfan |  |  | SO PT Esa Medika Mandiri
11000379 | 2025-01-24 | RS Regional dr. Hasri Ainun Habibie | Bubble CPAP | Draeger/Oxymixer | QM20246035 | Survey | Flow meter kemasukan air; Pemgurasan air pada kompresor dan flow meter | Ready to use |  | Nasrul |  |  | SO PT Esa Medika Mandiri
3000191 | 2025-01-24 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Survey | Error generator. Cek generato, cek monoblock, cek power supply board | Pending | Perlu penggantian monoblock | Imam |  |  | SO PT Esa Medika Mandiri
19000861 | 2025-01-24 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASMK0321 | Service | Pengukuran tidak akurat; Kalibrasi pressure | Ready to use |  | Nanda |  |  | SO PT Esa Medika Mandiri
11000378 | 2025-01-23 | RSUP Dr. Wahidin Sudirohusodo | Jaudience Meter | Draeger/JM105 | B7036828 | Survey | Alat mati total. Baterai drop, adaptor rusak | Pending |  | Nasrul |  |  | SO PT Esa Medika Mandiri
14000331 | 2025-01-23 | RSPAL dr. Ramlan | Laser Surgical | Jena Surgical/Multipulse HoPlus | J191150371 | Survey | Test fiber; Fungsi normal | Ready to use |  | Purwono, Bahar, Satria |  |  | SO PT Esa Medika Mandiri
9000539 | 2025-01-23 | RSUD Ainun Habibie | Ventilator | Draeger/Caleo | ASJH0030 | Survey | Humidifier tidak berfungsi; Pengecekan chamber humidifier | Ready to use | Disarankan penggantian 2MSI1320, kit water heater dan MV230v | Ruslan |  |  | SO PT Esa Medika Mandiri
3000362 | 2025-01-23 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Lifter error ET130; Plug and play connector DRV | Ready to use |  | Imam
12000164 | 2025-01-22 | RSUD ODSK Provinsi Sulawesi Utara | Pendant | Draeger/Movita | ASJD0327 | Service | Pemasangan spare part backup | Ready to use |  | Marshall
12000164 | 2025-01-22 | RSUD ODSK Provinsi Sulawesi Utara | Pendant | Draeger/Movita | ASJD0324 | Service | Pemasangan spare part backup | Ready to use |  | Marshall
12000164 | 2025-01-22 | RSUD ODSK Provinsi Sulawesi Utara | Pendant | Draeger/Movita | ASHF0154 | Service | Pemasangan spare part backup | Ready to use |  | Marshall
12000164 | 2025-01-22 | RSUD ODSK Provinsi Sulawesi Utara | Pendant | Draeger/Movita | ASHF0153 | Service | Pemasangan spare part backup | Ready to use |  | Marshall
4000479 | 2025-01-22 | RSUP Surabaya | C-arm | Apelem/Evo R+ | 551410231085 | Service | Uji Kesuaian | Ready to use |  | Satria, Bahar, Purwono
4000479 | 2025-01-22 | RSUP Surabaya | C-arm | Apelem/Evo R+ | 551410231087 | Service | Uji Kesuaian | Ready to use |  | Satria, Bahar, Purwono
9000537 | 2025-01-22 | RSUD Dr. Iwan Boking | Mesin Anasthesia | Draeger/Fabius Plus | ASKE0185 | Survey | Baterai fail, flow calibration fail, O2 failed. | Pending | Disarankan panggenatian service set 3 tahun dan kabel O2 housing | Ruslan
9000536 | 2025-01-22 | RSUD Dr. Iwan Boking | Ventilator | Draeger/Savina300 | ASJL0134 | Survey | Expiration flow sensor failed. Cek valve dan ganti O2 sensor | Ready to use | Disarankan penggantian kabel harness karna muncul error flow sensor dan service set 6 tahun | Ruslan
20000323 | 2025-01-22 | RSUD Kota Makassar | Mamografi | Metaltronica/Helianthus DBT | ILILHF/0627/CB2 | Survey | Error HV Generator. Cek fuse, cek inverter, cek generator; Inverter rusak | Pending |  | Sujana
19000875 | 2025-01-22 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Draeger/Fabius Plus | ARZM0051 | Survey | Nilai PEEP tidak akurat; Service mode, PEEP cell masih bagus | Ready to use | Unit masih berfungsi namun disarankan penggantian service set 2 tahun, keypad, cozy service set 2 tahun | Nanda
11000376 | 2025-01-21 | RSAU Dr. Dadi Sardjoto | Baby Incubator | Draeger/Babyleo TN500 | ASLL0012 | Survey | Parameter sensor failure; PBA Sensor | Pending |  | Nasrul
7000659 | 2025-01-21 | Klinik Graha Amanah | Jaudience Meter | Draeger/JM105 | 2401082 | Service | Error 05; Penggantian baterai | Ready to use |  | Wahyu
9000534 | 2025-01-21 | RSUD Ainun Habibie | MOT | SYAS | - | Service | Membran monolitich bocor; Penggantian membran monolitich dengan frame yg baru | Ready to use |  | Ruslan
16000274 | 2025-01-21 | GiGi Dental House | Dental Unit | Kavo/1058TM | 10207196 | Survey | Air pada unit suka tidak keluar, pengecekan jalur air dan angin, jalur aman | Ready to use | Unit berfungsi normal, kemungkinan rusak pada switch komrepssor | Ipan
18000564 | 2025-01-21 | RSUP Prof. Dr. R. D. Kandou Manado | Bor Tulang | Aesculap/Elan4 | 003976 | Survey | Bor hanpiece tidak berputar; Handpiece dibawa untuk dicek lebih lanjut | Pending |  | Irfan R
9000535 | 2025-01-21 | RSUD Ainun Habibie | Mesin Anasthesia | Draeger/Atlan A300 | ASMM0118 | Survey | O2 sensor fail; O2 sensor expired | Pending | Disarankan oenggantian O2 sensor | Ruslan
15000275 | 2025-01-21 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Service | Mobile xray terdapat bercak; Bersihkan board dengan contact cleaner | Ready to use |  | Sutan
8000625 | 2025-01-21 | RSUD Batang | Mesin Anasthesia | Draeger/Atlan A350 | ASTH0135 | Service | Penggantian ventilator assembly | Ready to use |  | Amna
6000892 | 2025-01-21 | RST Wijayakusuma | Ventilator | Draeger/Savina300 | ASEL0196 | Survey | baterai low, FiO2 calibration fail, baterai fail. | Pending | Disarankan penggantian baterai internal, external dan O2 cell | Dwi
6000892 | 2025-01-21 | RST Wijayakusuma | Ventilator | Draeger/Savina300 | ASEL0195 | Survey | baterai low, FiO2 calibration fail, baterai fail. | Pending | Disarankan penggantian baterai internal, external dan O2 cell | Dwi
6000892 | 2025-01-21 | RST Wijayakusuma | Ventilator | Draeger/Savina300 | ASEL0198 | Survey | baterai low, FiO2 calibration fail, baterai fail. | Pending | Disarankan penggantian baterai internal, external dan O2 cell | Dwi
6000892 | 2025-01-21 | RST Wijayakusuma | Ventilator | Draeger/Savina300 | ASEL0201 | Survey | baterai low, FiO2 calibration fail, baterai fail. | Pending | Disarankan penggantian baterai internal, external dan O2 cell | Dwi
6000892 | 2025-01-21 | RST Wijayakusuma | Ventilator | Draeger/Savina300 | ASEL0202 | Survey | baterai low, FiO2 calibration fail, baterai fail. | Pending | Disarankan penggantian baterai internal, external dan O2 cell | Dwi
11000377 | 2025-01-21 | RSUP Makassar | Mesin Anasthesia | Draeger/Primus | ASRH0038 | Service | Baterai internal failure, tekanan drop. Cek sistem, power supply, di charger manual | Ready to use |  | Nasrul
19000860 | 2025-01-20 | RS Telogorejo | Mesin Anasthesia | Draeger/Primus | ASDK0223 | Service | Mixer gas failure; Tarik error, pressure regulator rusak | Pending | Perlu penggantian regulator set | Nanda
12000166 | 2025-01-20 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Survey | Low KV. Unit disambung ke ups | Pending | UPS masih belum cover. Disarankan menggunakan UPS 20kVa single phase | Marshall
9000540 | 2025-01-20 | RSUD Ainun Habibie | Mesin Anasthesia | Draeger/Fabius Plus | ASLI0160 | Survey | O2 kalibrasi gagal; O2 sensor expired | Pending | Disarankan oenggantian O2 sensor | Ruslan
4000478 | 2025-01-20 | RSUP Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0146 | Service | Uji Kesesuaian dan Training user | Ready to use |  | Satria, Bahar, Purwono
7000657 | 2025-01-20 | RSUD Bagas Waras | Infant Warmer | Draeger/Resuscitaire | JY03385 | Service | Error 05; Penggantian heater assy | Ready to use |  | Wahyu
7000658 | 2025-01-20 | RS JIH | Mesin Anasthesia | Draeger/Atlan A350 | ASRJ0460 | Survey | O2 delivery failed, sample line expired. Cek error, cek gas monitoring module; Pump flow power rusak | Pending |  | Wahyu
8000624 | 2025-01-20 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0148 | Service | Alarm apnea pressure. General check, system check, valve Cozy check | Ready to use |  | Amna
10000223 | 2025-01-19 | RSUD Mas Amsyar | PEndant | SYAS | SY-PS2402004 | Survey | Gas break pendant tidak berfungsi; Kompressor tidak proper | Pending |  | Kreshna
6000891 | 2025-01-17 | RSU An Ni'mah | Ventilator | Draeger/Savina300 | ASPJ0377 | Maintenance | Preventive Maintenance | Ready to use | Direkomendasikan penggantian O2 cell | Dwi
3000361 | 2025-01-17 | RSUD Sidoarjo Barat | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0755 | Service | Table tidak berfungsi; Kalibrasi meja | Ready to use |  | Imam, Purwono
2000190 | 2025-01-17 | RSUD Kab. Temanggung | Bone Mineral Densitometer | Apelem/Medix DR | H237H1092 | Survey | Generator fault; Inverter dan tube monoblock rusak | Pending |  | Eddy
2000189 | 2025-01-17 | RSUD Kab. Temanggung | C-arm | Apelem/Evo R+ | 55115011469 | Service | KV error, cek kabel komunikasi; Perbaikan kabel dan cover kabel komunikasi WS console ke unit | Ready to use |  | Eddy
8000623 | 2025-01-17 | RSUD RA Kartini | Ventilator | Draeger/Savina300 | ASEL0163 | Service | I:E ratio tidak bisa konek ke pasien. System check, check expiratory valve | Ready to use |  | Amna
8000623 | 2025-01-17 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHL0168 | Survey | Baterai fail tidak bisa charge; | Pending | Ajukan claim warranty | Amna
11000374 | 2025-01-17 | RSUD Kota Makassar | Mamografi | Metaltronica/Lilyum | LILHF/0627/CB2 | Survey | HV drop; Blok inverter rusak | Pending |  | Nasrul
20000320 | 2025-01-16 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0145 | Service | table top berbunyi; pemasangan cover | Ready to use |  | Sujana
6000890 | 2025-01-16 | RSU Wiradadi Husada | Ventilator | Draeger/Evita V300 | ASNL0235 | Survey | air pressure dan O2 pressure gagal | Pending | sensor motor valve perlu diganti | Dwi
8000622 | 2025-01-16 | RSUD dr. R. Soetijono Blora | Ventilator | Draeger/Babylog VN500 | ASLL0066 | Survey | Battery fail dan kit cover inlet fail | Pending | Perlu penggantian baterai pack, kit gas inlet, neonatal flow sensor dan insert neonatal flow | Amna
9000533 | 2025-01-16 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 0920036 | Survey | Hasil USG tidak terbaca; Perbaiki posisi kabel S Video | Ready to use |  | Ruslan
14000330 | 2025-01-16 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 0720015 | Service | Hardware not connected. Board XR-5 dibawa untuk dicoba di unit lain | Pending |  | Imam, Purwono
18000565 | 2025-01-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN800 | 98711107 | Survey | Nebulizer tidak berfungsi, tarik error; Pressure regulator bocor | Pending | Dibuatkan penawaran harga | Irfan R
6000889 | 2025-01-16 | RSUD Banyumas | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0075 | Survey | Baterai low, gas sensor monitoring fail, vapor3000 leak. Cek keseluruhan alat | Pending | Butuh penggantian baterai, ILCA3, flow power | Dwi
19000859 | 2025-01-16 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Service | Shock tidak ada tp counting berjalan. Cek quick seasoning, cek emh aman. Kemungkinan HVPS rusak | Pending | HVPS dibawa untuk di cek lebih lanjut | Nanda
6000888 | 2025-01-15 | RSUD Ajibarang | Baby Incubator | Draeger/Isolette 8000 | AC09097 | Survey | Setting temperature dan air mode tidak sesuai. Cek heater, motor fan, sensor air tempt; Setting dan real temp sudah akurat | Pending |  | Dwi
6000888 | 2025-01-15 | RSUD Ajibarang | Baby Incubator | Draeger/Isolette 8000 | AC09098 | Survey | Setting temperature dan air mode tidak sesuai. Cek heater, motor fan, sensor air tempt; Setting dan real temp sudah akurat | Pending |  | Dwi
9000532 | 2025-01-15 | RS Gunung Maria Tomohon | Mesin Anasthesia | Draeger/Fabius Plus | ASRK0058 | Service | O2 calibration failed; Penggantian O2 sensor | Ready to use |  | Ruslan, Marshall
18000572 | 2025-01-15 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0413 | Survey | Troubleshoot perbaikan kabel harnes | Pending | Dibuatkan penawaran harga | Irfan R
15000273 | 2025-01-15 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Joystick table top bermasalah; Penggantian joystick baru | Ready to use |  | Sutan
20000322 | 2025-01-15 | RS Universitas Hasanuddin | Lampu Operasi | Draeger/Polaris 600 | ASPN0042 | Survey | pemasangan part Reg.6 | Ready to use |  | Sujana
20000321 | 2025-01-14 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0145 | Service | Penyambungan koneksi PACS dengan unit | Ready to use |  | Sujana
7000656 | 2025-01-14 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Service | deflate function bermasalah. Cek water system dan pressure regulator; Penggantian pressure sensor water | Ready to use |  | Wahyu
3000360 | 2025-01-14 | RSUD Gambiran | C-arm | Apelem/Evo R+ | 55121001850 | Survey | Unit mati total; Software tidak berjalan, Powersupply mati, HDD komputer rusak | Pending |  | Imam, Purwono
8000620 | 2025-01-14 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Primus | ASHM0012 | Service | Sensor check fail; Pressure sensor tidak berfungsi | Pending |  | Amna
6000887 | 2025-01-14 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASJH0125 | Survey | Air supply bermasalah. Cek proportional valve, self test dan service connect | Pending | Perlu penggantian proportional valve 2pcs | Dwi
6000886 | 2025-01-14 | RSUD Brebes | Ventilator | Draeger/Evita 3000 | ASHA0095 | Survey | failure device software unit. Cek service connect draeger | Pending | Perlu penggantian basic board (mainboard) | Dwi
19000858 | 2025-01-14 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista120 | V3SHA0042 | Service | Konektor SPO2 patah; konektor SPO2 di lem karna kondisi sensor masih baik | Ready to use | Unit masih bisa digunakan | Nanda
11000375 | 2025-01-13 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASAM0045 | Survey | Speaker fail, O2 sensor fail, flow sensor fail | Pending | Disarankan penggantian speaker set dan O2 cell | Nasrul
6000885 | 2025-01-13 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASNE0151 | Survey | O2 sensor error notification. Cek O2 sensor kit, selftest | Pending | Perlu penggantian O2 sensor PIA | Dwi
6000884 | 2025-01-13 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHM0013 | Survey | Unit mati total. Sensor check, power pack check | Pending | Perlu penggantian power pack power supply | Dwi
7000700 | 2025-01-13 | RSUD Merah Putih | Baby Incubator | Draeger/Isolette C2000 | DSRL0026 | Survey | Nilai kelembapan tidak sesuai. Cek heater humidifier, bersihkan | Ready to use |  | Wahyu
3000359 | 2025-01-13 | RSUD Dr. Saiful Anwar | Bone Mineral Densitometer | Apelem/Medix DR | J207H0735 | Service | Calculation error; Re install software | Ready to use |  | Imam, Purwono
15000272 | 2025-01-13 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Meja tidak bisa bergerak. Bersihkan PUBOX dan kencangkan soket soketnya | Ready to use |  | Sutan
04000521 | 2025-01-10 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Service | DICOM tidak berfungsi; Pengecekan koneksi LAN dan LAN Switch. Direkomendasikan penggantian Switch Hub | Ready to use |  | Satria, Imam
11000372 | 2025-01-10 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0145 | Survey | Koordinasi penyambungan PACS | Ready to use |  | Nasrul
11000373 | 2025-01-10 | RSUP Makassar | Mobile X-ray | Apelem/Rafale BDR | 97124084483 | Survey | Training user | Ready to use |  | Nasrul
6000883 | 2025-01-10 | RSUD dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Primus | ASRL0050 | Survey | Gas sensor fail, valve ventilator error. Self test, cek PGM, ILCA 3 error
10000222 | 2025-01-10 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186P0271 | Survey | LCD touchscreen console tidak bisa digunakan | Pending | Disarankan untuk mengganti touchscreen console | Kreshna
9000530 | 2025-01-10 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Fabius Plus | ASJH0080 | Survey | O2 calibration failed, system test failed | Pending | Disarankan penggantian service set 1 tahun dan O2 sensor | Ruslan
9000530 | 2025-01-10 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Fabius Plus | ASJL0008 | Survey | O2 calibration failed, system test failed | Pending | Disarankan penggantian service set 1 tahun dan O2 sensor | Ruslan
9000530 | 2025-01-10 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Fabius Plus | ASJL0005 | Survey | O2 calibration failed, system test failed | Pending | Disarankan penggantian service set 1 tahun dan O2 sensor | Ruslan
9000530 | 2025-01-10 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Fabius Plus | ASJK0228 | Survey | O2 calibration failed, system test failed | Pending | Disarankan penggantian service set 1 tahun dan O2 sensor | Ruslan
4000475 | 2025-01-10 | RSUD Abdoel Wahab Sjahranie | Instrument | Aesculap | 002282 | Instalasi | Instrument set lengkap | Ready to use |  | Satria
2000188 | 2025-01-10 | RSUP Dr. H. Adam Malik Medan | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/083/C1 | Survey | Error system detected low kv output | Pending | Perlu penggantian inverter, HTT dan cable HTT | Eddy
16000272 | 2025-01-10 | GiGi Dental House | Dental Unit | Kavo/1058TM | 10207196 | Service | Dental unit bunyi alarm, notif oxygenal empty; Pengisian ulang cairan oxygenal | Ready to use |  | Ipan
3000358 | 2025-01-10 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Meja tidak ada pergerakan, cek koneksi DRV, kalibrasi meja | Ready to use |  | Imam, Habibi, Kreshna
12000153 | 2025-01-10 | RSUD Toto Kabila | C-arm | Apelem/Evo R+ | 55121004853 | Service | Penggantian spare part travo, inverter, psu 24v dan power board | Ready to use |  | Marshall
9000531 | 2025-01-10 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Perseus A500 | ASKK0126 | Service | Gas monitoring failure; Penggantian flow power | Ready to use |  | Ruslan, Marshall
8000619 | 2025-01-10 | RSUD Dr. Loekmono Hadi | Infant Warmer | Draeger/Resuscitaire | AY11083 | Survey | Error 05; Heater rusak | Pending |  | Amna
19000874 | 2025-01-10 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0150 | Service | Error alarm modul. Tarik error dan kirim error log ke teknisi draeger | Pending |  | Nanda
20000318 | 2025-01-09 | RSUP Dr. Wahidin Sudirohusodo | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0495 | Survey | Error fluoros tidak bisa expose dan hasil gambar kurang bagus. Sudah kalibrasi detector, hasil gambar masih kurang bagus | Pending | Disarankan penggantian detektor | Sujana
6000882 | 2025-01-09 | RSI Banjarnegara | Ventilator | Draeger/Savina300 | ASSE0143 | Maintenance | Preventive Maintenance Total Care | Ready to use |  | Dwi
6000882 | 2025-01-09 | RSI Banjarnegara | Ventilator | Draeger/Savina300 | ASSE0142 | Maintenance | Preventive Maintenance Total Care | Ready to use |  | Dwi
6000882 | 2025-01-09 | RSI Banjarnegara | Ventilator | Draeger/Savina300 | ASSE0144 | Maintenance | Preventive Maintenance Total Care | Ready to use |  | Dwi
6000882 | 2025-01-09 | RSI Banjarnegara | Ventilator | Draeger/Savina300 | ASSE0141 | Maintenance | Preventive Maintenance Total Care | Ready to use |  | Dwi
8000618 | 2025-01-09 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Perseus A500 | ASNK0040 | Survey | LCD fail; Board Guluc rusak | Pending |  | Amna
18000571 | 2025-01-09 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0413 | Survey | Flour sensor tidak terbaca. Tarik error; Kabel harnas bermasalah | Pending |  | Irfan R
7000655 | 2025-01-09 | RSUD Saras Adyatma | Mesin Anasthesia | Draeger/Atlan A350 | ASTH0115 | Service | Ventilator failure; Penggantian piston ventilator kit | Ready to use |  | Wahyu
19000857 | 2025-01-09 | RSUD dr. Gunawan Mangunkusumo | Patient Monitor | Draeger/Vista120 | VTSLK0025 | Survey | Layar ngeblock garis; Front panel LCD ada yang rusak | Ready to use | Unit masih bisa digunakan namun disarankan penggantian layar LCD | Nanda
19000855 | 2025-01-08 | RSUD Ibu Fatmawati Surakarta | Infant Warmer | Draeger/Resuscitaire | DSNL0047 | Service | Error 05; Heater putus, sementara diperbaiki dengan disambung | Ready to use |  | Nanda
20000319 | 2025-01-08 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0145 | Service | Table top berbunyi; perbaikan cover | Ready to use |  | Sujana
6000881 | 2025-01-08 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASLD0042 | Survey | Gas sensor fail, inaccurate O2, leak system. Sensor ILCA3 problem | Pending | Perlu penggantian sensor ILCA3 dan membran piston | Dwi
4000474 | 2025-01-08 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Uji Kesuaian | Ready to use | Alat KSO | Satria, Bahar
16000271 | 2025-01-08 | GiGi Dental House | Dental Unit | Kavo/E50 Life | 30211888 | Survey | Water tank sulit dibuka dan angin tidak keluar; Water tank sudah bisa dibuka dan kompressor sudah diperbaiki | Ready to use |  | Ipan
8000617 | 2025-01-08 | RSUD Kayen Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASSH0043 | Handover | Unit baru digunakan menunggu dokter sp paru | Ready to use |  | Amna
11000370 | 2025-01-08 | RSUD Kota Makassar | Mamografi | Metaltronica/Lilyum | ILILHF/0627/CB2 | Survey | HV generator drop, fuse rusak | Pending |  | Nasrul
11000371 | 2025-01-07 | RSUP Tadjudin Chalid | Digital Radiografi Fluoroskopi | Apelem/Platinum | J236F0882 | Service | PACS disconnect; Setting ulang PACS | Ready to use |  | Nasrul
20000351 | 2025-01-07 | RSUP Jayapura | Instrument | Aesculap | x | Instalasi | Instrument set lengkap | Ready to use |  | Sujana
6000879 | 2025-01-07 | RSKBD Panti Nugroho | Ventilator | Draeger/Savina300 | ASLD0168 | Survey | Training user | Ready to use |  | Dwi
1000221 | 2025-01-07 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186P0271 | Survey | Error layar console | Pending | Diperlukan upgrade software | Kreshna
9000529 | 2025-01-07 | RSUP Prof. Dr. R. D. Kandou Manado | Pendant | Draeger/Movita | ASJD0324 | Survey | Valve block break tidak berfungsi; Pengecekan dan perbaikan minor | Ready to use | Sementara unit bisa digunakan | Ruslan, Marshall
6000880 | 2025-01-07 | RS Orthopaedi Purwokerto | Ventilator | Draeger/Savina300 | ASTC0258 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Dwi
8000616 | 2025-01-07 | RS Fastabiq Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASTH0060 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
3000357 | 2025-01-07 | RSUD Pringsewu | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0756 | Survey | Ambil Raw Data detektor | Ready to use |  | Imam
11000369 | 2025-01-06 | RS Universitas Hasanuddin | Lampu Operasi | Draeger/Polaris 600 | ASPN0042 | Service | Camera tidak bisa konek, penggantian recover box dan adaptor. Unit masih error | Pending | Menunggu part Vcodic trans reg. 6 | Nasrul
8000615 | 2025-01-06 | RSI Sultan Agung Semarang | Mesin Anasthesia | Draeger/Primus | ASLD0028 | Service | Penggantian service set 2 tahun (mpgm IE) | Ready to use |  | Amna
7000651 | 2025-01-06 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0091 | Service | Nebulizer tidak berfungsi; Penggantian pressure regulator | Ready to use |  | Wahyu
14000329 | 2025-01-06 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/EM1000 | 0720015 | Service | Hardware not connected, cek data, cek ip, cek console | Pending | Akan dijadwalkan pengecekan lebih lanjut | Purwono
19000851 | 2025-01-06 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | Busi bocor; Penggantian karet ring busi | Ready to use |  | Nanda
11000368 | 2025-01-03 | RSUP Makassar | Mobile X-ray | Apelem/Rafale BDR | 97124084483 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Nasrul
19000873 | 2025-01-03 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | QC fail; Penggantian koneksi kabel LAN, unit masih belum bisa. | Pending | Mainboard dibawa untuk dicek lebih lanjut | Nanda
6000877 | 2025-01-03 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASTJ0021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Dwi
6000877 | 2025-01-03 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASTD0075 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Dwi
6000877 | 2025-01-03 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASTE0026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Dwi
6000877 | 2025-01-03 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASTE0029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Dwi
6000877 | 2025-01-03 | RSUP Dr. Sardjito | Ventilator | Draeger/Isolette8000 | DSTJ0017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Dwi
6000877 | 2025-01-03 | RSUP Dr. Sardjito | Ventilator | Draeger/Isolette8000 | DSTJ0021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Dwi
6000877 | 2025-01-03 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Babyleo TN500 | ASTJ0047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Dwi
6000877 | 2025-01-03 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Babyleo TN500 | ASTJ0088 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Dwi
1000218 | 2025-01-03 | RSUD Sultan Suriansyah | X-ray Digital | Apelem/Camargue DR | CPD33483019 | Survey | Software dan monitor layar berwarna hitam | Pending | Dibuatkan penawaran harga | Kreshna
8000613 | 2025-01-03 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0043 | Service | O2 sensor fail; perlu penggantian O2 sensor capsule, Co2y | Pending |  | Amna
8000613 | 2025-01-03 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0108 | Service | O2 sensor fail; perlu penggantian O2 sensor capsule, Co2y | Pending |  | Amna
8000613 | 2025-01-03 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0147 | Service | O2 sensor fail; perlu penggantian O2 sensor capsule, Co2y | Pending |  | Amna
3000356 | 2025-01-03 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Lifter error, cek koneksi drv dan paramter axis contoller | Ready to use | Kondisi tegangan kurang stabil, disarankan menggunakan UPS | Imam
16000270 | 2025-01-02 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 2808201 | Survey | Handpiece mikromotor tidak memutar | Pending | Handpiece belum bisa dibuka karna keterbatasan tool | Ipan
17000208 | 2025-01-02 | RSUD Dr. Drajat Prawiranegara | X-ray Digital | Apelem/Camargue DR | 2510220562 | Service | Rotor fault, pengecekan koneksi rotor di tube; Kabel putus digigit tikus |  |  | Bahar
8000612 | 2025-01-02 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTL0299 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Amna
8000612 | 2025-01-02 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTL0302 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Amna
8000612 | 2025-01-02 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTK0258 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Amna
8000612 | 2025-01-02 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTL0300 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Amna
8000612 | 2025-01-02 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTL0301 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Amna
8000612 | 2025-01-02 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTL0298 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Amna
4000477 | 2025-01-01 | RS IKN Nusantara | Meja Operasi | SYAS/Epione | OP850XVZY00011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal, Farhat
4000476 | 2025-01-01 | RS IKN Nusantara | Lampu Operasi | STAS/Lumina | 80XWYY0009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal, Farhat
8000650 | 2024-12-31 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN800 | ASRH0028 | Service | Nilai PEEP tidak masuk setting kalibrasi. Penggantian membran valve ekspirasi | Ready to use |  | Nanda
3000355 | 2024-12-31 | RSUD Kardinah Tegal | X-ray Digital | Apelem/Camargue DR | IMA703962 | Service | Penggantian LED colimator | Ready to use |  | Imam
11000364 | 2024-12-30 | RSUP Makassar | Incubator Transport | Draeger/Isolette TI500 | DSFL0032 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
6000875 | 2024-12-30 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASLD0042 | Survey | Gas sensor fail; O2 inaccuvate, leak breathing system. Self test, cek membran piston dan ILCA3 sensor | Pending | Perlu penggantian ILCA 3 dan Breathing system piston membran | Dwi
8000610 | 2024-12-30 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0011 | Service | compliance test fail; Pemasangan service set 3 tahun dan O2 sensor | Ready to use |  | Amna
11000364 | 2024-12-30 | RSUP Makassar | Incubator Transport | Draeger/Isolette TI500 | DSTL0053 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
2000187 | 2024-12-30 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Service | Error table ET109, penggantian parameter DRV | Ready to use |  | Eddy
2000186 | 2024-12-30 | RSUD Kab. Temanggung | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/011/C0 | Survey | Network lost, pengecekan LAN CARD cpu console | Pending | Perlu penggantian LAN Card | Eddy
11000365 | 2024-12-30 | RSUP Makassar | Bubble CPAP | Draeger/Oxymixer | QM20246056 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000365 | 2024-12-30 | RSUP Makassar | Bubble CPAP | Draeger/Oxymixer | QM20246054 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000365 | 2024-12-30 | RSUP Makassar | Bubble CPAP | Draeger/Oxymixer | QM20246057 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
9000526 | 2024-12-27 | RSUD Dr. H. Chasan Boesoirie | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/089/C1 | Survey | Tray 2 tidak bisa print. Setting posisi tray | Ready to use |  | Ruslan
8000608 | 2024-12-27 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0033
20000317 | 2024-12-27 | RSUP Makassar | Laparascopy | Aesculap | 1023032 | Instalasi | Instrument set lengkap | Ready to use |  | Sujana
4000473 | 2024-12-27 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Meja error; Kalibrasi meja | Ready to use |  | Satria
110003666 | 2024-12-27 | RSUP Makassar | Mesin Anasthesia | Draeger/Primus | ASRH0038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
110003666 | 2024-12-27 | RSUP Makassar | Mesin Anasthesia | Draeger/Primus | ASRH0049 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
16000269 | 2024-12-27 | RSUD Kab. Bekasi | Bed Pediatric | SYAS/2 crank | 240526302000400041 | Survey | Handle rel bed patah | Pending | Pengajuan claim warranty | Ipan
8000608 | 2024-12-27 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0033 | Survey | Alat mati total. Cek baterai main voltage dan power peek | Pending | Disarankan penggantian service set 3 tahun | Amna
90000528 | 2024-12-24 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | 2510220561 | Survey | Error DROC, tidak bisa masuk expose; Perbaiki posisi dongle expose | Ready to use |  | Ruslan
6000874 | 2024-12-24 | RSU Wiradadi Husada | Ventilator | Draeger/Evita V300 | ASNL0230 | Survey | Device failure (6), gas supply sensor, O2 air supply, O2 sensor kalibrasi. Pengecekan keseluruhan device check, tarik data error ada masalah pada unit system | Pending | Akan dijadwalkan pengecekan lebih lanjut dengan teknisi Draeger | Dwi
14000328 | 2024-12-24 | RSUP Surabaya | Instrument | Aesculap/ELAN 4 | 7566 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Purwono
3000354 | 2024-12-24 | RSUD Pagelaran | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/136/C2 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Imam
7000654 | 2024-12-24 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN800 | ASTM0051 | Survey | Training user | Ready to use |  | Wahyu
7000654 | 2024-12-24 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN800 | ASTM0107 | Survey | Training user | Ready to use |  | Wahyu
7000654 | 2024-12-24 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Babyleo TN500 | ASTL0117 | Survey | Training user | Ready to use |  | Wahyu
7000654 | 2024-12-24 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Babyleo TN500 | ASTL0149 | Survey | Training user | Ready to use |  | Wahyu
8000607 | 2024-12-24 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHL0168 | Service | O2 sensor calibration dan battery fail. Cek proportional valve, PVT part, valve calibration, hit software, cek PATO stand alone. Device ready to use | Pending | Disarankan penggantian baterai | Amna
19000850 | 2024-12-24 | RSJ Prof Dr Soerojo Magelang | Ventilator | Draeger/Savina300 | ASMM0191 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
11000363 | 2024-12-24 | RSUP Makassar | Ventilator | Draeger/Babylog VN800 | ASTM0103 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
19000849 | 2024-12-24 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0044 | Survey | Ada kebocoran. Unit sudah dicek dan tidak ada kerusakan | Ready to use |  | Nanda
11000363 | 2024-12-24 | RSUP Makassar | Ventilator | Draeger/Babylog VN800 | ASTM0102 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
17000207 | 2024-12-23 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0498 | Survey | Joystick bermasalah, software lemot, kolimasi tidak presisi, Fan 4 mati sendiri | Pending |  | Bahar
7000648 | 2024-12-23 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN800 | ASTM0051 | Maintenance | Maintenance Warranty | Ready to use |  | Wahyu
7000648 | 2024-12-23 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN800 | ASTM0107 | Maintenance | Maintenance Warranty | Ready to use |  | Wahyu
7000648 | 2024-12-23 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Babyleo TN500 | ASTL0117 | Maintenance | Maintenance Warranty | Ready to use |  | Wahyu
7000648 | 2024-12-23 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Babyleo TN500 | ASTL0179 | Maintenance | Maintenance Warranty | Ready to use |  | Wahyu
8000602 | 2024-12-23 | RSI Sunan Kudus | Mesin Anasthesia | Draeger/Atlan A300XL | ASNL0042 | Survey | Service data alarm, flow sensor calibration req. System check, check flow sensor, penggantian water trap, sample line. | Pending | Unit perlu penggantian sample line | Amna
19000847 | 2024-12-23 | RSUD Ibu Fatmawati Surakarta | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0061 | Survey | Installasi unit demo | Ready to use |  | Nanda
8000601 | 2024-12-21 | RS Fastabiq Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASTH0060 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
4000472 | 2024-12-20 | RSUP Jayapura | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | H248L0157 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Eddy, Irfan
11000361 | 2024-12-20 | RS Regional dr. Hasri Ainun Habibie | Bubble CPAP | Draeger/Oxymixer | QM20246035 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
8000399 | 2024-12-20 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASMA0130 | Service | O2 sensor system check fail. System check, infologer, hit software, running test | Ready to use |  | Amna
6000873 | 2024-12-20 | RSUD dr. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEL0081 | Survey | Flow sensor failure, device failure, gas supply error, O2 air error. Kalibrasi O2 sensor, error check. | Pending | Unit akan dicek lebih lanjut | Dwi
9000524 | 2024-12-19 | RSUD Dr. H. Chasan Boesoirie | ESWL | Medispec/E3000 | 0922002 | Service | Penggantian busi | Ready to use |  | Ruslan
6000872 | 2024-12-19 | RSUD Kardinah (KSO) | CT Scan | Phillips | 989600656511 | Maintenance | Preventive Maintenance | Ready to use |  | Dwi
8000606 | 2024-12-19 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHL0168 | Service | O2 sensor calibration fail & baterai fail. Hit software - Zeroing PATO stand alone - O2 sensor calibration - PATO flus flow 100% dan 21 flush flow. Device not ready. Running test O2 1 hari | Pending |  | Amna
19000845 | 2024-12-19 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | Hasil shock tidak stabil. Kalibrasi quick seasoning | Ready to use | Busi diretur | Nanda
19000846 | 2024-12-19 | RSUD Budi Rahayu Magelang | Ventilator | Draeger/Evita V600 | ASTD0042 | Service | O2 callibration fail. Coba ditembakan dengan kompressor eksternal | Ready to use |  | Nanda
6000871 | 2024-12-18 | RSUD Dr. Soeselo Slawi | Lampu Operasi | SYAS/Lumina | OL9680/802401000055 | Survey | 2 Lampu room OK 8 mati, ada kebocoran air. Cek panel lampu | Pending | Ada koslet/short panel lampu, disarankan penggantian lampu | Dwi
20000354 | 2024-12-18 | RS Universitas Hasanuddin | Infant Warmer | Draeger/Resuscitaire | ASMH0034 | Survey | Pengembalian unit demo | Ready to use |  | Sujana
17000206 | 2024-12-18 | RSUD Gunung Jati Cirebon | X-ray Digital | Apelem/Camargue DR | IMA704014 | Service | Penggantian tensator spring | Ready to use |  | Bahar, Bilal
16000268 | 2024-12-18 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Service | Ada bocor angin pada dentist element di low speed. Penggantian selang ukuran 4mm pada low speed dentist element | Ready to use |  | Ipan
11000362 | 2024-12-18 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina 300 | ASJJ0404 | Survey | FIO2 fail, baterai external fail, device check | Ready to use | Disarankan penggantian baterai external | Nasrul
11000362 | 2024-12-18 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina 300 | ASJJ0408 | Survey | FIO2 fail, baterai external fail, device check | Ready to use | Disarankan penggantian baterai external | Nasrul
8000398 | 2024-12-18 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | Survey | System check failure. Check CO2Y, check valve, check breathing system | Ready to use |  | Amna
19000844 | 2024-12-18 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Service | Unit tidak bisa shocking. Cek counter, EMH sudah counter 5000x. Update lifetime shock | Ready to use | Disarankan penggantian EMH | Nanda
3000353 | 2024-12-18 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Error pergerakan meja, setting axis control | Ready to use |  | Imam
6000867 | 2024-12-17 | RS Orthopaedi Purwokerto | Ventilator | Draeger/Savina300 | ASTL0258 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Dwi
12000163 | 2024-12-17 | RSUP Prof. Dr. R. D. Kandou Manado | Pendant | Draeger/Movita | ASJD0324 | Survey | Board valve bermasalah. Part yg dikirim salah. Kembalikan ke Draeger | Pending | Unit masih belum diperbaiki dan menunggu spare part yang benar datang | Marshall
7000644 | 2024-12-17 | RSUD Merah Putih | X-ray Analog | Apelem/Camargue Analog | IMA703979 | Service | Vertical movement stuck; Penggantian tensator spring | Ready to use |  | Wahyu
8000605 | 2024-12-17 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHL0168 | Service | O2 sensor dan O2 supply system check fail. Hot software, reload software, check PATO SA. O2 supply passed. O2 sensor calibration dan battery failed | Pending |  | Amna
19000852 | 2024-12-17 | RS Ken Saras | Mesin Anasthesia | Draeger/Fabius GS Premium | ASCL0014 | Service | Mainboard control rusak. Troubleshoot dengan board lain | Pending |  | Nanda
8000604 | 2024-12-16 | RSUD RA Kartini | Ventilator | Draeger/Savina 300 | ASEL0163 | Service | Engine drive motor blower fail; Penggantian engine driver dan motor blower | Ready to use |  | Amna
18000568 | 2024-12-16 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASHD0033 | Survey | Tekanan O2 low; Pengecekan unit, konektor O2 bocor | Ready to use | Disarankan mengganti konektor O2 | Irfan R
19000853 | 2024-12-16 | RS UNS | Ventilator | Draeger/Babylog VN800 | ASNM0024 | Survey | Training user | Ready to use |  | Nanda
20000316 | 2024-12-14 | RSUP Makassar | Cabinet | SYAS | 241132923000400040 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
9000525 | 2024-12-14 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | 2510220561 | Service | Hasil foto miring, tray printer tidak bisa print; Perbaiki posisi tabung dan wall stand, setting tray | Ready to use |  | Ruslan
9000515 | 2024-12-14 | RSUD Toto Kabila | Dental Unit | Kavo/1058TM | 10205320 | Survey | air tidak lancar, pengecekan tank air. Air kotor membuat mesin rusak, disarankan pemasangan filter air bersih | Ready to use |  | Ruslan
20000353 | 2024-12-13 | RSUP Makassar | Cabinet | SYAS | 241132923000400031 | Instalasi | Installasi 12 unit. Unit berfungsi dengan baik | Ready to use |  | Sujana
19000840 | 2024-12-13 | RS UNS | Ventilator | Draeger/Babylog VN800 | ASNM0024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
11000360 | 2024-12-13 | RSUP Tadjudin Chalid | ESWL | Medispec/E3000 | 0922014 | Service | Monitor C-arm tidak bisa ON (Booting); Baterai cmos mati, penggantian baterai cmos | Ready to use |  | Nasrul
6000868 | 2024-12-13 | RSU An Ni'mah | Ventilator | Draeger/Babylog VN600 | ASSK0026 | Service | Error low pressure oxygen, cek supply oxygen central, cek input ventilator system; Drop pressure gas oxygen central | Ready to use |  | Dwi
8000603 | 2024-12-13 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHL0168 | Service | O2 sensor calibration dan O2 supply system check fail. Penggantian PATO SA, proportional valve, baterai dan O2 dan air filter, hit software, running test. Device still not ready to use | Pending | Dijadwalkan kembali pengecekan | Amna
6000870 | 2024-12-12 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASHA0095 | Service | Error software ventilator, reset system dan reprogram firmware | Pending | Akan dilakukan kunjungan dengan teknisi Draeger | Dwi
1600264 | 2024-12-12 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Service | Pengecekan dentist elemen dan kompressor, selang robek; Penggantian selang | Ready to use |  | Ipan
5000222 | 2024-12-12 | RSPPN Soedirman | Dental Unit | Kavo/1058TM | 10207194 | Service | Pengecekan dan perapihan selang pembuangan dan tabung penampungan air | Ready to use |  | Bilal
6000869 | 2024-12-12 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHM0013 | Service | Pemasangan service set 2 tahun dan distributing valve | Pending | Power pack, sensor check, apl kuning problem | Dwi
6000866 | 2024-12-12 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASHL0013 | Service | Error PATO oksigen; Penggantian PGM2 PATO | Ready to use |  | Dwi
7000639 | 2024-12-12 | RS UII | Ventilator | Draeger/Babylog VN600 | ASTM0039 | Instalasi | Unit berfungsi dengan baik | Ready to use | Aksesoris lengkap | Wahyu
8000614 | 2024-12-12 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHM0013 | Service | Pemasangan 2/2 distributing valve dan primus service set | Ready to use |  | Nanda
19000839 | 2024-12-12 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Service | Shock tidak ada dan sering loncat. Coba busi mode baru. Busi aman tp ketika diberi power 3 counter diatas 2000 shock hilang | Ready to use |  | Nanda
19000848 | 2024-12-12 | RSUD Ibu Fatmawati Surakarta | Mesin Anasthesia | Draeger/Primus | ASMF0112 | Maintenance |  | Ready to use | Perlu penggantian maintenance kit dan service set 6 thn | Nanda
527 | 2024-12-11 | RSUD Sanana | X-ray Digital | Apelem/Camargue DR | 251023012 | Survey | DAP masih belum bisa dipasang karna konektor beda ukuran | Pending |  | Ruslan
3000352 | 2024-12-11 | RSUD Provinsi NTB | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/032/C0 | Service | Penggantian PCB 240 dan PCB 215 | Ready to use |  | Imam, Satria
5000221 | 2024-12-11 | RSUP Persahabatan | Instrument | Aesculap/Abdominal Set | - | Instalasi | Instrument set lengkap | Ready to use |  | Bilal
5000221 | 2024-12-11 | RSUP Persahabatan | Instrument | Aesculap/Partus Set | - | Instalasi | Instrument set lengkap | Ready to use |  | Bilal
16000265 | 2024-12-11 | GiGi Dental House | Dental Unit | Kavo/E50 Life | 30211887 | Service | Pemasangan water tank baru | Ready to use |  | Ipan
4000471 | 2024-12-11 | RSUD Provinsi NTB | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/032/C0 | Survey | Komputer workstation gagal masuk windows | Pending | Perlu perbaikan software | Satria, Imam
6000865 | 2024-12-11 | RSUD Panembahan Senopati Bantul | Mesin Anasthesia | Draeger/Primus | ASHJ0223 | Survey | N2O drop input, cek supply gas input, cek breathing circuit, self test; Gauge regulator N2O central problem | Pending | Disarankan penggantian regulator | Dwi
6000864 | 2024-12-11 | RSUD Panembahan Senopati Bantul | Mesin Anasthesia | Draeger/Perseus A500 | ASND0096 | Survey | N2O drop input, cek supply gas input, cek breathing circuit, self test; Gauge regulator N2O central problem | Pending | Disarankan penggantian regulator | Dwi
8000609 | 2024-12-11 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ARZF0073 | Service | O2 sensor fail; Penggantian service set 3 tahun dan O2 sensor | Ready to use |  | Amna
19000838 | 2024-12-11 | RS Ken Saras | Mesin Anasthesia | Draeger/Fabius GS Premium | ASBC0054 | Service | NV ram fail, coba penggantian baterai cmos | Ready to use | Disarankan mengganti baterai cmos dan maintenance kit 3 tahun | Nanda
12000162 | 2024-12-11 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Survey | Pemindahan panel unit. Unit masih low KV | Pending | Disarankan penambahan stabilizer | Marshall
20000352 | 2024-12-10 | RSUP Dr. Wahidin Sudirohusodo | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0495 | Service | Terdapat artefak; Cek koneksi kabel dan koneksi ke generator | Pending |  | Sujana
18000570 | 2024-12-10 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette 8000 | AR35152 | Survey | Unit tidak bisa melakukan pengisian baterai; Baterai rusak | Pending |  | Irfan R
7000464 | 2024-12-10 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN800 | ASTM0051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000646 | 2024-12-10 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN800 | ASTM0107 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000653 | 2024-12-10 | RSUD Nyi Ageng Serang (KSO) | CT Scan | Siemens/Somatom | 120196 | Service | Computer workstation tidak terkoneksi dengan printer; Sesuaikan IP komputer dengan printer | Ready to use |  | Wahyu
15000270 | 2024-12-10 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Service | Pendampingan uji coba ke pasien | Ready to use |  | Sutan
6000863 | 2024-12-10 | RS Orthopaedi Purwokerto | Ventilator | Draeger/Babylog VN800 | ASPC0031 | Survey | Demo unit | Ready to use |  | Dwi
8000611 | 2024-12-10 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTL0299 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000611 | 2024-12-10 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTL0302 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000611 | 2024-12-10 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTK0258 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000611 | 2024-12-10 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTL0300 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000611 | 2024-12-10 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTL0301 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000611 | 2024-12-10 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 Select | ASTL0298 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
19000837 | 2024-12-10 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0150 | Survey | Alarm failure, baterry internal failure. Tarik error. Alarm modul, baterai internal rusak | Pending |  | Nanda
19000816 | 2024-12-09 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0180 | Survey | System leak; Balon bagging bocor | Pending |  | Nanda
19000816 | 2024-12-09 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHH0122 | Survey | System leak; Balon bagging bocor | Pending |  | Nanda
19000816 | 2024-12-09 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0179 | Survey | System leak; Balon bagging bocor | Pending |  | Nanda
19000816 | 2024-12-09 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0016 | Survey | System leak; Balon bagging bocor | Pending |  | Nanda
18000561 | 2024-12-09 | RSUD Samratulangi Tondano | Ventilator | Draeger/Savina300 | ASJM0392 | Service | Touchscreen tidak berfungsi; Pengecekan dan pembersihan kable konektor LCD | Ready to use |  | Irfan R
7000638 | 2024-12-09 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Evita V300 | ASPC0205 | Survey | Pengecekan rutin, cek kondisi filter - mikrofilter - internal baterai | Ready to use | Disarankan penggantian baterai internal | Wahyu
7000638 | 2024-12-09 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Evita V300 | ASPC0206 | Survey | Pengecekan rutin, cek kondisi filter - mikrofilter - internal baterai | Ready to use | Disarankan penggantian baterai internal | Wahyu
7000637 | 2024-12-09 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina300 | ASPJ0061 | Survey | Pengecekan rutin, cek kondisi filter - mikrofilter - internal baterai | Ready to use | Disarankan penggantian service set 2 tahun | Wahyu
7000637 | 2024-12-09 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina300 | ASPJ0063 | Survey | Pengecekan rutin, cek kondisi filter - mikrofilter - internal baterai | Ready to use | Disarankan penggantian service set 2 tahun | Wahyu
7000637 | 2024-12-09 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina300 | ASPJ0064 | Survey | Pengecekan rutin, cek kondisi filter - mikrofilter - internal baterai | Ready to use | Disarankan penggantian service set 2 tahun | Wahyu
7000637 | 2024-12-09 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina300 | ASPJ0065 | Survey | Pengecekan rutin, cek kondisi filter - mikrofilter - internal baterai | Ready to use | Disarankan penggantian service set 2 tahun | Wahyu
6000862 | 2024-12-09 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASPL0085 | Maintenance | General check, system check, running test | Ready to use |  | Dwi
6000862 | 2024-12-09 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASNH0254 | Maintenance | General check, system check, running test | Ready to use |  | Dwi
15000269 | 2024-12-09 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Service | Penggantian water pressure | Ready to use |  | Sutan
16000267 | 2024-12-09 | RSPPN Soedirman | Dental Unit | Kavo/1058TM | 10207195 | Survey | Pendampingan pengecekan dental unit | Ready to use |  | Ipan
17000200 | 2024-12-06 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Service | Meja tidak bergerak, cek DRV | Ready to use |  | Bahar
8000396 | 2024-12-06 | RSUD RA Kartini | Tiang Infus | SYAS | - | Instalasi | Installasi 30 unit dan alat berfungsi dengan baik | Ready to use |  | Amna
4000470 | 2024-12-06 | RS Bhayangkara Anton Soedjarwo | Kursi Roda | SYAS | 240531313000400963 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000469 | 2024-12-06 | RS Bhayangkara Anton Soedjarwo | Infus Pump | Bbraun/Compact Plus P | 050033 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000469 | 2024-12-06 | RS Bhayangkara Anton Soedjarwo | Infus Pump | Bbraun/Compact Plus P | 050038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000469 | 2024-12-06 | RS Bhayangkara Anton Soedjarwo | Infus Pump | Bbraun/Compact Plus P | 050047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000469 | 2024-12-06 | RS Bhayangkara Anton Soedjarwo | Infus Pump | Bbraun/Compact Plus P | 050055 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000468 | 2024-12-06 | RS Bhayangkara Anton Soedjarwo | Syringe Pump | Bbraun/Compact Plus P | 102570 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000468 | 2024-12-06 | RS Bhayangkara Anton Soedjarwo | Syringe Pump | Bbraun/Compact Plus P | 102538 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000468 | 2024-12-06 | RS Bhayangkara Anton Soedjarwo | Syringe Pump | Bbraun/Compact Plus P | 102555 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
16000266 | 2024-12-06 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Survey | Pengecekan handpiece | Ready to use |  | Ipan
11000354 | 2024-12-06 | RSUD Andi Makassau | Ventilator | Draeger/Babylog VN800 | ASSF0003 | Service | FIO2 failure | Ready to use | Unit bisa digunakan, namun direkomendasikan untuk penggantian maintenance kit | Nasrul
18000599 | 2024-12-06 | RS Mata Sulawesi Utara | Mesin Anasthesia | Draeger/Perseus A500 | ASJM0178 | Maintenance | Preventice Maintenance Total Care | Ready to use |  | Irfan R
4000468 | 2024-12-06 | RS Bhayangkara Anton Soedjarwo | Syringe Pump | Bbraun/Compact Plus P | 102554 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
3000351 | 2024-12-06 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Imam
8000395 | 2024-12-05 | RSUD Kayen Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASSH0043 | Survey | General check | Ready to use |  | Amna
7000640 | 2024-12-05 | RSJ Grhasia | Ventilator | Draeger/Evita V600 | ASTM0008 | Instalasi | Unit berfungsi dengan baik | Ready to use |  | Wahyu
6000860 | 2024-12-05 | RSU An Ni'mah | Ventilator | Draeger/Babylog VN600 | ASSK0026 | Survey | Baterai test, penggantian neo natal flow sensor | Ready to use |  | Dwi
6000861 | 2024-12-05 | RSU An Ni'mah | Bubble CPAP | Draeger/Oxymixer | QM20215459 | Survey | Pengecekan flow high, flow low, konsentrasi sensor O2, cek breathing circuit, humidifier dan aksesoris | Ready to use |  | Dwi
19000835 | 2024-12-05 | RSJ Prof Dr Soerojo Magelang | Patient Monitor | Draeger/Vista 120 | VTSKF0078 | Survey | Unit stuck pada logo draeger, reload software masih belum bisa; Mainboard rusak | Pending |  | Nanda
7000636 | 2024-12-05 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Service | Shockwave jumping; Bersihkan konektor busi. Konektor dibawa untuk dibubut | Ready to use |  | Wahyu
18000569 | 2024-12-04 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0181 | Survey | Flow sensor tidak terbaca; Kabel herner rusak |  |  | Irfan R
11000353 | 2024-12-04 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASAM0044 | Survey | Lisensi terhapus; Oring piston rusak | Pending | Disarankan penggantian service set 3 tahun | Wahyu
7000652 | 2024-12-04 | RSUD Nyi Ageng Serang (KSO) | CT Scan | Siemens/Somatom | 120196 | Survey | Training penggunaan | Ready to use | Unit KSO | Wahyu
16000262 | 2024-12-04 | GiGi Dental House | Dental Unit | Kavo/1058TM | 1002854 | Service | Penggantian 1 set arm rest | Ready to use |  | Ipan
5000218 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Infus Pump | Bbraun/Compact Plus P | 11078 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000218 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Infus Pump | Bbraun/Compact Plus P | 11087 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000218 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Infus Pump | Bbraun/Compact Plus P | 11088 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000218 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Infus Pump | Bbraun/Compact Plus P | 11094 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000218 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Infus Pump | Bbraun/Compact Plus P | 11097 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000218 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Infus Pump | Bbraun/Compact Plus P | 11098 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000218 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Infus Pump | Bbraun/Compact Plus P | 11100 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000218 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Infus Pump | Bbraun/Compact Plus P | 11103 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000218 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Infus Pump | Bbraun/Compact Plus P | 11105 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000218 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Infus Pump | Bbraun/Compact Plus P | 11111 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000217 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Syringe Pump | Bbraun/Compact Plus P | 102527 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000217 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Syringe Pump | Bbraun/Compact Plus P | 102534 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000217 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Syringe Pump | Bbraun/Compact Plus P | 102541 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000217 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Syringe Pump | Bbraun/Compact Plus P | 102559 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000217 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Syringe Pump | Bbraun/Compact Plus P | 102579 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000217 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Syringe Pump | Bbraun/Compact Plus P | 102600 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000217 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Syringe Pump | Bbraun/Compact Plus P | 102610 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000217 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Syringe Pump | Bbraun/Compact Plus P | 102611 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000217 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Syringe Pump | Bbraun/Compact Plus P | 102617 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000217 | 2024-12-04 | RSUD Datu Kandang Haji Balangan | Syringe Pump | Bbraun/Compact Plus P | 102618 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
16000261 | 2024-12-04 | RSPPN Soedirman | Dental Unit | Kavo/1058TM | 10207194 | Service | Penggantian mainboard | Ready to use |  | Ipan
19000834 | 2024-12-04 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0089 | Survey | Gas supply unit fail; Reload software, blower sudah bisa digunakan | Ready to use |  | Nanda
9000522 | 2024-12-04 | RSUD Tobelo | Ventilator | Draeger/Babylog VN800 | ASTF0020 | Survey | FIO2 low; Cek ISO PIA O2 sensor. Unit bisa digunakan jika saturasi 95% | Ready to use |  | Ruslan
9000523 | 2024-12-04 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0011 | Survey | PEEP tinggi; Pengecekan M.1.3 sensor | Ready to use |  | Ruslan
9000523 | 2024-12-04 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0012 | Survey | PEEP tinggi; Pengecekan M.1.3 sensor | Ready to use |  | Ruslan
6000859 | 2024-12-04 | RS Dadi Keluarga | Ventilator | Draeger/Babylog VN800 | ASPF0030 | Survey | O2 supply failed, baterai low, flow sensor error; System test, running test. Penggantian flow sensor dan breathing circuit | Ready to use |  | Dwi
12000157 | 2024-12-04 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Service | PCB tidak berfungsi; Mengganti fuse pcb | Ready to use |  | Marshall
20000307 | 2024-12-03 | RSUP Dr. Wahidin Sudirohusodo | Overbed Table | SYAS | - | Instalasi | Installasi sebanyak 102 unit | Ready to use |  | Sujana
20000307 | 2024-12-03 | RSUP Dr. Wahidin Sudirohusodo | Cabinet | SYAS | - | Instalasi | Installasi sebanyak 102 unit | Ready to use |  | Sujana
8000394 | 2024-12-03 | RS Mitra Bangsa Pati | Baby Incubator | Draeger/Babylog VN600 | ASTB0024 | Service | Breathing circuit check fail; Pengecekan expiratory valve, connectivity breathing system | Ready to use |  | Amna
11000352 | 2024-12-03 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | EMH tidak bisa mengembang; Penggantian linear actuator | Ready to use | Direkomendasikan penggantian water chamber dan terminal board | Wahyu
20000309 | 2024-12-03 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASAM0044 | Survey | Ventilator fail; Pengecekan unit, license hilang | Pending |  | Sujana
6000858 | 2024-12-03 | RSUD Prof. Dr. Margono Soekarjo | Patient Monitor | Draeger/Infinity Delta XL | 6006966670 | Survey | Pembacaan SPO2, EKG dan NIBP tidak akurat; Penggantian aksesoris | Pending |  | Dwi
15000268 | 2024-12-02 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Service | Penggantian water pressure sensor. | Pending | Saat uji coba, membran EMH pecah | Sutan
4000467 | 2024-12-02 | RSUP Jayapura | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | H248L0157 | Pra-Install | Alat sudah diinsya | Ready to use |  | Satria, Imam, Bahar, Richard
19000833 | 2024-12-02 | RSUD Dr. Soediran Mangun Sumarso | Patient Monitor | Draeger/Vista 120 CMS | 35297369247 | Survey | central monitor tidak konek ke patient monitor;ganti IP | Ready to use |  | Nanda
20000306 | 2024-12-02 | RSUD Andi Makassau | ESWL | Medispec/EM1000 | 0622173 | Service | Unit c-arm tidak booting; Pengecekan cpu unit c-arm | Ready to use |  | Sujana
6000857 | 2024-12-02 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0048 | Service | O2 sensor fail, penggantian O2 sensor baru | Ready to use |  | Dwi
12000195 | 2024-11-30 | RSUP Jayapura | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | H248L0157 | Instalasi | Penginstallan unit | Pending | Belum ada listrik, belum bisa uji fungsi dan training user | Marshall
11000350 | 2024-11-30 | RSUP Tadjudin Chalid | Mesin Anasthesia | Draeger/Primus | ASJF0037 | Survey | Ventilator valve failed; Pengecekan VGC, system check | Pending | Disarankan penggantian ekectrovalve MV3 dan MV2 | Nasrul
20000308 | 2024-11-30 | RSUP Dr. Wahidin Sudirohusodo | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0495 | Service | Unit fail; Pengecekan koneksi kabel | Pending | Unit perlu penggantian kabel koneksi | Sujana
18000595 | 2024-11-29 | RSUD Ainun Habibie | Patient Monitor | Draeger/C700 | ASKM0926 | Survey | Training user | Ready to use |  | Irfan
11000351 | 2024-11-29 | RSUD Tenriawaru Bone | X-ray Digital | Apelem/Camargue DR | CPD26084L16 | Survey | Ceilling keras saat naik turun, komputer lemot; Pembersihan memori | Pending | Disarankan penggantian tensator spring | Nasrul
6000856 | 2024-11-29 | RSUD Ajibarang | Bubble CPAP | Draeger/Oxymixer | QM20220254 | Survey | Tidak masuk range kalibrasi O2, pengecekan running test flow high and low O2 | Pending |  | Dwi
7000645 | 2024-11-29 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Babyleo TN500 | ASTL0117 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000645 | 2024-11-29 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Babyleo TN500 | ASTL0149 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
19000832 | 2024-11-28 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0089 | Service | Gas supply fail; coba ganti blower backup masih belum bisa | Pending | Unit sementara belum bisa | Nanda
8000393 | 2024-11-28 | RSUD RA Kartini | Ventilator | Draeger/Babylog VN500 | ASMK0029 | Service | O2 sensor calibration fail | Ready to use | Unit bisa digunakan namun O2 sensor perlu diganti | Amna
19000831 | 2024-11-28 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRH0031 | Survey | Unit sering rusak; | Pending | Disarankan adanya udara tekan central agar unit tidak mudah rusak | Nanda
6000855 | 2024-11-28 | RSI Banjarnegara | Ventilator | Draeger/Savina300 | ASSE0144 | Survey | Humidifier F&P mr850 error, pengecekan seluruh unit; Humidifier error | Pending |  | Dwi
6000855 | 2024-11-28 | RSI Banjarnegara | Ventilator | Draeger/Savina300 | ASSE0143 | Survey | Humidifier F&P mr850 error, pengecekan seluruh unit; Humidifier error | Pending |  | Dwi
8000392 | 2024-11-26 | RSUD Batang | Mesin Anasthesia | Draeger/Atlan A300 | ASTH0153 | Service | Ventilator fail; System check, cek piston, cek diapram piston | Ready to use |  | Amna
15000267 | 2024-11-26 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Expose gambar tidak tampil;Penggantian tombol expose dan kalibrasi | Ready to use |  | Sutan
15000267 | 2024-11-26 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Expose gambar tidak tampil;Penggantian tombol expose dan kalibrasi | Ready to use |  | Sutan
18000567 | 2024-11-25 | RSUD Sele Be Solu | UPS | Yoshiga/160 GPI | E2408CO12260721368 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
6000853 | 2024-11-25 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0048 | Survey | Apnea pressure dan O2 fail, cek O2 sensor dan cek kebocoran; Sensor O2 expired | Pending | Perlu penggantian O2 sensor | Dwi
6000854 | 2024-11-25 | RSUD Banyumas | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0075 | Survey | Baterai low, gas sensor failure; Cek baterai, running test dan cek kebocoran | Pending | Disarankan penggantian baterai pack kit maintenance breathing system dan gas modil ILCA3 | Dwi
7000632 | 2024-11-25 | RSUD Muntilan | Ventilator | Draeger/Babylog VN800 | ASRF0013 | Service | Layar mati; Cek error, cek koneksi layar; Kencangkan semua sambungan ke layar | Ready to use |  | Wahyu
8000391 | 2024-11-25 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0138 | Service | O2 sensor vail, VT not detected; System cek. Perlu penggantian O2 sensor dan flow sensor serta waterlock | Pending |  | Amna
8000391 | 2024-11-25 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0161 | Service | O2 sensor vail, VT not detected; System cek. Perlu penggantian O2 sensor dan flow sensor serta waterlock | Pending |  | Amna
8000391 | 2024-11-25 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0120 | Service | O2 sensor vail, VT not detected; System cek. Perlu penggantian O2 sensor dan flow sensor serta waterlock | Pending |  | Amna
8000391 | 2024-11-25 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Primus | ARYJ0111 | Service | O2 sensor vail, VT not detected; System cek. Perlu penggantian O2 sensor dan flow sensor serta waterlock | Pending |  | Amna
19000830 | 2024-11-25 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0089 | Survey | Gas supply dan nebulizer fail; System check, error check. Blower rusak | Pending | Disarankan penggantian blower | Nanda
19000829 | 2024-11-25 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus | ASPH0124 | Survey | Layar tidak responsif; Cek mainboard, reload software. Layar touchscreen rusak | Pending | Disarankan penggantian layar touchscreen | Nanda
20000315 | 2024-11-25 | RSUP Dr. Wahidin Sudirohusodo | Overbed Table | SYAS | - | Instalasi | Installasi unit | Ready to use |  | Sujana
19000828 | 2024-11-23 | RSUD Budi Rahayu Magelang | Ventilator | Draeger/Evita V600 | ASTD0032 | Survey | Training user | Ready to use |  | Nanda
19000827 | 2024-11-22 | RSUP dr. Kariadi | Dental Unit | Kavo/1058TM | 1002722 | Service | Lampu mati; Fuse putus, penggantian fuse | Ready to use |  | Nanda
7000634 | 2024-11-22 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Maintenance | Total care Maintenance | Ready to use |  | Wahyu
7000633 | 2024-11-22 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Maintenance | Total care Maintenance | Ready to use |  | Wahyu
8000390 | 2024-11-22 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASJE0013 | Service | Bunyi kipas agak kasar; Cek kipas, bersihkan | Ready to use |  | Amna
6000852 | 2024-11-22 | RSUD CIlacap | Bubble CPAP | Draeger/Oxymixer | QM20230175 | Survey | High flow tidak bisa diatur, kebocoran pada cover flow high dan low. Cek knop high low | Pending | Perlu penggantian high flow set knop dan cover low | Dwi
6000851 | 2024-11-21 | RSKBD Panti Nugroho | Ventilator | Draeger/Evita V300 | ASLA0168 | Survey | G5500 off, baterai internal cek, cek setting, cek filter dan cek error log; Baterai rusak | Ready to use | Perlu penggantian baterai internal dan kit air filter | Dwi
19000826 | 2024-11-21 | RSUD Bung Karno | Mesin Anasthesia | Draeger/Primus | ASMA0192 | Survey | System check, unit check; Unit OK | Ready to use | Disarankan penggantian service set 6 tahun dan CO2y 2 tahun | Nanda
19000826 | 2024-11-21 | RSUD Bung Karno | Mesin Anasthesia | Draeger/Primus | ASMA0195 | Survey | System check, unit check; Unit OK | Ready to use | Disarankan penggantian service set 6 tahun dan CO2y 2 tahun | Nanda
19000826 | 2024-11-21 | RSUD Bung Karno | Mesin Anasthesia | Draeger/Primus | ASMA0194 | Survey | System check, unit check; Unit OK | Ready to use | Disarankan penggantian service set 6 tahun dan CO2y 2 tahun | Nanda
19000826 | 2024-11-21 | RSUD Bung Karno | Mesin Anasthesia | Draeger/Primus | ASMA0132 | Survey | System check, unit check; Unit OK | Ready to use | Disarankan penggantian service set 6 tahun dan CO2y 2 tahun | Nanda
16000260 | 2024-11-21 | RSUD Karawang | Dental Unit | Kavo/Primus 1058 | 10205314 | Survey | Pengecekan dental unit, air tidak keluar pada handpiece. Pembersihan jalur handpiece | Ready to use | Unit normal namun arm lampu sudah los pegasnya | Ipan
20000314 | 2024-11-21 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Perbaikan EMH | Ready to use |  | Sujana
16000259 | 2024-11-20 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Survey | Tekanan handpiece kurang kencang; Pengecekan handpiece. Unit berfungsi normal | Ready to use |  | Ipan
7000635 | 2024-11-20 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Survey | Shockwave jumping; Bersihkan konektor busi. Konektor dibawa untuk dibubut | Pending |  | Wahyu
8000389 | 2024-11-20 | RS Islam Pati | Lampu Operasi | Draefer/Polaris 600 | ASRK1077 | Maintenance | System check, General check | Ready to use |  | Amna
8000389 | 2024-11-20 | RS Islam Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0114 | Maintenance | System check, General check | Ready to use |  | Amna
1900825 | 2024-11-20 | RSU Banyumanik 2 | Ventilator | Draeger/Savina300 | ASJP0110 | Survey | Leak pada breathing circuit; System check OK. | Ready to use | Unit aman, disarankan penggantian breathing circuit | Nanda
12000156 | 2024-11-19 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASHD0005 | Survey | Baterai bermasalah | Pending | Disarankan penggantian baterai dan service set 1 tahun | Marshall
1900824 | 2024-11-19 | RSUD dr. Gunawan Mangunkusumo | Ventilator | Draeger/Savina300 | ASLA0071 | Survey | Internal baterai fail | Pending | Disarankan penggantian service set + baterai internal | Nanda
7000647 | 2024-11-19 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Service | Console disconnect; Penggantian motherboard c-arm eswl | Ready to use |  | Wahyu
1900823 | 2024-11-19 | RSUD dr. Gunawan Mangunkusumo | Patient Monitor | Draeger/Vista 120 | VTSLK0025 | Survey | Laya LCD rusak | Pending | Disarankan penggantian LCD | Nanda
1900823 | 2024-11-19 | RSUD dr. Gunawan Mangunkusumo | Patient Monitor | Draeger/Vista 120 | VTSLK0031 | Survey | Laya LCD rusak | Pending | Disarankan penggantian LCD | Nanda
1900823 | 2024-11-19 | RSUD dr. Gunawan Mangunkusumo | Patient Monitor | Draeger/Vista 120 | VTSLK0071 | Survey | Laya LCD rusak | Pending | Disarankan penggantian LCD | Nanda
7000628 | 2024-11-19 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASNL0106 | Maintenance | Total care Maintenance | Ready to use |  | Wahyu
5000215 | 2024-11-19 | RSPAD Gatot Subroto | Laser Surgical | Jena Surgical | J191150377 | Survey | Fiber laser pangkalnya meleleh terbakar dan nempel pada konektor | Ready to use | Fiber sudah tidak dapat digunakan | Bilal
7000627 | 2024-11-19 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASNL0107 | Service | Pemasangan internal baterai | Ready to use |  | Wahyu
11000339 | 2024-11-19 | RSUD Kota Kendari | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/127/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Nasrul
20000297 | 2024-11-19 | RSUD Andi Djemma Masamba | Baby Incubator | Draeger/Isolette C2000 | CF32803 | Survey | Unit gagal booting | Pending | Unit disarankan penggantian mainboard MU25448 | Sujana
20000297 | 2024-11-19 | RSUD Andi Djemma Masamba | Baby Incubator | Draeger/Isolette C2000 | CF32806 | Survey | Unit gagal booting | Pending | Unit disarankan penggantian mainboard MU25448 | Sujana
12000156 | 2024-11-19 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASHD0005 | Survey | Baterai sering bermasalah | Pending | Disarankan penggantian baterai dan service set 1 tahun | Marshall
14000324 | 2024-11-18 | RSUD Sidoarjo Barat | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0755 | Service | Kolimator miring; Setting kolimator | Ready to use |  | Purwono
20000298 | 2024-11-18 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Survey | Kalibrasi unit | Ready to use |  | Sujana
11000338 | 2024-11-18 | RSUD Kota Kendari | ESWL | Medispec/E3000 | 0922003 | Service | Monitor C-arm tidak menyala; Penggantian baterai CMOS | Ready to use |  | Nasrul
8000388 | 2024-11-18 | RS PKU Muhammadiyah Jepara | Ventilator | Draeger/Babylog VN600 | ASSJ0178 | Maintenance |  | Ready to use |  | Amna
5000214 | 2024-11-18 | RSAU Dr. Esnawan Antariksa | Mobile X-ray | Apelem/Rafale | 005/18/01391 | Survey | Tidak masuk menu booting; Pengecekan board, ad ajalur yang putus | Pending | Perlu penggantian board R.2.25 | Bilal
17000199 | 2024-11-18 | RSUD Dr. Drajat Prawiranegara | ESWL | Medispec/EM1000 | 0617048 | Service | EMH tidak stabil; Tarik unit untuk pengecekan lebih lanjut | Pending |  | Bahar
12000159 | 2024-11-18 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Service | Low kv. Pemindahan panel sumber listrik | Pending | Unit masih low kv | Marshall
8000387 | 2024-11-16 | RSI Sultan Hadlirin | Ventilator | Draeger/Evita V300 | ASPH0033 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000387 | 2024-11-16 | RSI Sultan Hadlirin | Ventilator | Draeger/Evita V300 | ASPH0031 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000387 | 2024-11-16 | RSI Sultan Hadlirin | Ventilator | Draeger/Evita V300 | ASNM0051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000387 | 2024-11-16 | RSI Sultan Hadlirin | Ventilator | Draeger/Evita V300 | ASPH0032 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000387 | 2024-11-16 | RSI Sultan Hadlirin | Ventilator | Draeger/Evita V300 | ASPC0426 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000387 | 2024-11-16 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina300 | ASSD0199 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
4000466 | 2024-11-16 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Error meja; Perbaikan driver motor dan software | Ready to use |  | Satria, Imam
18000597 | 2024-11-15 | RSUD Dr. H. Chasan Boesoirie | ESWL | Medispec/E3000 | 0922002 | Service | Pemasangan standard electrode | Ready to use |  | Irfan
9000517 | 2024-11-15 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Evita V300
17000198 | 2024-11-15 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Survey | Pendampingan pasien event RS | Ready to use |  | Bahar
11000337 | 2024-11-15 | RSUP Tadjudin Chalid | Laser Surgical | Jena Surgical | J191150292 | Service | Penggantian fiber optic | Ready to use |  | Nasrul
9000516 | 2024-11-15 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Evita V300 | ASLD0171 | Survey | Air supply GS500 tidak aktif, device check system check; Kompressor kurang baik dan baterai internal bermasalah | Ready to use | Unit bisa digunakan namun direkomendasikan untuk mengganti kompressor internal/eksternal dan baterai internal | Ruslan
19000822 | 2024-11-15 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Service | Shock tidak keluar tapi counter tetap jalan; Reconnect HVPS | Ready to use |  | Nanda
11000336 | 2024-11-14 | RSUP Tadjudin Chalid | Mamografi | Metaltronica/Helianthus DBT | HELIBDT/133/C2 | Survey | Software crash | Pending |  | Nasrul
9000513 | 2024-11-14 | RSUD Ainun Habibie | Instrument | Aesculap/Hysterectomy Set | - | Instalasi | Unit lengkap | Ready to use |  | Ruslan
8000386 | 2024-11-14 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASHA0059 | Service | Device failure 140.0072; Update windows, update field unit, reload software | Pending | Perlu penggantian M16 | Amna
19000821 | 2024-11-14 | RSUD Dr. Soediran Mangun Sumarso | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/134/C2 | Service | Cahaya kolimator miring; Adjustment kolimator | Ready to use |  | Nanda
19000836 | 2024-11-13 | RSUD Tugurejo | Baby Incubator | Draeger/Caleo | ASHK0035 | Survey | Layar bergaris dan blur; Panel LCD rusak | Pending |  | Nanda
9000514 | 2024-11-13 | RSUD Toto Kabila | Lampu Operasi | Draeger/Polaris 600 | ASML1180 | Service | Arm dan kamera tidak aktif; Pengecekan power supply | Ready to use |  | Ruslan
16000258 | 2024-11-13 | RS Dr. Suyoto | Dental Unit | Kavo/Primus 1058 | 10207194 | Service | Penggantian waterblock | Ready to use |  | Ipan
7000631 | 2024-11-12 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Maintenance | PM Total Care, water filter cleaning dan refill air | Ready to use |  | Wahyu
8000385 | 2024-11-12 | RSUD Batang | Central Monitor | Draeger/Vista 120 CMS | VLSMF0146 | Service | Software central crash; Reload software | Ready to use |  | Amna
18000598 | 2024-11-12 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0111 | Service | Pembacaaan O2 unit low; O2 sentral kurang supply | Ready to use |  | Irfan R
11000359 | 2024-11-11 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0067 | Survey | Ventilator fail, O2 sensor fail, flow sensor fail; System check, device check, calibration test | Pending | Disarankan penggantian service set 3 tahun, O2 sensor dan flow sensor | Nasrul
16000257 | 2024-11-11 | GiGi Dental House | Dental Unit | Kavo/E50 Life | 30211887 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Ipan
8000384 | 2024-11-11 | RS Budi Agung | Mesin Anasthesia | Draeger/Fabius Plus | ASTK0003 | Service | Pemasangan hose PEEP Co2y | Ready to use |  | Amna
19000818 | 2024-11-11 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0150 | Survey | Alarm failure, power supply check; Power supply, baterai internal dan modul alarm rusak | Pending |  | Nanda
20000295 | 2024-11-11 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASAM0045 | Survey | Ventilator vail; O2 sensor dan mainboard rusak | Pending | Disarankan penggantian service set 3 tahun, O2 sensor dan mainboard | Sujana
18000562 | 2024-11-10 | RSUD Toto Kabila | Lampu Operasi | Draeger/Polaris 600 | ASML1180 | Service | Lampu operasi jatuh; Penggantian lampu, sementara hanya 1 dari 2 lampu yang menyala | Ready to use |  | Irfan R
9000512 | 2024-11-09 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina300 | ASTK0074 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000512 | 2024-11-09 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina300 | ASTK0073 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000512 | 2024-11-09 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina300 | ASTK0078 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000512 | 2024-11-09 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina300 | ASTK0148 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000512 | 2024-11-09 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina300 | ASTK0077 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000512 | 2024-11-09 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina300 | ASTK0076 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000512 | 2024-11-09 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina300 | ASTK0075 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
20000299 | 2024-11-08 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Penggantian EMH | Ready to use |  | Sujana
19000817 | 2024-11-08 | RS Fastabiq Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASSM0078 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
9000510 | 2024-11-08 | RS Budi Setia Langowan | Ventilator | Draeger/Savina300 | ASTK0072 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
8000383 | 2024-11-08 | RSUD RA Kartini | C-arm | Apelem/Evo R+ | 55120027813 | Survey | Pendampingan ukes dengan BPFK Solo | Ready to use |  | Amna
17000197 | 2024-11-08 | RSUD Dr. Drajat Prawiranegara | ESWL | Medispec/EM1000 | 0617048 | Service | Fungsi EMH tidak normal, suara terlalu keras, selang rapuh; Penggantian EMH dan selang | Ready to use |  | Bahar
9000510 | 2024-11-08 | RS Budi Setia Langowan | Ventilator | Draeger/Savina300 | ASTK0071 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
12000193 | 2024-11-07 | RSUD Labuha | DR | Chamargut/Apelem | IMA 704030 | Service | Pergantian UPS Console, setting ulang tube yang tidak rata, dan service PM | Ready to use |  | Marshall
19000814 | 2024-11-07 | RSUD Dr. Soediran Mangun Sumarso | Patient Monitor | Draeger/Vista 120 | VGSSF0028 | Service | Pemindahan unit ke tempat baru (tembok lebih kokoh) | Ready to use |  | Nanda
8000382 | 2024-11-07 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASNB0020 | Service | Ventilator fail; Cek piston cup, cek diapgram piston, system test | Ready to use |  | Amna
7000626 | 2024-11-07 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Maintenance | Preventive Maintenance | Ready to use |  | Wahyu
7000625 | 2024-11-06 | RSUP dr. Soeradji Tirtonegoro | Central Monitor | Draeger/Vista 120CMS | VMSPC0056 | Service | Central monitor disconnect; Atur ulang ip address dan install ulang cms | Ready to use |  | Wahyu
4000461 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9887 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000461 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9807 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000461 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9847 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000461 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9783 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000461 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9878 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000461 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9897 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000461 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9875 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000461 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9897 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000461 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9882 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000460 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9880 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000460 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9824 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000460 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9855 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000460 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9890 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000460 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9853 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000460 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9903 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000460 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9835 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000460 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9782 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000460 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9894 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000464 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9836 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000464 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9748 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000464 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9758 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000464 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9866 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000464 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9863 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000464 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9845 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000464 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9822 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000464 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9869 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000464 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9876 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000464 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9838 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000462 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9847 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000462 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9896 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000462 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9793 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000462 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9785 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000462 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9805 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000462 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9765 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000462 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9867 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000462 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9771 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000462 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9790 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000462 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9891 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000463 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9766 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000463 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9900 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000463 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9759 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000463 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9779 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000463 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9860 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000463 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9858 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000463 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9852 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000463 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9827 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000463 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9859 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000463 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 9851 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000465 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 634641 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000465 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 634647 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000465 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 634648 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000465 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 634658 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000465 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 634663 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000465 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 634685 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000465 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 634694 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000465 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 634715 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000465 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 668970 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
4000465 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Infus Pump | Bbraun/Infusomat Compact Plus P | 671859 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000213 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Instrument | Aesculap/Curetage pilatron | - | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100699 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100618 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100620 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100716 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100611 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100709 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100715 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100603 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100503 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100713 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100512 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100698 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100707 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100527 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100518 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100704 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100693 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000207 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100528 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100615 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100515 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100719 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100600 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100621 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100493 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100496 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100695 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100533 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100526 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100601 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100596 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100619 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100711 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100720 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100499 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100598 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100593 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100727 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000209 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100497 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100717 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100700 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100597 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100612 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100718 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100617 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100513 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100522 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100607 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100605 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100414 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100624 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100532 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100602 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100495 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100622 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100514 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000206 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100419 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100628 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100411 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100542 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100427 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100417 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100431 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100426 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100439 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100425 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100436 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100441 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100438 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100433 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100440 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100399 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100437 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100398 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000205 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Syringe Pump | Bbraun/Purfoser compact plus | 100403 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000212 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Instrument | Aesculap/Tubectomy set | - | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
5000211 | 2024-11-06 | RSUP Prof. dr. IGNG Ngoerah | Instrument | Aesculap/Vasectomy Set | - | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Bilal
19000808 | 2024-11-06 | RSUD Pandan Arang | Incubator Transport | Draeger/Isolette TI500 | BH12345 | Service | Alarm no voltage; Cek unit, penggantian mainboard supply | Ready to use |  | Nanda
8000381 | 2024-11-06 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASHA0095 | Survey | Failure 0769.32927; Reload software infinity installer | Pending |  | Amna
20000313 | 2024-11-05 | RSUP Tadjudin Chalid | Digital Radiografi Fluoroskopi | Apelem/Platinum | J236F0882 | Service | Connect PACS | Ready to use |  | Sujana
11000335 | 2024-11-05 | RS Regional dr. Hasri Ainun Habibie | Ventilator | Draeger/Babylog VN800 | ASSM0042 | Survey | Neoflow sensor failed; Device check, system check, penggantian flow sensor | Ready to use |  | Nasrul
8000380 | 2024-11-05 | RSUD Batang | Mesin Anasthesia | Draeger/Atlan A300 | ASTH0135 | Service | Gas delivery fault; General check, air calibration valve, check hit software, infologger | Ready to use |  | Amna
8000378 | 2024-11-04 | RSJD Dr. Amino Gondohutomo | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0145 | Service | Pemasangan maintenance kit CO2y | Ready to use |  | Amna
3000343 | 2024-11-04 | RSUD Malingping | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/131/C2 | Service | Pemasangan display RAM | Ready to use | Perlu dilakukan kalibrasi | Imam
12000194 | 2024-11-01 | RSUD Toto Kabila | Patient Monitor Router | Draeger | VLSNF0251 | Survey | Pengecekan router unit | Pending | Disarankan untuk mengganti router | Marshall
18000563 | 2024-11-01 | RSU GMIM Pancaran Kasih Manado | Ventilator | Draeger/Savina300 | ASTJ0019 | Service | Pemasangan regulator oxygen | Ready to use |  | Irfan R
19000807 | 2024-11-01 | RS Nasional Diponegoro | Infant Warmer | Draeger/Resuscitaire | MF10231 | Survey | Error 05; Cek alat, cek board SSR, kabel heater putus. Perbaikan kabel heater | Ready to use | Disarankan penggantian heater assy | Nanda
20000294 | 2024-11-01 | RSUD Morowali | Mesin Anasthesia | Draeger/Atlan A350 | ASPH0022 | Survey | Ekspirasi dan inspirasi fail; Pengecekan unit flow sensor | Pending | Perlu penggantian flow sensor dan service set 3 tahun | Sujana
20000292 | 2024-11-01 | RSUD Morowali | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/126/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Sujana
11000349 | 2024-10-31 | RS Universitas Hasanuddin | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/086/C1 | Survey | Daily check failed; Kalibrasi detektor | Ready to use |  | Nasrul
3000342 | 2024-10-31 | RSUD Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Service | PC Workstation WSI 05 trash | Ready to use |  | Imam, Ruslan
9000509 | 2024-10-31 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 10131182 | Survey | Cek kelengkapan aksesoris; Tidak ada keyboard, mouse dan karet membran | Pending | Unit tidak dinyalakan | Ruslan
8000379 | 2024-10-31 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | ASML0039 | Service | O2 sensor fail | Pending |  | Amna
8000379 | 2024-10-31 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | ASML0137 | Service | O2 sensor fail | Pending |  | Amna
19000806 | 2024-10-31 | RSUD Dr. Soehadi Prijonegoro | Ventilator | Draeger/Babylog VN800 | ASTH0035 | Survey | Pemasangan infus tray | Ready to use |  | Nanda
7000624 | 2024-10-31 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista120 | VGSTE0001 | Survey | Training user | Ready to use |  | Wahyu
11000340 | 2024-10-31 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Penggantian motor actuator | Ready to use |  | Nasrul
20000293 | 2024-10-31 | RSUD Morowali | CPAP | Draeger/Oxymixer | QM20230004 | Survey | Indikasi rusak; Pengecekan breathing circuit, cleaning unit dan calibrasi unit serta oksigen sensor. | Ready to use | Unit baik dan keluaran oksigen sesuai dengan setting | Sujana
19000820 | 2024-10-30 | RS Telogorejo | Mesin Anasthesia | Draeger/Primus | ASCL0014 | Survey | O2 monitoring tidak terbaca; System check | Pending | Disarankan penggantian sensor PATO | Nanda
14000323 | 2024-10-30 | RSUP Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0146 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000322 | 2024-10-30 | RSUP Surabaya | C-arm | Apelem/Evo R+ | 551240231085 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000321 | 2024-10-30 | RSUP Surabaya | C-arm | Apelem/Evo R+ | 551240251087 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Purwono
4000458 | 2024-10-30 | RSUP Makassar | C-arm | Apelem/Evo R+ | 551240241086 | Survey | Uji kesesuaian dan Uji paparan | Ready to use |  | Satria, Bahar, Sujana, Nasrul
12000143 | 2024-10-30 | RSUD Ainun Habibie | MoT | SYAS | - | Survey | lampu MoT tidak bisa; Reload controller via microsoft | Ready to use |  | Marshall
8000377 | 2024-10-30 | RSUD K.R.M.T Wongsonegoro | Dental Unit | Kavo/Primus 1058 | 10502318 | Survey | Display blank; Penggantian mainboard PCBA dentist 3002 dan reload software | Ready to use |  | Amna
7000623 | 2024-10-30 | RS Panti Rini | Ventilator | Draeger/Savina300 | ARSC0003 | Survey | Vt tidak tercapai; Cek flow sensor, kabel harnes, kalibrasi flow sensor dan pressure sensor | Pending | Unit perlu penggantian service set 2thn dan expiration valve | Wahyu
8000376 | 2024-10-29 | RS Budi Agung | Mesin Anasthesia | Draeger/Fabius Plus | ASTK0003 | Test Function | Uji fungsi dan training user | Ready to use |  | Amna
4000459 | 2024-10-29 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0145 | Survey | Uji kesesuaian dan Uji paparan | Ready to use |  | Satria, Bahar, Sujana, Nasrul
16000257 | 2024-10-29 | GiGi Dental House | Dental Unit | Kavo/E50 Life | 30211887 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan, Bilal
19000804 | 2024-10-29 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0124 | Survey | Layar IACS hang tidka bisa masuk system OS; Restart - reconnect cable communication | Ready to use | Unit sudah bisa dan akan dipantau dalam beberapa hari | Nanda
7000622 | 2024-10-29 | RSUD dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Atlan A300 | ASRK0304 | Service | Penggantian O2 sensor PIA2 | Ready to use |  | Wahyu
7000621 | 2024-10-28 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Service | Layar mati; Penggantian motherboard cpu c-arm | Ready to use |  | Wahyu, Imam
9000508 | 2024-10-28 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | 2510220561 | Service | Hasil gambar tidak presisi; Setting tube dan wall stand | Ready to use |  | Ruslan
16000253 | 2024-10-28 | Signature Dental Care | Dental Unit | Kavo/Primus 1058 | 10207194 | Service | Pengembalian board | Ready to use |  | Ipan
19000811 | 2024-10-28 | RSUD Ibu Fatmawati Surakarta | Hospital Furniture | SYAS | - | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
11000328 | 2024-10-27 | RSUD Konawe | Ventilator | Draeger/Babylog VN500 | ASND0073 | Survey | Air supply indikator merah, FIO2 error; Device check - unit bisa digunakan dalam keadaan tertentu | Ready to use | Disarankan penggantian oksigen sensor PATO | Nasrul
17000196 | 2024-10-26 | RSUD Gabriel Manek | Digital Radiografi Fluoroskopi | Apelem/Platinum | E226F0723 | Service | Artefak pada gambar hasil expose; Kalibrasi detektor | Pending | Disarankan penggantian detektor | Bahar, Satria
15000266 | 2024-10-25 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Penggantian keypad | Ready to use |  | Sutan
8000375 | 2024-10-25 | RS Budi Agung | Mesin Anasthesia | Draeger/Fabius Plus | ASTK0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
11000344 | 2024-10-25 | RSUD Lakipadada | Ventilator | Draeger/Babylog VN800 | ASTE0042 | Survey | Training user | Ready to use |  | Nasrul
20000300 | 2024-10-25 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Babylog VN600 | ASMH2450 | Survey | Penarikan unit demo |  |  | Sujana
3000341 | 2024-10-25 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Backup software PC Workstation | Ready to use |  | Imam
12000160 | 2024-10-24 | RSUD ODSK Provinsi Sulawesi Utara | Baby Incubator | Draeger/Babyleo TN500 | ASTH0041 | Survey | Pengecekan HCA | Ready to use |  | Marshall
7000619 | 2024-10-24 | RSJD Dr. RM. Soedjarwadi | Mesin Anasthesia | Draeger/Atlan A350 | ASPH0045 | Survey | Gas monitoring faulty; Cek error log, penggantian sampling line dan water lock | Ready to use |  | Wahyu
7000620 | 2024-10-24 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista120 | VGSTE0001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000620 | 2024-10-24 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista120 | VGSTE0021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
11000327 | 2024-10-24 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Penggantian motor actuator | Ready to use |  | Nasrul
16000251 | 2024-10-24 | RSUD K.R.M.T Wongsonegoro | Dental Unit | Kavo/Primus 1058 | 10205318 | Survey | Display pada dentist element blank | Pending | Disarankan penggantian mainboard dentist element | Ipan
8000374 | 2024-10-23 | RSUD Kayen Pati | Infant Warmer | Draeger/Resuscitaire | BU14106 | Service | Saklar on/off valve S-way fail; Cleaning valve | Ready to use |  | Amna
3000339 | 2024-10-23 | RSUD Dr. Drajat Prawiranegara (KSO) | Cathlab | Siemens/Artis ONE | 82377 | Service | Ganti mccb dan kontaktor panel | Ready to use |  | Imam
3000340 | 2024-10-23 | RSUD Sumedang | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/132/C2 | Service | Daily check failed; Kalibrasi detektor | Ready to use |  | Imam
9000519 | 2024-10-22 | RSUD Sanana | X-ray Digital | Apelem/Camargue DR | 251023012 | Service | Setting rem roda ceilling | Ready to use |  | Ruslan
16000252 | 2024-10-22 | Signature Dental Care | Dental Unit | Kavo/Primus 1058 | 10207194 | Service | Peminjaman board untuk troubleshoot |  |  | Ipan
11000334 | 2024-10-22 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina300 | ASJL0396 | Survey | Expirasi flow sensor failed; Device check, kabel harnes rusak | Pending | Disarankan penggantian service set 6 thn | Nasrul
4000457 | 2024-10-22 | RSUD Bangil | C-arm | Apelem/Evo R+ | 55121016912 | Survey | Uji kesesuaian | Ready to use |  | Satria, Bahar, Purwono
12000140 | 2024-10-22 | RSUD Kotamobagu | C-arm | Apelem/Evo R+ | 55118031654 | Survey | Power supply error, blue screen; HDD rusak terkena air | Pending | Disarankan penggantian HDD dan reload software | Marshall
8000373 | 2024-10-22 | RSI Harapan Anda | Ventilator | Draeger/Babylog VN500 | ASJC0050 | Service | System check nebulizer fail; Penggantian pressure regulator | Ready to use |  | Amna
7000618 | 2024-10-22 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Survey | Layar c-arm mati; Indikasi kerusakan motherboard CPU C-arm | Pending | Pengajuan claim warranty | Wahyu
7000617 | 2024-10-22 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Survey | Unit mati; Cek koneksi kabel, fuse. Indikasi rusak pada power supply | Pending | Pengajuan claim warranty | Wahyu
9000518 | 2024-10-21 | RSUD Dr. H. Chasan Boesoirie | Mesin Anasthesia | Draeger/Perseus A500 | ASSK0089 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00067 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00064 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00049 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00062 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00061 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00068 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00046 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00035 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00031 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000305 | 2024-10-21 | RSUP Makassar | Meja Operasi | SYAS/Lumina | OL9680/80XVU00014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
11000345 | 2024-10-21 | RS Muhammadiyah Makassar | CPAP | Draeger/Oxymixer | QM20215173 | Survey | Bunyi berdengung, cek pressure air dan oksigen; Penggantian tabung O2 | Ready to use |  | Nasrul
7000615 | 2024-10-21 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | VGSSF0082 | Service | Pemasangan braket monitor | Ready to use |  | Wahyu
7000615 | 2024-10-21 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | VGSSF0105 | Service | Pemasangan braket monitor | Ready to use |  | Wahyu
15000265 | 2024-10-21 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Unit tidak bisa digunakan; Tombol emergency terkekan, sudah diperbaiki | Ready to use |  | Sutan
4000456 | 2024-10-21 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/EM1000 | 0720015 | Survey | Uji kesesuaian | Ready to use |  | Satria, Bahar, Purwono
19000805 | 2024-10-20 | RSUP dr. Kariadi | Baby Incubator | Draeger/Isolette C2000 | CR35569 | Service | Laya tombol keyboard tidak bisa dan pengukuran tidak akurat | Pending | Disarankan penggantian keypad rusak dan sensor module rusak | Nanda
11000326 | 2024-10-19 | RSIA Sentosa | Ventilator | Draeger/Babylog VN800 | ASRN0085 | Service | Alarm air supply pressure low; Kompresor berair, di bersihkan | Ready to use |  | Nasrul
11000332 | 2024-10-18 | RSUD Lakipadada | Ventilator | Draeger/Babylog VN800 | ASTE0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
8000372 | 2024-10-18 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Babylog VN600 | ASSK0028 | Survey | Refresh training | Ready to use |  | Amna
8000372 | 2024-10-18 | RS PKU Muhammadiyah Mayong | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ0081 | Survey | Refresh training | Ready to use |  | Amna
15000264 | 2024-10-18 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Joystick tidak bisa digunakan; Penggantian joystick sementara | Ready to use |  | Sutan
9000506 | 2024-10-18 | RSUD Samratulangi Tondano | Ventilator | Draeger/Evita V300 | ASNL0255 | Service | Nilai VT tidak mencapai setting; Device check dan pembersihan valve inspirasi | Ready to use |  | Ruslan, Irfan, Marshall
11000332 | 2024-10-18 | RSUD Lakipadada | Ventilator | Draeger/Babylog VN800 | ASTE0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
20000291 | 2024-10-18 | RSUD Andi Makassau | ESWL | Medispec/EM1000 | 0623173 | Service | EMH tidak ngembang maksimal; Pengecekan tekanan dan penambahan air EMH | Ready to use |  | Sujana
19000800 | 2024-10-18 | RSUP dr. Kariadi | Dental Unit | Kavo/Primus 1058 | 1002722 | Service | Lampu dental unit mati; Penggantian lampu dan fuse | Ready to use |  | Nanda
11000333 | 2024-10-17 | RSUD Kota Makassar | Ventilator | Draeger/Evita V300 | ASNM0196 | Service | Baterai internal failure
11000333 | 2024-10-17 | RSUD Kota Makassar | Ventilator | Draeger/Evita V300 | ASNM0196 | Survey | Baterai internal failure, PS500 rusak | Pending | Perlu disarankan penggantian baterai PS500 | Nasrul
19000799 | 2024-10-17 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | QC fail, setting alat | Ready to use |  | Nanda
7000616 | 2024-10-17 | RSUD dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0091 | Survey | Nebulizer fail; Cek error log, cek pressure regulator | Pending | Pengajuan claim warranty pressure regulator | Wahyu
7000614 | 2024-10-16 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0011 | Survey | Inspiration valve tidak attach; Cek error, bersihkan inspiration valve | Ready to use | Disarankan penggantian A temsy vbar 2year | Wahyu
7000613 | 2024-10-16 | RSUD Ajibarang | Infant Warmer | Draeger/Resuscitaire | DSNJ0048 | Service | Error 05; Pemasangan heater assy | Ready to use |  | Wahyu
4000454 | 2024-10-16 | RSUD Merah Putih | X-ray Analog | Apelem/Camargue Analog | IMA703979 | Service | Penggantian tube handle | Ready to use |  | Satria, Bahar
4000455 | 2024-10-16 | RSUD Merah Putih | X-ray Analog | Apelem/Camargue Analog | IMA703979 | Service | Pengecekan ceilling; Tensator Spring rusak | Pending | Perlu dilakukan penggantian tensator spring | Satria, Bahar
8000371 | 2024-10-16 | RSU Sebening Kasih | Meja Operasi | SYAS | 240852512100400017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
3000338 | 2024-10-16 | RSUD Dr. Drajat Prawiranegara (KSO) | Cathlab | Siemens/Artis ONE | 82377 | Service | Penggantian DVD RW | Ready to use |  | Imam
12000155 | 2024-10-15 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0410 | Survey | Power supply error. Troubleshoot dengan PSU unit lain, alat berfungsi | Pending | Disarankan penggantian PSU | Marshall
11000358 | 2024-10-15 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Error thermal sensor; Pembresihan terminal board | Ready to use |  | Nasrul
19000785 | 2024-10-15 | RSUD Dr. Soediran Mangun Sumarso | Patient Monitor | Draeger/Vista120 | VGSSF0029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
8000370 | 2024-10-15 | RSUD Kayen Pati | Infant Warmer | Draeger/Resuscitaire | GV14025 | Service | Penggantian valve control mini s- way | Ready to use |  | Amna
19000785 | 2024-10-15 | RSUD Dr. Soediran Mangun Sumarso | Patient Monitor | Draeger/Vista120 | VGSSF0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000785 | 2024-10-15 | RSUD Dr. Soediran Mangun Sumarso | Patient Monitor | Draeger/Vista120 | VGSSF0026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000785 | 2024-10-15 | RSUD Dr. Soediran Mangun Sumarso | Patient Monitor | Draeger/Vista120 | VGSSF0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
7000611 | 2024-10-15 | RSU Wiradadi Husada | Ventilator | Draeger/Evita V300 | ASNL0235 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
19000795 | 2024-10-14 | RSU Banyumanik 2 | Ventilator | Draeger/Savina300 | ASPJ0113 | Survey | Katup pengamanan fail; Cek unit - bypass circuit, ada kerusakan pada testlung yang tidak lentur | Ready to use |  | Nanda
19000795 | 2024-10-14 | RSU Banyumanik 2 | Ventilator | Draeger/Savina300 | ASPJ0110 | Survey | Katup pengamanan fail; Cek unit - bypass circuit, ada kerusakan pada testlung yang tidak lentur | Ready to use |  | Nanda
18000551 | 2024-10-14 | RSUD ODSK Provinsi Sulawesi Utara | Baby Incubator | Draeger/Babyleo TN500 | ASTH0041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
18000551 | 2024-10-14 | RSUD ODSK Provinsi Sulawesi Utara | Baby Incubator | Draeger/Babyleo TN500 | ASTH0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
8000369 | 2024-10-14 | RSI Sultan Hadlirin | Ventilator | Draeger/Babylog VN600 | ASTD0005 | Survey | Training | Ready to use |  | Amna
19000796 | 2024-10-14 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN800 | ASRH0028 | Service | Nebulizer fail; Penggantian pressure regulator | Ready to use | Claim warranty | Nanda
19000797 | 2024-10-14 | RS Nasional Diponegoro | Infant Warmer | Draeger/Resuscitaire | AF10306 | Service | Bunyi dengung dari gas supply; Ada indikasi pada pressure regulator yg rusak dan saklar on/off gas | Ready to use | Unit sementara bisa digunakan | Nanda
20000304 | 2024-10-14 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Penggantian motor actuator | Ready to use |  | Sujana
9000500 | 2024-10-11 | RSU GMIM Pancaran Kasih Manado | Ventilator | Draeger/Savina300 | ASTJ0018 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000500 | 2024-10-11 | RSU GMIM Pancaran Kasih Manado | Ventilator | Draeger/Savina300 | ASTJ0020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000500 | 2024-10-11 | RSU GMIM Pancaran Kasih Manado | Ventilator | Draeger/Savina300 | ASTJ0021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000500 | 2024-10-11 | RSU GMIM Pancaran Kasih Manado | Ventilator | Draeger/Savina300 | ASTJ0019 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
8000368 | 2024-10-11 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0010 | Service | Software corrupt; Reload software - system check | Ready to use |  | Amna
3000337 | 2024-10-11 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Survey | Tegangan drop | Ready to use | Disarankan menggunakan stabilizer | Imam
9000499 | 2024-10-10 | RS GMIM Bethsaida Tomohon | Ventilator | Draeger/Savina300 | ASTJ0023 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000499 | 2024-10-10 | RS GMIM Bethsaida Tomohon | Ventilator | Draeger/Savina300 | ASTJ0022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000499 | 2024-10-10 | RS GMIM Bethsaida Tomohon | Ventilator | Draeger/Savina300 | ASTJ0024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000499 | 2024-10-10 | RS GMIM Bethsaida Tomohon | Ventilator | Draeger/Savina300 | ASTJ0025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
4000453 | 2024-10-10 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0917012 | Survey | ESWL tidak dapat menembak, kalibrasi gagal | Pending | Unit belum dapat digunakan | Satria, Bahar
3000336 | 2024-10-10 | RSUD Gunung Jati Cirebon | X-ray Digital | Apelem/Camargue DR | IMA704014 | Service | Penggantian tensator spring | Ready to use |  | Imam, Bahar
19000795 | 2024-10-10 | RSU Banyumanik 2 | Ventilator | Draeger/Savina300 | ASPJ0113 | Service | Katup pengamanan fail; Cek unit - bypass circuit, ada kerusakan pada testlung yang tidak lentur | Ready to use |  | Nanda
19000795 | 2024-10-10 | RSU Banyumanik 2 | Ventilator | Draeger/Savina300 | ASPJ0110 | Service | Katup pengamanan fail; Cek unit - bypass circuit, ada kerusakan pada testlung yang tidak lentur | Ready to use |  | Nanda
7000610 | 2024-10-10 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Service | Unit double shockwave; Kalibrasi HVG, perbaiki koneksi HVG | Ready to use |  | Wahyu
7000609 | 2024-10-09 | RSUD Sleman | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0128 | Survey | O2 gas supply fail; Cek error log, adjust software | Ready to use | Unit berfungsi dengan terbatas. Disarankan penggantian pba optimix status display | Wahyu
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00067 | Pra-Install |  |  |  | Sujana
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00064 | Pra-Install |  |  |  | Sujana
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00049 | Pra-Install |  |  |  | Sujana
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00037 | Pra-Install |  |  |  | Sujana
12000141 | 2024-10-09 | RSUD Kotamobagu | C-arm | Apelem/Evo R+ | 55118031654 | Survey | Power supply error; Menambahkan power board, trafo, PSU 24V untuk pengecekan | Pending |  | Marshall
15000262 | 2024-10-09 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Platinum mati total; Pengecekan PU BOX dan Generator | Ready to use |  | Sutan
8000367 | 2024-10-09 | RS Aisyiyah Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0028 | Service | O2 sensor fail; O2 sensor expired | Pending |  | Amna
19000794 | 2024-10-09 | RSUD Simo Boyolali | Mesin Anasthesia | Draeger/Fabius Plus | ASTC0004 | Survey | Pendampingan kalibrasi | Ready to use |  | Nanda
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00062 | Pra-Install |  |  |  | Sujana
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00061 | Pra-Install |  |  |  | Sujana
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00030 | Pra-Install |  |  |  | Sujana
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00068 | Pra-Install |  |  |  | Sujana
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00046 | Pra-Install |  |  |  | Sujana
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00035 | Pra-Install |  |  |  | Sujana
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00031 | Pra-Install |  |  |  | Sujana
20000312 | 2024-10-09 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVU00014 | Pra-Install |  |  |  | Sujana
9000498 | 2024-10-09 | RS Siloam Sonder | Ventilator | Draeger/Savina300 | ASTJ0070 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan, Irfan
9000498 | 2024-10-09 | RS Siloam Sonder | Ventilator | Draeger/Savina300 | ASTJ0072 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan, Irfan
9000498 | 2024-10-09 | RS Siloam Sonder | Ventilator | Draeger/Savina300 | ASTJ0071 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan, Irfan
9000498 | 2024-10-09 | RS Siloam Sonder | Ventilator | Draeger/Savina300 | ASTJ0069 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan, Irfan
20000301 | 2024-10-08 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0145 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Sujana
9000505 | 2024-10-08 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | VGSSF0108 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000505 | 2024-10-08 | RSUD ODSK Provinsi Sulawesi Utara | Patient Monitor | Draeger/Vista 120 | VGSSF0022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000504 | 2024-10-08 | RSUD ODSK Provinsi Sulawesi Utara | Ventilator | Draeger/Babylog VN800 | ASTH0036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000496 | 2024-10-08 | RSUP Prof. Dr. R. D. Kandou Manado | Instrument | Aesculap/ELAN4 | 003976 | Service | Food switch tidak berfungsi; Penggantian food switch dan kabel handpiece | Ready to use |  | Ruslan
18000556 | 2024-10-08 | RS GMIM Kalooran Amurang | Ventilator | Draeger/Savina300 | ASJTJ0017 | Instalasi | Alat berfungsi dengan baik | Ready to use | Disarankan penggantian O2 sensor capsule | Irfan
18000556 | 2024-10-08 | RS GMIM Kalooran Amurang | Ventilator | Draeger/Savina300 | ASTJ0013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
18000556 | 2024-10-08 | RS GMIM Kalooran Amurang | Ventilator | Draeger/Savina300 | ASTJ0016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
18000556 | 2024-10-08 | RS GMIM Kalooran Amurang | Ventilator | Draeger/Savina300 | ASTJ0012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
18000553 | 2024-10-08 | RS GMIM Kalooran Amurang | Ventilator | Draeger/Savina300 | ASTJ0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
18000553 | 2024-10-08 | RS GMIM Kalooran Amurang | Ventilator | Draeger/Savina300 | ASTH0014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
18000553 | 2024-10-08 | RS GMIM Kalooran Amurang | Ventilator | Draeger/Savina300 | ASTJ0010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
18000553 | 2024-10-08 | RS GMIM Kalooran Amurang | Ventilator | Draeger/Savina300 | ASTJ0015 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
19000786 | 2024-10-07 | RSUD Budi Rahayu Magelang | Ventilator | Draeger/Evita V600 | ASTD0042 | Survey | Training user | Ready to use |  | Nanda
8000365 | 2024-10-07 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0010 | Service | Penggantian O2 sensor, mainboard PB3, service set 3 tahun | Ready to use |  | Amna
3000335 | 2024-10-07 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Penggantian detektor | Ready to use |  | Imam, Bahar
4000451 | 2024-10-07 | RSPAD Gatot Subroto | Laser Surgical | Jena Surgical | J191150377 | Service | Socket fiber tidak bisa dicabut; Pemotongan sutter yang sudah lengket | Ready to use |  | Bilal, Bahar
19000812 | 2024-10-04 | RSUP Dr. Sardjito | Project ISDB | Draeger | - | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
4000452 | 2024-10-04 | Signature Dental Care | Dental Unit | Kavo/1058TM | 10207194 | Survey | Penggantian base plat | Ready to use |  | Satria, Ipan
18000554 | 2024-10-04 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
18000554 | 2024-10-04 | RSUD Tobelo | Ventilator | Draeger/Evita V600 | ASTF0012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
8000364 | 2024-10-04 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina300 | ASTJ007 | Survey | Training user | Ready to use |  | Amna
9000497 | 2024-10-04 | RSU GMIM Pancaran Kasih Manado | Ventilator | Draeger/Savina300 | ASTJ0011 | Pra-Install | Unit belum diinsyall dan uji fungsi karna tabung O2 RSU kosong |  |  | Ruslan
3000334 | 2024-10-04 | RSUD Sumedang | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/132/C2 | Survey | Ambil part yg rusak untuk diclaim warranty | Ready to use |  | Imam
18000552 | 2024-10-03 | RSUD Tobelo | Ventilator | Draeger/Babylog VN800 | ASTF0020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
7000607 | 2024-10-03 | RS Mata Purwokerto | Mesin Anasthesia | Draegr/Fabius Plus | ASTH0001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
19000793 | 2024-10-03 | RSUD dr. Gunawan Mangunkusumo | Incubator Transport | Draeger/Isolette TI500 | KY11399 | Survey | Air flow failure; Indikasi flow fan sudah lemah | Pending |  | Nanda
3000333 | 2024-10-03 | RSUD Dr. Drajat Prawiranegara (KSO) | Cathlab | Siemens/Artis ONE | 82377 | Service | Cicruit breaker triph; Cek tegangan | Ready to use |  | Imam
9000493 | 2024-10-02 | RSUD Otohana | Ventilator | Draeger/Babylog VN800 | ASTE0041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000494 | 2024-10-02 | RSUD Otohana | Patient Monitor | Draeger/Vista 120S | 52SNM0821 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
11000331 | 2024-10-02 | RSUD Lakipadada | Mesin Anasthesia | Draeger/
8000363 | 2024-10-02 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina300 | ASTJ0006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000363 | 2024-10-02 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina300 | ASTJ0007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
11000331 | 2024-10-02 | RSUD Lakipadada | Mesin Anasthesia | Draeger/Primus | ASMJ0188 | Survey | O2 sensor fail, gas sensor fail; Self test, failed | Pending | Disarankan penggantian PATO, ILCA, Flow power, Flow sensor, waterlock, service set 2 tahun | Nasrul
1000206 | 2024-10-02 | RSUD Sultan Suriansyah | Dental Unit | Kavo/Primus 1058 | 10205321 | Survey | Pengisian cairan oxygenal | Ready to use |  | Heru
18000555 | 2024-10-01 | RSUD Weda | Bubble CPAP | Draeger/Oxymixer | MO00690 | Survey | Unit dipakai ke pasien bayi dan terjadi kebiruan pada pasien, di cek dengan flow analyser, unit dalam keadan baik | Ready to use |  | Irfan
12000138 | 2024-10-01 | RSUD Pratama Bitung | Mobile X-ray | Apelem/Rafale DR | 8112416337 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Marshall
19000792 | 2024-10-01 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN800 | ASTF0051 | Service | Breathing circuit berair; Setting humidity | Ready to use |  | Nanda
2000288 | 2024-10-01 | RS PKU Muhammadiyah Unismuh | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0127 | Service | Breathing error; Cleaning VGA dan cek flow sensor | Ready to use |  | Sujana
19000813 | 2024-10-01 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN800 | ASTF0051 | Service | Humidity berair; Setting humidifier | Ready to use |  | Nanda
15000261 | 2024-09-30 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Error ET12 longitudinal panel; motion block error, penggantian kabel LAN DRV | Ready to use |  | Sutan
18000539 | 2024-09-30 | RSUD Otohana | Ventilator | Draeger/Babylog VN800 | ASTF0045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
19000791 | 2024-09-30 | RSUD Dr. Soehadi Prijonegoro | Ventilator | Draeger/Babylog VN800 | ASTH0035 | Service | Pendampingan kalibrasi | Ready to use |  | Nanda
18000539 | 2024-09-30 | RSUD Otohana | Ventilator | Draeger/Babylog VN800 HFO | ASTH0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000539 | 2024-09-30 | RSUD Otohana | Ventilator | Draeger/Babylog VN800 HFO | ASTH0037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000539 | 2024-09-30 | RSUD Otohana | Ventilator | Draeger/Babylog VN800 HFO | ASTH0038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000539 | 2024-09-30 | RSUD Otohana | Ventilator | Draeger/Evita V600 | ASTE0093 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
20000286 | 2024-09-30 | RSUP Makassar | C-arm | Apelem/Evo R+ | 551240241086 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
11000329 | 2024-09-29 | RSUD Kota Kendari | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/127/C2 | Maintenance | Preventice Maintenance | Ready to use |  | Nasrul
11000330 | 2024-09-29 | RSUD Kota Kendari | ESWL | Medispec/E3000 | 0922003 | Maintenance | Preventice Maintenance | Ready to use |  | Nasrul
9000488 | 2024-09-28 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Babylog VN800 | ASTH0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000488 | 2024-09-28 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Babylog VN800 | ASTF0045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000488 | 2024-09-28 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Babylog VN800 | ASTH0037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000488 | 2024-09-28 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Babylog VN800 | ASTH0038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000487 | 2024-09-28 | RSUD Dr. H. Chasan Boesoirie | Ventilator | Draeger/Evita V600 | ASTE0093 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
1000205 | 2024-09-27 | RSI Sultan Agung Banjarbaru | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0126 | Service | Error leakage; remote user dengan pengecekan inspirasi dan eksperasi set teslang | Ready to use |  | Heru
7000606 | 2024-09-27 | RSUD Saras Adyatma | Mesin Anasthesia | Draeger/Atlan A300 | ASTH0115 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000606 | 2024-09-27 | RSUD Saras Adyatma | Patient Monitor | Draeger/Vista 120 | VGSTC0078 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
8000362 | 2024-09-27 | RSUD Batang | Mesin Anasthesia | Draeger/Atlan A300 | ASTH0152 | Survey | Koordinasi pemasangan stop kontak dan selang gas sentral |  |  | Amna
8000362 | 2024-09-27 | RSUD Batang | Mesin Anasthesia | Draeger/Atlan A300 | ASTH0153 | Survey | Koordinasi pemasangan stop kontak dan selang gas sentral |  |  | Amna
3000332 | 2024-09-27 | RSUD Kardinah Tegal | X-ray Analog | Apelem/Camargue Analog | IMA703962 | Survey | Lampu kolimator mati | Pending | Disarankan penggantian lampu kolimator | Imam
11000342 | 2024-09-26 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Uji fungsi motor actuator | Ready to use |  | Nasrul
20000284 | 2024-09-26 | RSUP Dr. Wahidin Sudirohusodo | Dental Unit | Kavo/Primus 1058 | 10205313 | Survey | Waterblock pressure dan leakage water | Pending | Disarankan penggantian kabel HDMI untuk video oral | Sujana
7000605 | 2024-09-26 | RS Orthopaedi Purwokerto | Ventilator | Draeger/Savina 300 | ASSM0109 | Survey | Uninstall unit |  | Unit dikirim ke kantor draeger | Wahyu
9000486 | 2024-09-26 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | 2510220561 | Service | Hasil gambar miring; Setting tube pada posisi 0 | Ready to use |  | Ruslan
3000331 | 2024-09-26 | RSUD RA Kartini | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/006/C0 | Service | time out high starter; HSS board fail | Pending | Perlu penggantian HSS Board | Imam, Satria, Amna
10000204 | 2024-09-26 | RSUD H. Boejasin | MoT | SYAS/MoT | - | Service | Penggantian pcb alarm gas medis | Ready to use |  | Heru
11000356 | 2024-09-25 | RSUP Makassar | Meja Operasi | SYASTECH/EPIONE | OP8500XWYX00006 | Survey | Installasi aksesoris | Ready to use |  | Nasrul
12000135 | 2024-09-25 | RSD Liun Kendage Tahuna | Laparascopy | Aesculap | 2407CEO312 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Marshall
20000311 | 2024-09-25 | RSUP Makassar | Meja Operasi | SYAS/Epione | OP8500XWYX0014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
7000604 | 2024-09-25 | RS Raffa Medika | Ventilator | Draeger/Savina 300 | ASTJ0009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000603 | 2024-09-25 | RS Raffa Medika | Ventilator | Draeger/Babylog VN600 | ASTE0067 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
9000484 | 2024-09-25 | RSUD Sanana | X-ray Digital | Apelem/Camargue DR | 251023012 | Service | Film tray 2 tidak keluar saat print; Setting printer | Ready to use |  | Ruslan
15000259 | 2024-09-25 | RSUD Hasanuddin Damrah | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/138/C2 | Service | Error 1226; Kalibrasi detector, daily cek. Baterai ups rusak 16pcs | Ready to use | Perlu dilakukan penggantian baterai ups | Sutan
8000361 | 2024-09-25 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHD0188 | Service | O2 supply fail; Cal gas supply | Ready to use |  | Amna
7000602 | 2024-09-25 | RS Raffa Medika | Mesin Anasthesia | Draeger/Fabius Plus | ASTD0062 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
19000809 | 2024-09-25 | RS Indriati Solo | Ventilator | Draeger/Babylog VN800 | ASPH0042 | Survey | Installasi unit demo | Ready to use |  | Nanda
9000485 | 2024-09-24 | RSUD Sanana | Dental Unit | Kavo/E30TM | 10111150 | Service | Scaller tidak berfungsi; Handpiece dan board scaller rusak | Pending | Diajukan penawaran harga | Ruslan
8000360 | 2024-09-24 | RSUD Dr. R. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | ASLJ0092 | Service | Alarm O2 sensor fail; O2 sensor rusak | Pending | Disarankan penggantian O2 sensor | Amna
8000360 | 2024-09-24 | RSUD Dr. R. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0041 | Service | Alarm O2 sensor fail; O2 sensor rusak | Pending | Disarankan penggantian O2 sensor | Amna
8000360 | 2024-09-24 | RSUD Dr. R. Soeprapto Cepu | Ventilator | Draeger/Savina 300 | ASJK0211 | Service | Baterai internal failure | Pending | Disarankan penggantian baterai internal | Amna
7000601 | 2024-09-24 | RS Duta Mulya | Ventilator | Draeger/Savina 300 | ASTJ0005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000601 | 2024-09-24 | RS Duta Mulya | Ventilator | Draeger/Savina 300 | ASTJ0004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
3000330 | 2024-09-24 | RSUD Balaraja | C-arm | Apelem/C-arm | 55118034657 | Service | software tidak running; install microsoft visual | Ready to use |  | Imam
7000599 | 2024-09-23 | RS PKU Muhammadiyah Kota Yogyakarta | Mesin Anasthesia | Draeger/Atlan A300 | ASSL0253 | Service | Penggantian kit motor ventilator | Ready to use |  | Wahyu
19000789 | 2024-09-23 | RSUD Simo Boyolali | Jaudience Meter | Draeger/JM105 | H3505875 | Service | Error 05; Penggantian baterai | Ready to use |  | Nanda
19000790 | 2024-09-23 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | Qc fail, gantry tidak bergerak; Setting unit BMD dan ups | Ready to use |  | Nanda
16000250 | 2024-09-23 | GiGi Dental House | Dental Unit | Kavo/Primus 1058 | 10207196 | Service | Waterblock bocor air; Penggantian waterblock | Ready to use |  | Ipan
3000329 | 2024-09-23 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Service | installasi ups | Ready to use |  | Imam
7000598 | 2024-09-20 | RSUD Sleman | Mesin Anasthesia | Draeger/Perseus A500 | ASTH0075 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
17000194 | 2024-09-20 | RSUD Dr. Drajat Prawiranegara | ESWL | Medispec/EM1000 | 0917011 | Survey | EMH mengeluarkan cahaya dan suara yang keras; Plat EMH rusak | Pending | EMH dibawa ke kantor untuk diperbaiki lebih lanjut | Bahar
19000788 | 2024-09-20 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN800 | ASTF0051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
8000359 | 2024-09-20 | RS Islam Pati | UPS | Yoshiga/HPT1 | 501101213200KB200041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
20000283 | 2024-09-19 | RSUD Andi Makassau | ESWL | Medispec/EM1000 | 0622173 | Service | Penggantian EMH | Ready to use |  | Sujana
18000538 | 2024-09-19 | RSUD Kotamobagu | C-arm | Apelem/Evo R+ | 55118031654 | Survey | Power supply error; Board unit sementara dibawa untuk diperiksa lebih lanjut | Pending |  | Irfan R
19000787 | 2024-09-19 | RSUD Dr. Soehadi Prijonegoro | Ventilator | Draeger/Babylog VN800 | ASTH0035 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
9000482 | 2024-09-19 | RSP Fam Dofa | Finiset bed | SYAS/SBF5 | 230774007000400014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
8000358 | 2024-09-19 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina300 | ASTJ0006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000358 | 2024-09-19 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina300 | ASTJ0007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
7000597 | 2024-09-18 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASAC0044 | Survey | Auxilary alarm fault; Auxilarry alarm rusak | Pending | Proses pengajuan claim warranty | Wahyu
9000466 | 2024-09-18 | RSUD Fam Dofa | Mobile X-ray | Apelem/Rafale DR | 1931312200003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000467 | 2024-09-18 | RSUD Fam Dofa | Trolley | SYAS/SC11 | 230769101000441683 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
19000786 | 2024-09-18 | RSUD Budi Rahayu Magelang | Ventilator | Draeger/EvitaV600 | ASTD0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
9000483 | 2024-09-18 | RSP Fam Dofa | Gynaecologi Table | SYAS/SGT2 | 230734191000900083 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000480 | 2024-09-18 | RSP Fam Dofa | Overbed table | SYAS/STD3-GS | 230761103000401031 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000479 | 2024-09-18 | RSP Fam Dofa | Cabinet | SYAS/Cabinet 99 | 230633132000400906 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000478 | 2024-09-18 | RSP Fam Dofa | Dirty Linen Trolley | SYAS | 2306696001000400115 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000477 | 2024-09-18 | RSP Fam Dofa | Food Trolley | STAS/SEF1 | 230835213000400045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000476 | 2024-09-18 | RSP Fam Dofa | Stretcher Standard | SYAS/SST1 | 23066830200400138 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000475 | 2024-09-18 | RSP Fam Dofa | Examination Table | SYAS/Standar SET1 | 230764101000400633 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000474 | 2024-09-18 | RSP Fam Dofa | Meja Operasi | SYAS/Electrohydraoulic | 230867221180400037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000473 | 2024-09-18 | RSP Fam Dofa | Infused Standard | SYAS/4Hook | 230836163000403562 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000472 | 2024-09-18 | RSP Fam Dofa | Wheel Chair | SYAS/SW01-SS | 23081313000401732 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000471 | 2024-09-18 | RSP Fam Dofa | Emergency Trolley | STAS/SCE2 | 230869801000400455 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000469 | 2024-09-18 | RSP Fam Dofa | Bed Patient | SYAS/ 3 crank | 230962404000400302 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000470 | 2024-09-18 | RSP Fam Dofa | Baby Bed | SYAS/JBB2 | 2306262110004230 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
9000468 | 2024-09-18 | RSP Fam Dofa | Medicine Trolley | SYAS/SCM20 | 2.3077e+17 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
11000343 | 2024-09-17 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Survey | Pengecekan EMH dan linear actuator | Pending | Linear actuator rusak | Nasrul
19000785 | 2024-09-17 | RSUD Dr. Soediran Mangun Sumarso | Patient Monitor | Draeger/Vista 120 CMS | VGSCF0029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000785 | 2024-09-17 | RSUD Dr. Soediran Mangun Sumarso | Patient Monitor | Draeger/Vista 120 CMS | VGSCF0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000785 | 2024-09-17 | RSUD Dr. Soediran Mangun Sumarso | Patient Monitor | Draeger/Vista 120 CMS | VGSCF0026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000785 | 2024-09-17 | RSUD Dr. Soediran Mangun Sumarso | Patient Monitor | Draeger/Vista 120 CMS | VGSCF0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
8000357 | 2024-09-17 | RSUD Batang | Mesin Anasthesia | Draeger/Atlan A300 | ASTH0152 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000357 | 2024-09-17 | RSUD Batang | Mesin Anasthesia | Draeger/Atlan A300 | ASTH0153 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
7000596 | 2024-09-17 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | VGSSF0082 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000596 | 2024-09-17 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | VGSSK0089 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
4000449 | 2024-09-17 | Signature Dental Care | Dental Unit | Planmeca/Compact i5 | UPCV574532 | Maintenance | Preventice Maintenance | Ready to use |  | Satria, Ipan
7000596 | 2024-09-17 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | VGSSF0090 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000596 | 2024-09-17 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | VGSSF0105 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP850XWYX00135 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP850XWYX00138 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP850XWYX00145 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP850XWYX00002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP850XWYX00139 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP850XWYX00141 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP8500XWYX0006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP8500XWYX00143 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP8500XWYX00142 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP8500XWYX00144 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP8500XWYX0005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
2000290 | 2024-09-13 | RSUP Makassar | Meja Operasi | SYAS | OP8500XWYX0004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
10000203 | 2024-09-13 | RSUD H. Boejasin | MoT | - | - | Service | Monitor tidak menyala dan tidak auto ON; Setting IP PLC, bios auto ON | Ready to use |  | Heru
8000356 | 2024-09-13 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 edition | ASEL0142 | Survey | O2 setting alarm low | Pending | Perlu penggantian O2 sensir capsule, Neo flow insert, inspiration block, expiration valve dan cable flow sensor neo | Amna
8000356 | 2024-09-13 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 edition | ASEL0143 | Survey | O2 setting alarm low | Pending |  | Amna
16000249 | 2024-09-12 | Dosy dental care | Dental Unit | Kavo/Primus 1058 | 1003976 | Survey | Tekanan air di scaller bermasalah; Handpiece di scaller tidak bergetar namun air normal | Pending | Perlu dilakukan penggantian handpiece | Ipan
20000296 | 2024-09-12 | RSUP Tadjudin Chalid | ESWL | Medispec/EM1000 | 0922014 | Service | UPS mati total; Pengecekan connector UPS, perbaikan colokan kabel | Ready to use |  | Sujana
20000310 | 2024-09-12 | RSUP Tadjudin Chalid | Digital Radiografi Fluoroskopi | Apelem/Platinum | J236F0882 | Service | Connecting PACS | Pending | Unit belum bisa tersambung karena IT belum bisa | Sujana
19000784 | 2024-09-12 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Atlan | ASPH0141 | Maintenance | Maintenance unit dan refresh training | Ready to use |  | Nanda
17000193 | 2024-09-12 | RSUP Dr. Wahidin Sudirohusodo | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0495 | Service | Hasil gambar kurang bagus; Pengambilan data untuk Apelem | Pending |  | Bahar, Satria
17000192 | 2024-09-12 | RSUP Dr. Wahidin Sudirohusodo | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0495 | Service | Ambil data kelengkapan revisi UKES | Ready to use |  | Bahar, Satria
20000282 | 2024-09-11 | RSKD Gigi dan Mulut Makassar | Dental Unit | Kavo/1058TM | 10207188 | Service | Error PCBA; Pemasangan PCBA unit | Ready to use | Unit berfungsi namun triway syringe bocor dan disarankan ganti | Sujana
9000464 | 2024-09-11 | RSUD Sanana | UPS | Yoshiga/GPI 33-80 | 501212003370N380001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
19000783 | 2024-09-10 | RSIA Umi Barokah Boyolali | Baby Incubator | Draeger/Isolette C2000 | DSSK0008 | Maintenance | Maintenance unit dan refresh training | Ready to use |  | Nanda
19000783 | 2024-09-10 | RSIA Umi Barokah Boyolali | Baby Incubator | Draeger/Isolette C2000 | DSRL0029 | Maintenance | Maintenance unit dan refresh training | Ready to use |  | Nanda
5000199 | 2024-09-10 | RSUD Dr. Drajat Prawiranegara | Meja Operasi | SYAS/Electrohydraoulic | 2211672211840048 | Service | Meja tidak berfungsi total; Penggantian remote | Ready to use |  | Bilal, Ipan
5000199 | 2024-09-10 | RSUD Dr. Drajat Prawiranegara | Meja Operasi | SYAS/Electrohydraoulic | 2211672211840044 | Service | Meja tidak berfungsi total; Penggantian board | Ready to use |  | Bilal, Ipan
5000199 | 2024-09-10 | RSUD Dr. Drajat Prawiranegara | Meja Operasi | SYAS/Electrohydraoulic | 2211672218400047 | Service | Meja tidak berfungsi total; Penggantian board | Ready to use |  | Bilal, Ipan
18000525 | 2024-09-10 | RSUD Tobelo | Neurosurgery | Aesculap | 2563 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
8000355 | 2024-09-10 | RS Islam Pati | UPS | Yoshiga/HPT1 | 501101213200KB200041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
11000325 | 2024-09-10 | RSUP Tadjudin Chalid | ESWL | Medispec/EM1000 | 0922014 | Service | Monitor c-arm mati, cek koneksi kabel; Penggantian baterai CMOS | Ready to use |  | Nasrul
3000328 | 2024-09-10 | RSUP Dr. H. Adam Malik Medan | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/083/C1 | Survey | Low kv output level; Cek tegangan input, cek tegangan trafo. Generator, inverter dan HTT rusak | Pending |  | Imam
9000465 | 2024-09-09 | RSUD Sanana | UPS | Yoshiga/GPI-33 | 5012122003370N380001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
19000782 | 2024-09-09 | RSUD Tugurejo | Baby Incubator | Draeger/Caleo | ASHK0030 | Service | Water humidity low, selenoid valve rusak | Ready to use |  | Nanda
11000355 | 2024-09-06 | RS Universitas Hasanuddin | Lampu Operasi | Draeger/Polaris 600 | ASRF1087 | Service | Kamera mati, kabel power rusak |  |  | Nasrul
18000524 | 2024-09-06 | RSUD Dr. H. Chasan Boesoirie | Mesin Anasthesia | Draeger/Perseus A500 | ASSK0059 | Survey | Training user | Ready to use |  | Irfan R
19000781 | 2024-09-06 | RSUD Tugurejo | Baby Incubator | Draeger/Caleo | ASHK0030 | Survey | Humidity tidak bisa naik | Pending | Unit akan dicek lebih lanjut | Nanda
8000600 | 2024-09-06 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | Service | O2 sensor failed; Penggantian O2 sensor | Ready to use |  | Amna
8000600 | 2024-09-06 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Service | O2 sensor failed; Penggantian O2 sensor | Ready to use |  | Amna
3000327 | 2024-09-06 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Survey | UPS unit bunyi alarm; Cek input, output, output dengan bebas UPS | Ready to use | UPS berfungsi normal, disarankan memperbaiki input UPS | Imam
18000526 | 2024-09-05 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 120 | VGSNE2298 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000526 | 2024-09-05 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 120 | VGSRF0245 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000526 | 2024-09-05 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista 120 | VGSRF0247 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
7000593 | 2024-09-05 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Survey | Water filler cleaning | Ready to use |  | Wahyu
16000248 | 2024-09-05 | GiGi Dental House | Dental Unit | Kavo/Primus 1058 | 10207196 | Survey | Suction lemah; Pengecekan selang dan filter. Unit normal | Ready to use |  | Ipan, Satria
11000341 | 2024-09-05 | RSUP Makassar | Lampu Operasi | SYASTECH/Lumina | OL9680/80XVZW00037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul, Sujana
18000524 | 2024-09-04 | RSUD Dr. H. Chasan Boesoirie | Mesin Anasthesia | Draeger/Perseus A500 | ASSK0059 | Instalasi | Uji Fungsi | Ready to use |  | Irfan R
19000779 | 2024-09-04 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN800 | ASRH0028 | Survey | Nebu fail; Proportional valve dan pressure regulator rusak | Pending |  | Nanda
4000448 | 2024-09-04 | RSUD Dr. Drajat Prawiranegara | Meja Operasi | Syas/Hydraulic | 2211672211840048 | Service | Remote tidak berfungsi; Penggantian remote | Pending | Unit masih tidak berfungsi, akan dijadwalkan untuk visit teknisi MAK | Bahar, Satria
4000448 | 2024-09-04 | RSUD Dr. Drajat Prawiranegara | Meja Operasi | Syas/Hydraulic | 2211672211840044 | Service | Remote tidak berfungsi; Penggantian remote | Pending | Unit masih tidak berfungsi, akan dijadwalkan untuk visit teknisi MAK | Bahar, Satria
4000448 | 2024-09-04 | RSUD Dr. Drajat Prawiranegara | Meja Operasi | Syas/Hydraulic | 2211672218400047 | Service | Remote tidak berfungsi; Penggantian remote | Pending | Unit masih tidak berfungsi, akan dijadwalkan untuk visit teknisi MAK | Bahar, Satria
8000599 | 2024-09-04 | RSUD Dr. Soeselo Slawi | Lampu Operasi | SYAS/Lumina | OL9680/80240100055 | Service | Setting arm | Ready to use |  | Amna
7000592 | 2024-09-04 | RSUD dr. Soeradji Tirtonegoro | Central Monitor | Draeger/Vista 120 CMS | VMSPC0056 | Service | Central monitor disconnect; Atur ulang ip address dan install ulang cms | Ready to use |  | Wahyu
10000186 | 2024-09-03 | RSI Sultan Agung Banjarbaru | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0126 | Service | Volume tidar (VT) tidak sesuai; Kalibrasi internal dan penggantian water lock | Pending | Unit harus dilakukan penggantian sample line | Heru
19000774 | 2024-09-03 | RS St. Elizabeth | Lampu Operasi | Draeger/Polaris 600 | ASRK1018 | Maintenance | Adjustment arm lampu operasi | Ready to use |  | Nanda
6000744 | 2024-09-03 | RSUP Surabaya | Lampu Operasi | SYAS/Lumina | OL9680/80XVZW00044 | Survey | Spring arm patah; Perbaikan hardware spring arm | Ready to use | Alat berfungsi dengan baik | Afrizal
19000810 | 2024-09-02 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN800 | ASTF0051 | Survey | Layar mati dan tidak bisa di pencet; Kalibrasi touchscreen | Ready to use |  | Nanda
10000192 | 2024-09-02 | RSUD Ulin | Laser Surgical | Jena Surgical/Multipulse HoPlus | J191150176 | Maintenance | Preventice Maintenance | Ready to use | Unit ready namun fiber optic 272 tidak ready | Heru
8000598 | 2024-09-02 | RSI Sultan Hadlirin | Ventilator | Draeger/Babylog VN600 | ASTD0005 | Service | Humidity over temperature; Setting HC | Ready to use |  | Amna
7000595 | 2024-08-30 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Maintenance | PM Total Care; Penggantian service set 2 tahun | Ready to use |  | Wahyu
7000594 | 2024-08-30 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Maintenance | PM Total Care; Penggantian service set 2 tahun | Ready to use |  | Wahyu
6000742 | 2024-08-30 | RS Raffa Medika | Ventilator | Draeger/Babylog VN800 | ASPC0081 | Survey | Installasi backup unit | Ready to use |  | Afrizal
17000191 | 2024-08-30 | RSUD Sele Be Solu | Mobile X-ray | Apelem/Rafale DR | 81124111332 | Survey | Uji kesesuaian dan Uji paparan | Ready to use |  | Bahar
18000547 | 2024-08-30 | RSUD Sele Be Solu | UPS | Yoshiga/GPI33 | E2404CO12260721392 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Irfan R
4000447 | 2024-08-29 | Signature Dental Care | Mesin Milling | XTCERA/D5 | - | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria, Ipan
20000287 | 2024-08-29 | RSUP Makassar | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0145 | Instalasi |  | Ready to use |  | Eddy, Imam, Bilal, Nasrul, Sujana
6000741 | 2024-08-29 | RSUD dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASNB0071 | Survey | Sensor check kuning, valve check kuning; Perlu penggantian flow sensor | Ready to use | Penggantian flow sensor milik RS | Afrizal
12000113 | 2024-08-29 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0081 | Survey | Flow power dan ILCA3 bermasalah | Pending |  | Marshall
18000548 | 2024-08-29 | RSUD Sele Be Solu | UPS | Yoshiga/GPI33 | E2404CO12260721391 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
19000773 | 2024-08-29 | RSUD Dr. Soehadi Prijonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASKE0184 | Maintenance | Unit checking | Ready to use |  | Nanda
19000771 | 2024-08-29 | RSUD Dr. Soehadi Prijonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASKE0187 | Maintenance | Unit checking | Ready to use | Unit masih bisa digunakan tanpa monitoring O2 | Nanda
19000773 | 2024-08-29 | RSUD Dr. Soehadi Prijonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASKE0184 | Maintenance | Unit checking | Ready to use |  | Nanda
19000770 | 2024-08-29 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Service | Shock lompat dari counter; Kalibrasi HV quick | Ready to use |  | Nanda
19000772 | 2024-08-29 | RSUD Dr. Soehadi Prijonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0047 | Maintenance | Unit checking | Ready to use |  | Nanda
10000190 | 2024-08-29 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Survey | Kunjungan unit | Ready to use |  | Heru
6000740 | 2024-08-28 | RS Bhakti Asih Brebes | Ventilator | Draeger/Savina300 | ASHD0011 | Survey | Alat mati total; Perlu penggantian part mainboard dengan yang baru | Pending |  | Afrizal
6000739 | 2024-08-28 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHM0014 | Survey | Baterai fail; Penggantian baterai interal dan external | Pending |  | Afrizal
6000739 | 2024-08-28 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHL0170 | Survey | Baterai fail; Penggantian baterai interal dan external | Pending |  | Afrizal
6000739 | 2024-08-28 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHL0172 | Survey | Baterai fail; Penggantian baterai interal dan external | Pending |  | Afrizal
6000739 | 2024-08-28 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHL0173 | Survey | Baterai fail; Penggantian baterai interal dan external | Pending |  | Afrizal
12000114 | 2024-08-28 | RSUD Toto Kabila | C-arm | Apelem/Evo R+ | 55121004853 | Service | Troubleshoot dengan part pinjaman dari Kandou | Pending |  | Marshall
4000466 | 2024-08-27 | Dosy dental care | Dental Unit | Kavo/1058TM | 20051003975 | Survey | Coupling micromotor tidak berputar, tip scaler patah dan air tidak keluar | Pending | Perlu penggantian coupling micromotor dan scaller handpiece | Satria, Ipan
9000460 | 2024-08-27 | RSUD Otohana | Ventilator | Draeger/Evita V800 | ASPL0095 | Service | Pressure regulator bocor; Penggantian pressure regulator | Ready to use |  | Ruslan
6000738 | 2024-08-27 | RS Bhakti Asih Brebes | Mesin Anasthesia | Draeger/Primus | ASHM0013 | Survey | Sensor check kuning, valve check kuning; Perlu penggantian MV2 dan baterai | Pending |  | Afrizal
8000595 | 2024-08-27 | RS PKU Aisyah Jepar | Mesin Anasthesia | Draeger/Fabius Plus | ASND0058 | Service | Gas supply not detected; Cek pressure regulator, valve vontrol N20 | Ready to use |  | Amna
110000323 | 2024-08-26 | RSUP Tadjudin Chalid | ESWL | Medispec/EM1000 | 0922014 | Service | Error #301 dan error #06 thermal sensor; Cek mcu board, sensor thermal, terminal pcb. Terminal pcb diperbaiki | Ready to use |  | Nasrul
6000737 | 2024-08-26 | RSUD Kardinah Tegal | Printer | Drypix | - | Survey | cek error; setting software | Ready to use |  | Afrizal
8000594 | 2024-08-23 | RS PKU Muhammadiyah Mayong | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ0081 | Survey | O2 sensor fail; O2 sensor capsule check | Ready to use |  | Amna
8000593 | 2024-08-22 | RSU Sebening Kasih | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ0082 | Maintenance | Preventice Maintenance | Ready to use |  | Amna
4000445 | 2024-08-22 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Service | Meja tidak bergerak, komputer console tidak menyala otomatis; Pengecekan kabel LAN, penggantian baterai cmos | Ready to use |  | Bahar & Satria
2000285 | 2024-08-21 | RSUP Dr. Wahidin Sudirohusodo | Meja Operasi | SYAS | - | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
8000597 | 2024-08-21 | RSI Sultan Hadlirin | Ventilator | Draeger/Babylog VN600 | ASTD0005 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Amna
19000769 | 2024-08-21 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Survey | Error kita survey; Set ups, kalibrasi internal | Ready to use |  | Nanda
16000247 | 2024-08-21 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Maintenance |  | Ready to use |  | Ipan
20000281 | 2024-08-20 | RSUD Madani Palu | Ventilator | Draeger/Savina300 | ASSL0013 | Instalasi | Pemasangan aksesoris humidifier | Ready to use |  | Sujana
20000281 | 2024-08-20 | RSUD Madani Palu | Ventilator | Draeger/Babylog VN800 | ASTA0015 | Instalasi | Pemasangan aksesoris humidifier | Ready to use |  | Sujana
4000444 | 2024-08-20 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Service | Uji Kesesuaian | Ready to use |  | Bahar & Satria
19000768 | 2024-08-20 | RS JIH Surakarta | Vaporizer | Draeger/Vapor 2000 | ASLK0616 | Service | Adjusment vapo | Ready to use |  | Nanda
11000324 | 2024-08-20 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVZW00037 | Instalasi | Aksesoris lengkap, Alat berfungsi dengan baik | Ready to use | Belum uji fungsi | Nasrul, Sujana
11000324 | 2024-08-20 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVZU00062 | Instalasi | Aksesoris lengkap, Alat berfungsi dengan baik | Ready to use | Belum uji fungsi | Nasrul, Sujana
11000324 | 2024-08-20 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVZU00061 | Instalasi | Aksesoris lengkap, Alat berfungsi dengan baik | Ready to use | Belum uji fungsi | Nasrul, Sujana
11000324 | 2024-08-20 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVZW00036 | Instalasi | Aksesoris lengkap, Alat berfungsi dengan baik | Ready to use | Belum uji fungsi | Nasrul, Sujana
11000324 | 2024-08-20 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVZU00068 | Instalasi | Aksesoris lengkap, Alat berfungsi dengan baik | Ready to use | Belum uji fungsi | Nasrul, Sujana
11000324 | 2024-08-20 | RSUP Makassar | Lampu Operasi | SYAS/Lumina | OL9680/80XVZW00046 | Instalasi | Aksesoris lengkap, Alat berfungsi dengan baik | Ready to use | Belum uji fungsi | Nasrul, Sujana
8000592 | 2024-08-20 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHL0168 | Survey | O2 supply, air supply dan O2 system check fail; Cek PATA sensor, kit proporsional valve | Pending | Perlu penggantian PATO Sensor dan proportional valve | Amna
8000591 | 2024-08-19 | RS Mitra Bangsa Pati | Ventilator | Draeger/Babylog VN600 | ASTB0024 | Service | Leakage dan compliance test fail; Cek neo flow dan breathing circuit | Ready to use |  | Amna
16000246 | 2024-08-19 | RSUD Kab. Bekasi | Bed Patient | SYAS | 240526302000400044 | Service | Handle patah; Penggantian handle | Ready to use |  | Ipan
19000767 | 2024-08-19 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0149 | Survey | Layar ventilator tiba tiba mati dan restart, cek unit dan tarik error. Di cek running sampai besok | Ready to use |  | Nanda
7000590 | 2024-08-19 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Survey | Layar bekedip; cek koneksi kabel display kit-kencangkan sambungan kabel | Ready to use |  | Wahyu
7000590 | 2024-08-19 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Survey | PCM error; kalibrasi internal | Ready to use |  | Wahyu
4000442 | 2024-08-18 | RSUP Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0146 | Instalasi | Alat sudah diinstall namun belum ada supply listrik | Pending |  | Satria, Eddy, Purwono
8000584 | 2024-08-16 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASSC0025 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000584 | 2024-08-16 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0049 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000584 | 2024-08-16 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0015 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000584 | 2024-08-16 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0148 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
7000589 | 2024-08-15 | RSUD Sleman | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0128 | Survey | Central O2 supply fault; Cek error log, pressure sensor disconnect | Pending | Indikasi rusak pcb optimix pressure sensor | Wahyu
18000543 | 2024-08-15 | RSUD Tambulilato | UPS | Yoshiga/GPI40 | E2404C012260631393 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000542 | 2024-08-15 | RSUD Tambulilato | UPS | Yoshiga/GPI40 | E2404C012260631394 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Irfan R
11000317 | 2024-08-15 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASMF0003 | Survey | Alarm system failure, knop rotasi tidak berfungsi | Pending | Disarankan penggantian loudspeaker set dan board rotary knop | Nasrul
19000765 | 2024-08-14 | RS JIH Surakarta | Vaporizer | Draeger/Vapor 2000 | ASLK0616 | Survey | Kalibrasi fail; Adjust dan setting dan kalibrasi | Ready to use |  | Nanda
19000766 | 2024-08-14 | RS JIH Surakarta | Ventilator | Draeger/Babylog VN600 | ASSK0027 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
8000590 | 2024-08-14 | RSUD Dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0109 | Survey | O2 sensor capsule expired | Pending | Perlu penggantian O2 sensor capsule | Amna
3000322 | 2024-08-14 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Software duet sering crash; Reload software, setting worklist | Ready to use |  | Imam
7000584 | 2024-08-13 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASHL0055 | Survey | Device failure; Cek error | Pending | Perlu penggantian Mobi, CPU VGL, CPU Mixer, Restricfer, baterai internal | Wahyu
7000585 | 2024-08-13 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASKM0402 | Survey | Device failure; Tarik error log, cek pgm & cek baterai | Pending | Perlu penggantian mPGM set dan internal baterai | Wahyu
8000589 | 2024-08-13 | RSUD RAA Soewondo Pati | Infant Warmer | Draeger/Resuscitaire | KC13741 | Survey | Error 05 & valve SW error, cek heater dan valve | Pending | Perlu penggantian heater assy dan valve 5 | Amna
19000764 | 2024-08-13 | RSUD Dr. Soediran Mangun Sumarso | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/134/C2 | Maintenance | Preventice Maintenance | Ready to use |  | Nanda
19000763 | 2024-08-13 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 319201 | Survey | Cek unit, ada kabel meja pasien yang putus; Koneksi ulang | Ready to use |  | Nanda
19000762 | 2024-08-12 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 09221164 | Service | Probe USG kendor bagian mounting; ganti karet pada mounting | Ready to use | r | Nanda
11000320 | 2024-08-12 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASNL0248 | Service | Baterai internal fail; Pemasangan maintenance kit | Ready to use |  | Nasrul
11000320 | 2024-08-12 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASNL0249 | Service | Baterai internal fail; Pemasangan maintenance kit | Ready to use |  | Nasrul
11000320 | 2024-08-12 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASND0141 | Service | Baterai internal fail; Pemasangan maintenance kit | Ready to use |  | Nasrul
11000321 | 2024-08-12 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina300 | ASNP0009 | Survey | Baterai internal tidak ada; Pemasangan service set 2 tahun | Ready to use |  | Nasrul
4000439 | 2024-08-11 | RSUP Surabaya | Meja Operasi | SYAS | OP85OXWYX00136 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000439 | 2024-08-11 | RSUP Surabaya | Lampu Operasi | SYAS | OL9680/80XVZW00047 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000439 | 2024-08-11 | RSUP Surabaya | Pendant Surgery | SYAS | SYPS2406011 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000439 | 2024-08-11 | RSUP Surabaya | Pendant Anasthesi | SYAS | SYPA2402010 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000440 | 2024-08-11 | RSUP Surabaya | Meja Operasi | SYAS | OP85OXWYX00134 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000440 | 2024-08-11 | RSUP Surabaya | Lampu Operasi | SYAS | OL9680/80XVZW00049 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000440 | 2024-08-11 | RSUP Surabaya | Pendant Surgery | SYAS | SYPS2406009 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000440 | 2024-08-11 | RSUP Surabaya | Pendant Anasthesi | SYAS | SYPA2402009 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000441 | 2024-08-11 | RSUP Surabaya | Meja Operasi | SYAS | OP85OXWYX00003 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000441 | 2024-08-11 | RSUP Surabaya | Lampu Operasi | SYAS | OL9680/80XVZW00044 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000441 | 2024-08-11 | RSUP Surabaya | Pendant Surgery | SYAS | SYPS2406010 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000441 | 2024-08-11 | RSUP Surabaya | Pendant Anasthesi | SYAS | SYPA2402008 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000438 | 2024-08-11 | RSUP Surabaya | Meja Operasi | SYAS | OP85OXWYX00137 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000438 | 2024-08-11 | RSUP Surabaya | Lampu Operasi | SYAS | OL9680/80XVZW00032 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000438 | 2024-08-11 | RSUP Surabaya | Pendant Surgery | SYAS | SYPS2406012 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
4000438 | 2024-08-11 | RSUP Surabaya | Pendant Anasthesi | SYAS | SYPA2402007 | Instalasi | Alat berfungsi dengan baik | Ready to use | Unit belum dilakukan uji fungsi | Satria, Purwono
10000202 | 2024-08-10 | RSUD H. Boejasin | Meja Operasi | SYAS/EPIONE | OP850XVZY00001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Heru
9000457 | 2024-08-09 | RSUD Sanana | X-ray Digital | Apelem/Camargue DR | 251023012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan, Bahar
8000588 | 2024-08-09 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0014 | Survey | Power fail alarm; Check power pack, power pack fail | Pending | Perlu penggantian power pack | Amna
10000201 | 2024-08-08 | RSUD H. Boejasin | Lampu Operasi | SYAS/Lumina | OL9680/80XVZW00033 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Heru
8000587 | 2024-08-08 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Service | Flow tube setting dan lpm maksimal; Setting O2 flow tube | Ready to use | Disarankan penggantian O2 sensor | Amna
8000587 | 2024-08-08 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0014 | Service | Flow tube setting dan lpm maksimal; Setting O2 flow tube | Ready to use | Disarankan penggantian O2 sensor | Amna
7000582 | 2024-08-08 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASJN0023 | Maintenance | Preventive Maintenance Total Care | Ready to use |  | Wahyu
17000188 | 2024-08-08 | RSUD Sanana | X-ray Digital | Apelem/Camargue DR | 251023012 | Survey | Uji kesesuaian dan uji paparan | Ready to use |  | Bahar
7000583 | 2024-08-08 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASNL0107 | Maintenance | Preventive Maintenance Total Care | Ready to use |  | Wahyu
15000257 | 2024-08-08 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Survey | Water sensor tidak terbaca; Sensor water dibawa ke Jakarta untuk dicek lebih lanjut | Pending |  | Sutan
3000320 | 2024-08-08 | RSUD Dr. Drajat Prawiranegara (KSO) | Cathlab | Siemens/Artis ONE | 82377 | Service | Unit tidak menyala; Cek supply tegangan ke unit dan circuit breaker posisi tripped | Ready to use |  | Imam
4000443 | 2024-08-07 | RSUP Surabaya | Digital Radiografi Fluoroskopi | Apelem/Platinum Neo | G248L0146 | Survey | Survey ruangan untuk pra install | Ready to use |  | Satria, Purwono
5000203 | 2024-08-07 | RSUD Benda | Operating table | SYAS/SE02 | 230952512100400053 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal, Ipan
12000126 | 2024-08-07 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | 2510220561 | Service | Setting ukuran dan suction tray printer fujifilm | Ready to use |  | Marshall
19000757 | 2024-08-07 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Primus | ASJE0011 | Service | 2/2 distributing valve error | Ready to use |  | Nanda
5000203 | 2024-08-07 | RSUD Benda | Operating table | SYAS/SE02 | 230952512100400054 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal, Ipan
18000545 | 2024-08-07 | RSUD Ainun Habibie | Meja Operasi |  | 202406001217 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000544 | 2024-08-07 | RSUD Ainun Habibie | UPS | Yoshiga/GPI33 | E2405C012260751478 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000546 | 2024-08-07 | RSUD Ainun Habibie | Meja Operasi | SYAS/Shadowless | HF20240604 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
8000596 | 2024-08-07 | RSI Sultan Hadlirin | Ventilator | Draeger/Babylog VN600 | ASTD0005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
3000319 | 2024-08-07 | RSUD Malingping | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/131/C2 | Survey | Memory display rusak | Pending |  | Imam
8000586 | 2024-08-06 | RSUD RA Kartini | Ventilator | Draeger/Savina300 | ASEK0197 | Service | Error 42000; Penggantian engine drive, blow engine unit | Ready to use |  | Amna
6000731 | 2024-08-06 | RS Raffa Medika | Ventilator | Draeger/Babylog VN800 | x | Instalasi | Demo unit | Ready to use |  | Afrizal
6000731 | 2024-08-06 | RS Raffa Medika | Ventilator | Draeger/Evita V300 | ASNL0235 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
15000256 | 2024-08-06 | RSUP Dr. Mohammad Hoesin (KSO) | CT Scan | Siemens/M-CT172 | 659212171 | Service | Fan filter detector mati; Penggantian fan filter detektor | Ready to use |  | Sutan
8000585 | 2024-08-05 | RSUD Kayen Pati | Infant Warmer | Draeger/Resuscitaire | GV14025 | Survey | O2 pneumonatic leaking; Pneumatic fail | Pending | Perlu dilakukan penggantian O2 pneumonatic | Amna
12000125 | 2024-08-05 | RSUD Sanana | X-ray Digital | Apelem/Camargue DR | 251023012 | PraInstall |  |  |  | Marshall
19000752 | 2024-08-05 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0091 | Service | Penggantian blower | Ready to use |  | Nanda
9000456 | 2024-08-05 | RSUD Sanana | Dental Unit | Kavo/E30TM | 20190910111150 | Service | Air ke handpiece sering macet; Penggantian control valve unit dan GSB Spray | Ready to use | Direkomendasikan penggantian selang scaller dan scallernya | Ruslan
19000780 | 2024-08-05 | RSUP dr. Kariadi | Infant Warmer | Draeger/Resuscitaire | KC13734 | Service | Error 05; Penggantian heater | Ready to use |  | Nanda
6000728 | 2024-08-04 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0043 | Survey | Ventilation check | Ready to use |  | Afrizal
6000730 | 2024-08-04 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0042 | Survey | Ventilation check | Ready to use |  | Afrizal
6000729 | 2024-08-04 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASLD0042 | Survey | Ventilation check | Ready to use |  | Afrizal
3000321 | 2024-08-02 | RSUD Pagelaran (KSO) | CT Scan | Philips/Ingenuity | 320121 | Service | PC CIRS mati; Perbaikan pc cirs, namun masih ada error resource allocation for scan failed | Pending |  | Imam
6000736 | 2024-08-01 | RSUD dr. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEL0081 | Survey | Gas supply calibration error, O2 calibration error; Cek software dan hardware | Pending | Perlu penggantian proportional valve, one way valve, baterai internal | Afrizal
19000749 | 2024-08-01 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Survey | Error saat QC; Setting adjustment | Ready to use |  | Nanda
15000254 | 2024-08-01 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Install ulang aplikasi | Ready to use |  | Sutan
10000200 | 2024-08-01 | RSUD H. Boejasin | Pendant Anasthesi | SYAS/Pegasus | SYPS2402004 | Instalasi | Alat berfungsi dengan baik dan siap digunakan | Ready to use |  | Heru
10000199 | 2024-08-01 | RSUD H. Boejasin | Pendant Surgery | SYAS/Pegasus | SYPS2402005 | Instalasi | Alat berfungsi dengan baik dan siap digunakan | Ready to use |  | Heru
20000262 | 2024-07-31 | RSKD Gigi dan Mulut Makassar | Dental Unit | Kavo/Primus 1058 | 10207188 | Survey | Unit mati total; Pengecekan dentis element | Pending | Perlu penggantian dentist element | Sujana
6000716 | 2024-07-31 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHL0151 | Service | Pemasangan distributing valve, service set breathing system 2thn, prmus service set 3 tahun | Ready to use |  | Afrizal
8000583 | 2024-07-31 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius | ARYS0111 | Service | Penggantian O2 flus assembly | Ready to use |  | Amna
19000748 | 2024-07-31 | RSUD Pandan Arang | Incubator Transport | Draeger/Isolette TI500 | CF32769 | Service | Penggantian pengunci pintu | Ready to use |  | Nanda
19000747 | 2024-07-31 | RSUD Pandan Arang | Incubator Transport | Draeger/Isolette TI500 | BH12345 | Service | Penggantian baterai | Ready to use |  | Nanda
7000581 | 2024-07-31 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus | ASKF0077 | Service | Error gas monitoring; Penggantian flow power | Ready to use |  | Wahyu
15000253 | 2024-07-30 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Penggantian lampu kolimator | Ready to use |  | Sutan
15000252 | 2024-07-30 | RSUP Dr. Mohammad Hoesin (KSO) | CT Scan | Siemens | 659212171 | Service | Over head hantry; Pembersihan filter generator | Ready to use |  | Sutan
13000173 | 2024-07-30 | RSU Haji Medan | ESWL | Medispec/E3000 | 0719138 | Survey | Training dan pendampingan pasien | Ready to use |  | Sofian
19000746 | 2024-07-30 | RSUD dr. Adhyatma Semarang | Mesin Anasthesia | Draeger/Primus | ASHJ0179 | Service | Sensor check error; Penggantian absokute dan cylinder pressure sensor | Ready to use |  | Nanda
8000582 | 2024-07-30 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASMK0052 | Survey | Alarm cookpit; Hit software | Ready to use |  | Amna
7000580 | 2024-07-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0007 | Survey | Vaporizer salah diisi liquid | Pending | Unit perlu dibersihkan | Wahyu
7000580 | 2024-07-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0023 | Survey | Vaporizer salah diisi liquid | Pending | Unit perlu dibersihkan | Wahyu
19000745 | 2024-07-29 | RS Ken Saras | Mesin Anasthesia | Draeger/Fabius GS Premium | ASCL0014 | Service | Baterai fail; Baterai internal rusak, pressure regulator fail | Pending | Disarankan penggantian baterai internal dan pressure regulator | Nanda
16000245 | 2024-07-29 | RSUD Kab. Bekasi | Bed Patient | SYASTECH | 248526302000400044 | Survey | Tiang rail pegangannya patah dan karet tiang rusak | Pending | Perlu penggantian rail dan karet ganjalan | Ipan
4000434 | 2024-07-29 | RSUD Selaparang | Mobile X-ray | Apelem/Rafale | 1876552200001 | Survey | Uji paparan | Ready to use |  | Satria
19000754 | 2024-07-29 | RSUD Budi Rahayu Magelang | Ventilator | Draeger/Babylog VN500 | x | Instalasi | Installasi unit demo | Ready to use |  | Nanda
8000581 | 2024-07-29 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0110 | Survey | Leak dan compliance test fail; Setting piston ventilator | Ready to use |  | Amna
7000579 | 2024-07-29 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina300 | ASSN0009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000579 | 2024-07-29 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina300 | ASSN0010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
12000115 | 2024-07-27 | RSUD Toto Kabila | UPS | Yoshiga/GPH33 | E2405C012260751474 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Marshall
11000322 | 2024-07-26 | RSUD Tenriawaru Bone | X-ray Digital | Apelem/Camargue DR | CPD26084L16 | Survey | Disk full, tidak bisa print. gambar bergaris; Cleaning memory, kalibrasi detektor | Ready to use | Alat berfungsi namun gambar bergaris. Disarankan penggantian detektor | Nasrul
19000743 | 2024-07-26 | RSUP dr. Kariadi | Dental Unit | Kavo/1058 | 1002719 | Service | Air kumur bocor; Pembersihan selenoid valve | Ready to use |  | Nanda
19000744 | 2024-07-26 | RSUP dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0300 | Service | Pressure measurement tidak akurat; Internal kalibrasi | Ready to use |  | Nanda
15000251 | 2024-07-26 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Lampu kolimator mati | Pending | Disarankan penggantian lampu kolimator | Sutan
3000318 | 2024-07-26 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226P6745 | Survey | Hasil gambar terdapat artefak | Pending | Menunggu hasil analisa principle | Imam
17000187 | 2024-07-25 | RSUD Tidar Magelang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0496 | Survey | Uji Kesesuaian | Ready to use |  | Bahar
19000742 | 2024-07-25 | RSUP dr. Kariadi | Infant Warmer | Draeger/Babytherm 8004 | ARUC0014 | Service | Error 03; Board WT sensor rusak | Ready to use | Disarankan penggantian board WT sensor | Nanda
13000172 | 2024-07-25 | RSU Haji Medan | Dental Unit | Kavo/1058TM | 20008319 | Service | Isi ulang cairan oxygenal dan penggantian filter filamen | Ready to use |  | Sofian
8000580 | 2024-07-25 | RSI Sunan Kudus | Ventilator | Draeger/Babylog VN800 | ASPF0019 | Service | neonatal flow sensor fail; Penggantian neonatal flow sensor | Ready to use |  | Amna
7000591 | 2024-07-24 | RSUP dr. Soeradji Tirtonegoro | Bubble CPAP | Draeger/Seattle PAP | QM20230155 | Service | O2 mixer bunyi; Pemasangan pressure regulator tambahan | Ready to use |  | Wahyu
7000591 | 2024-07-24 | RSUP dr. Soeradji Tirtonegoro | Bubble CPAP | Draeger/Seattle PAP | QM20230157 | Service | O2 mixer bunyi; Pemasangan pressure regulator tambahan | Ready to use |  | Wahyu
18000540 | 2024-07-24 | RSUD Sele Be Solu | Mobile X-ray | Apelem/ Rafale DR | 8112411332 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Irfan R
19000753 | 2024-07-24 | RSUD Dr. Moewardi | All unit draeger | Draeger | x | Survey | Pengecekan alat untuk pengajaun maintenance tahun depan | Ready to use | Ada kerusakan engsel pada incubator | Nanda
2000185 | 2024-07-24 | RSUD Pringsewu | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0756 | Service | Setting ulang kolimasi | Ready to use |  | Eddy
9000452 | 2024-07-24 | RSUD Ainun Habibie | Infant Warmer | Draeger/Resuscitaire | DY03437 | Survey | Unit mati total; Pengecekan board power supply | Ready to use |  | Ruslan
8000579 | 2024-07-24 | RSUD Kab Batang | Incubator Transport | Draeger/TI500 | CF11157 | Service | Penggantian baterai | Ready to use |  | Amna
18000540 | 2024-07-24 | RSUD Sele Be Solu | Mobile X-ray | Apelem/ Rafale DR | 8112411332 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
6000726 | 2024-07-24 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASTF0005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
20000279 | 2024-07-23 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASMF0003 | Survey | Rotari knop error; Perlu dilakukan penggantian rotari knop | Pending |  | Sujana
18000541 | 2024-07-23 | RSUD Sele Be Solu | Meja Operasi | SYAS/OT Elektrohidrolik | 6234030004240300241 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
7000578 | 2024-07-23 | RSPAU dr. S. Hardjolukito | Mesin Anasthesia | Draeger/Fabius Plus | ASDL0084 | Survey | O2 fail, battery fail, gas bocor; Cek koneksi selang | Ready to use | Unit bisa digunakan dengan keterbatasan fungsi. Disarankan penggantian service set 3 tahun dan O2 cell | Wahyu
19000741 | 2024-07-23 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Service | Penggantian probe USG | Ready to use |  | Nanda
9000432 | 2024-07-23 | RSUD Toto Kabila | Central Monitor | Draeger/Vista 120 CMS | VLGNF0251 | Survey | Tidak terdeteksi di monitor; Router dan kabel LAN rusak | Pending | Disarankan penggantian router dan kabel | Ruslan
8000577 | 2024-07-23 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Babylog VN600 | ASSK0028 | Service | Kondensasi terlalu tinggi; Setting humidifier | Ready to use |  | Amna
6000727 | 2024-07-23 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Babylog VN800 | ASTE0047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
11000316 | 2024-07-23 | BPAFK Makassar | Baby Incubator | Draeger/Isolette 8000 Plus | DSSJ0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
3000317 | 2024-07-23 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Imam
20000277 | 2024-07-22 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina300 | ASJJ0403 | Service | Baterai external dan internal fail | Pending | Perlu penggantian service set 2 tahun dan baterai external | Sujana
20000277 | 2024-07-22 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina300 | ASJJ0405 | Service | Baterai external dan internal fail | Pending | Perlu penggantian service set 2 tahun dan baterai external | Sujana
20000277 | 2024-07-22 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina300 | ASJJ0404 | Service | Baterai external dan internal fail | Pending | Perlu penggantian service set 2 tahun dan baterai external | Sujana
20000277 | 2024-07-22 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina300 | ASJJ0409 | Service | Baterai external dan internal fail | Pending | Perlu penggantian service set 2 tahun dan baterai external | Sujana
20000277 | 2024-07-22 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina300 | ASHJ0007 | Service | Baterai external dan internal fail | Pending | Perlu penggantian service set 2 tahun dan baterai external | Sujana
20000278 | 2024-07-22 | RS Universitas Hasanuddin | Lampu Operasi | Draeger/Polaris 600 | ASRF1088 | Service | Camera tidak berfungsi, pengecekan unit wareless dan wireless reciever | Pending | Disarankan penggantian wireless video, reciever box, dan adaptor TV LG | Sujana
9000541 | 2024-07-22 | RSUD Tambulilato | Bed Patient | SYAS/Premium 3 crank | 24036240400241 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
19000759 | 2024-07-22 | RSUD Simo Boyolali | Mesin Anasthesia | Draeger/Fabius Plus XL | ASTC0003 | Survey | Pendampingan ke pasien | Ready to use |  | Nanda
19000759 | 2024-07-22 | RSUD Simo Boyolali | Mesin Anasthesia | Draeger/Fabius Plus XL | ASTC0004 | Survey | Pendampingan ke pasien | Ready to use |  | Nanda
6000725 | 2024-07-22 | RSUD dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASJF0034 | Survey | Display mati; MoBi display dibersihkan, unit sudah menyala | Pending | Unit menyala namun tidak optimal berfungsi; Disarankan penggantian display kit, primus service set 6 tahun, primus service set br system 2 thn | Afrizal
18000541 | 2024-07-22 | RSUD Sele Be Solu | Meja Operasi | SYAS/OT Elektrohidrolik | 6234030004240300241 | Test Function | Alat berfungsi dengan baik | Ready to use |  | Irfan R
8000576 | 2024-07-22 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASNH0240 | Survey | O2 sensor fail, gas sensor fail; Flow power rusak | Pending | Perlu diganti flow power | Amna
10000198 | 2024-07-21 | RSUD Mas Amsyar | Pendant Surgery | SYAS/Pegasus | SYPS2402004 | Instalasi | Alat berfungsi dengan baik dan siap digunakan | Ready to use |  | Heru
12000117 | 2024-07-21 | RSUD Toto Kabila | C-arm | Apelem/Evo R+ | 55121004853 | Survey | Overload voltage; Trafo, power board, power supply, inverter rusak | Pending |  | Marshall
2000181 | 2024-07-21 | RSUD Pagelaran (KSO) | CT Scan | Philips | 445002805 | Service | Pergeseran posisi gantry dan install ulang | Ready to use |  | Eddy
10000197 | 2024-07-19 | RSUD Mas Amsyar | Pendant Anasthesi | SYAS/Pegasus | SYPA2402003 | Instalasi | Alat berfungsi dengan baik dan siap digunakan | Ready to use |  | Heru
20000266 | 2024-07-19 | RSUD Madani Palu | Ventilator | Draeger/Savina300 | ASSL0013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000268 | 2024-07-19 | RSUD Madani Palu | Ventilator | Draeger/Babylog VN800 | ASTA0015 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
19000751 | 2024-07-19 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0091 | Survey | Blower fail; Blower rusak | Pending | Diajukan claim warranty penggantian blower | Nanda
7000577 | 2024-07-19 | RSUD Bagas Waras | Central Monitor | Draeger/Vista 120 CMS | VMSNG0015 | Survey | CPU Mati; Pengecekan power supply, ganti kabel SATA | Pending | CPU central perlu di service atau upgrade CPU | Wahyu
8000575 | 2024-07-19 | RSJD Dr. Amino Gondohutomo | Mesin Anasthesia | Draeger/Atlan A300 | ASTF0110 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
19000739 | 2024-07-19 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0091 | Service | Gas supply unit fail; Blower rusak | Pending | Disarankan penggantian blower. Proses claim warranty | Nanda
7000576 | 2024-07-18 | RSUD Wonosari | ESWL | Medispec/E3000 | 722027 | Service | Layar c-arm mati; Kencangkan koneksi power dan display | Ready to use |  | Wahyu
6000717 | 2024-07-18 | RSUD Cilacap | Bubble CPAP | Draeger/Seattle PAP | QM20230175 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
11000314 | 2024-07-18 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Survey | EMH tidak mengembang & mengempis; Motor aktuator error | Pending | Perlu penggantian motor actuator | Nasrul
11000315 | 2024-07-17 | RSUP Dr. Wahidin Sudirohusodo | Infant Warmer | Draeger/8004 | ASEL0092 | Survey | Error 35; Heater rusak | Pending | Disarankan penggantian heater | Nasrul
11000318 | 2024-07-17 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASJJ0106 | Survey | Baterai internal failure, air filter kotor, partikel filter kotor | Ready to use | Disarankan penggantian maintenance kit | Nasrul
11000318 | 2024-07-17 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASLA0080 | Survey | Baterai internal failure, air filter kotor, partikel filter kotor | Ready to use | Disarankan penggantian maintenance kit | Nasrul
11000318 | 2024-07-17 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASJE0004 | Survey | Baterai internal failure, air filter kotor, partikel filter kotor | Ready to use | Disarankan penggantian maintenance kit | Nasrul
11000318 | 2024-07-17 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASLA0079 | Survey | Baterai internal failure, air filter kotor, partikel filter kotor | Ready to use | Disarankan penggantian maintenance kit | Nasrul
11000318 | 2024-07-17 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASJJ0107 | Survey | Baterai internal failure, air filter kotor, partikel filter kotor | Ready to use | Disarankan penggantian maintenance kit | Nasrul
11000318 | 2024-07-17 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASLA0081 | Survey | Baterai internal failure, air filter kotor, partikel filter kotor | Ready to use | Disarankan penggantian maintenance kit | Nasrul
11000318 | 2024-07-17 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASED0284 | Survey | Baterai internal failure, air filter kotor, partikel filter kotor | Ready to use | Disarankan penggantian maintenance kit | Nasrul
9000455 | 2024-07-17 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASNN0144 | Survey | Device check fail, ventilator piston fail, baterai drop; Board PBA VGC power rusak dan baterai internal rusak | Pending | Perlu penggantian baterai internal dan board PBA VGC | Ruslan
19000738 | 2024-07-17 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE5098 | Service | Modul wifi tidak bisa; Kanibal board wifi dengan unit yang rusak LCD nya | Ready to use |  | Nanda
3000315 | 2024-07-17 | RSUD RA Kartini | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/006/C0 | Survey | 3D Tomo tidak muncul, layar display mati; Setting ulang dicom | Pending | Menunggu hasil analisa pihak metaltronica | Imam
18000536 | 2024-07-16 | RSUD Dr. H. Chasan Boesoirie | UPS Online | Delta 80kVA | 163080T0029 | Service | Baterai masih bermasalah; Penggantian baterai 30pcs | Ready to use |  | Irfan R
19000758 | 2024-07-16 | RSUD Simo Boyolali | Mesin Anasthesia | Draeger/Fabius Plus XL | ASTC0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000758 | 2024-07-16 | RSUD Simo Boyolali | Mesin Anasthesia | Draeger/Fabius Plus XL | ASTC0004 | Instalasi | alat berfungsi dengan baik | Ready to use |  | Nanda
7000575 | 2024-07-16 | RSUD KOTA YOGYAKARTA | Ventilator | Draeger/Babylog VN800 | ASTE0030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
19000740 | 2024-07-16 | RSUP dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0303 | Service | Expriation flow sensor fail calibration; Penggantian kabel flow sensor | Ready to use |  | Nanda
7000574 | 2024-07-15 | RSUD Muntilan | Ventilator | Draeger/Babylog VN800 | ASRF0013 | Service | Layar mati; Pengencangan konektor ECD adopter | Ready to use |  | Wahyu
6000718 | 2024-07-15 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0010 | Survey | NVRAM error, APL fail, no fresh gas; Pba control board rusak | Pending | Perlu penggantian pba control 3, fabius gs service set 3 tahun, maintenance kit, O2 sensir capsule | Afrizal
19000737 | 2024-07-15 | RSUP dr. Kariadi | Paykel | Fisher/MR850 | 210831585865 | Service | Error E51, fuse putus; Penggantian fuse | Ready to use |  | Nanda
6000718 | 2024-07-15 | RSI Harapan Anda | Ventilator | Draeger/Babylog VN500 | ASJC0050 | Survey | Pengecekan software dan hardware | Pending | Perlu penggantian pressure regulator, kit air filter | Afrizal
10000189 | 2024-07-15 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Penggantian connector LAN, setting kv - mAs console | Ready to use |  | Heru
7000573 | 2024-07-12 | RS PKU Muhammadiyah Bantul | Mesin Anasthesia | Draeger/Fabius Premium | ASEH0212 | Service | Pemasangan service set 2 tahun | Ready to use |  | Wahyu
19000736 | 2024-07-12 | RSUD K.R.M.T Wongsonegoro | Dental Unit | Kavo/Primus 1058 | 10205318 | Service | Filter air kotor; Pembersihan filter air | Ready to use |  | Nanda
8000573 | 2024-07-12 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASJF0033 | Service | Pemasangan service set 6tahun dan breathing Cozy, install diapram piston, cup dan filter | Ready to use |  | Amna
8000573 | 2024-07-12 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Service | Pemasangan service set 6tahun dan breathing Cozy, install diapram piston, cup dan filter | Ready to use |  | Amna
8000573 | 2024-07-12 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Service | Pemasangan service set 6tahun dan breathing Cozy, install diapram piston, cup dan filter | Ready to use |  | Amna
10000188 | 2024-07-11 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Uji fungsi setelah upgrade software | Ready to use |  | Heru
18000533 | 2024-07-11 | RSUD Labuha | UPS | Yoshiga/20kVA | E2211C012260621353 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000533 | 2024-07-11 | RSUD Labuha | UPS | Yoshiga/80kVA | 501201173120M880002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
16000244 | 2024-07-11 | Signature Dental Care | Dental Unit | Kavo/Primus 1058 | 10207194 | Service | Conexio pada dental unit tidak terkoneksi dengan foot control, pengecekan LAN dan sambungan conexio | Ready to use |  | Ipan
6000715 | 2024-07-11 | RSUD Ajibarang | Ventilator | Draeger/Babylog VN500 | ASMK0054 | Survey | Kerusakan pada flow sensor connecting cable; Pemasangan kabel koneksi flow sensor | Ready to use |  | Afrizal
9000416 | 2024-07-11 | RSUD Toto Kabila | Dental Unit | Kavo/Primus 1058 | 10205320 | Service | Muncul error red request dan water block pressure; Reset pada menu service dan bersihkan jalur air dental unit | Ready to use |  | Ruslan
7000572 | 2024-07-11 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Wahyu
19000760 | 2024-07-11 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Atlan A300 | ASPH0141 | Survey | General check | Ready to use |  | Nanda
3000314 | 2024-07-11 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | UPS rusak, detector tidak konek; Upgrade software dan perbaikan UPS | Ready to use |  | Imam
6000714 | 2024-07-10 | RSUD Prembun | CPAP | Draeger/Seattle PAP | QM20230153 | Survey | Pemasangan konektor CPAP | Ready to use |  | Afrizal
6000714 | 2024-07-10 | RSUD Prembun | CPAP | Draeger/Seattle PAP | QM20230154 | Survey | Pemasangan konektor CPAP | Ready to use |  | Afrizal
9000449 | 2024-07-10 | RSUD Tani dan Nelayan | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0080 | Service | Baterai internal fail; Penggantian baterai internal | Ready to use |  | Ruslan
18000534 | 2024-07-09 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista120 | VGSNE2298 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000534 | 2024-07-09 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista120 | VGSRF0245 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000534 | 2024-07-09 | RSUD Dr. H. Chasan Boesoirie | Patient Monitor | Draeger/Vista120 | VGSRF0247 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000448 | 2024-07-09 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0480 | Service | Baterai internal fail; Penggantian baterai internal | Ready to use |  | Ruslan
9000448 | 2024-07-09 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0476 | Service | Baterai internal fail; Penggantian baterai internal | Ready to use |  | Ruslan
9000448 | 2024-07-09 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0481 | Service | Baterai internal fail; Penggantian baterai internal | Ready to use |  | Ruslan
9000448 | 2024-07-09 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASJI0135 | Service | Baterai internal fail; Penggantian baterai internal | Ready to use |  | Ruslan
17000185 | 2024-07-09 | RSUD Dr. H. Chasan Boesoirie | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/089/C1 | Service | Unit tidak bisa menyala otomatis; Perbaikan sambungan kabel tegangan | Ready to use |  | Bahar
20000275 | 2024-07-09 | RSUP Tadjudin Chalid | ESWL | Medispec/EM1000 | 0922014 | Service | MCU error 06 - Thermal safety; MCU board rusak | Pending | Penggantian board MCU | Sujana
20000274 | 2024-07-08 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina300 | ASEF0116 | Service | Mati total; Penggantian power supply | Ready to use |  | Sujana
9000447 | 2024-07-08 | RSUD Ainun Habibie | Ventilator | Draeger/Evita V800 | ASTC0006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
6000719 | 2024-07-08 | RSU An Ni'mah | Ventilator | Draeger/Babylog VN600 | ASSK0026 | Survey | Training dengan user | Ready to use |  | Afrizal
17000184 | 2024-07-08 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | 2510220561 | Service | Kabel ceilling tidak bisa bergerak, wall stand tidak stabil; Perbaikan kabel ceilling, adjustment pembelat wall stand | Ready to use |  | Bahar
6000720 | 2024-07-08 | RSUD Prof. Dr. Margono Soekarjo | Jaudience Meter | Draeger/JM105 | B3503802 | Service | Pemasangan baterai internat | Ready to use |  | Afrizal
8000572 | 2024-07-08 | RSUD Sunan Kalijaga | Baby Incubator | Draeger/Isolette C2000 | CF32801 | Service | Air tempurature alarm low. Reset heater | Ready to use |  | Amna
19000733 | 2024-07-06 | RSUD Simo Boyolali | Mesin Anasthesia | Draeger/Fabius Plus XL | ASTC0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000733 | 2024-07-06 | RSUD Simo Boyolali | Mesin Anasthesia | Draeger/Fabius Plus XL | ASTC0004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
7000569 | 2024-07-05 | RSUP dr. Soeradji Tirtonegoro | Bubble CPAP | Draeger/Oxymixer | QM20230155 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000569 | 2024-07-05 | RSUP dr. Soeradji Tirtonegoro | Bubble CPAP | Draeger/Oxymixer | QM20230157 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
2000184 | 2024-07-05 | RSUD Nyi Ageng Serang (KSO) | CT Scan | Siemens/Somatom | 120196 | Instalasi | Pemasangan kabel power ke box panel, kabel power ke UPS dan CT Scan dan tanam baut gantry | Ready to use | Unit sudah normal | Eddy
18000535 | 2024-07-05 | RSUD Dr. H. Chasan Boesoirie | Mesin Anasthesia | Draeger/Perseus A500 | ASSK0059 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
12000123 | 2024-07-05 | RSUD Ainun Habibie | Ventilator | Draeger/Evita V800 | ASTC0006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Marshall
8000571 | 2024-07-05 | RSI Sultan Agung Banjarbaru | Mesin Anasthesia | Draeger/Atlan A300XL | ASTA0171 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
6000713 | 2024-07-05 | RS JIH | CPAP | Draeger/Seattle PAP | QM20230185 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
12000122 | 2024-07-04 | RSUD Toto Kabila | C-arm | Apelem/Evo R+ | 55121004853 | Survey | Overload voltage; Trafo, power board, power supply, inverter rusak | Pending |  | Marshall
19000731 | 2024-07-04 | RSJ Prof Dr Soerojo Magelang | Incubator Transport | Draeger/TI500 | CH12771 | Survey | Tidak bisa charger; Baterai internal drop | Pending | Disarankan penggantian baterai internal | Nanda
19000730 | 2024-07-04 | RSJ Prof Dr Soerojo Magelang | Patient Monitor | Draeger/Vista 120 | VTSKF0078 | Service | Layar stuck logo draeger/Reconnect kabel koneksi | Ready to use |  | Nanda
6000712 | 2024-07-04 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASKK0221 | Survey | Error 10, baterai drop | Pending | Perlu penggantian baterai internal, baterai eksternal, filter air, filter O2, filter electronic | Afrizal
15000249 | 2024-07-04 | RSUD Raden Mattaher | X-ray Analog | Apelem/Camargue Analog | IMA703863 | Service | Tensator spring rusak; Penggantian tensator spring | Ready to use |  | Sutan
20000278 | 2024-07-03 | RSKD Gigi dan Mulut Makassar | Dental Unit | Kavo/Primus 1058 | 10207188 | Survey | Unit mati total; PCBA 3001 rusak | Pending | Perlu penggantian PCBA Unit | Sujana
12000121 | 2024-07-03 | RSUD Toto Kabila | C-arm | Apelem/Evo R+ | 55121004853 | Survey | Overload voltage (Masih proses troubleshoot) | Pending |  | Marshall
19000729 | 2024-07-03 | BKIM Provinsi Jawa Tengah | Patient Monitor | Draeger/Vista120 | VGSRC0031 | Maintenance | Preventive Maintenance; General Check | Ready to use |  | Nanda
19000729 | 2024-07-03 | BKIM Provinsi Jawa Tengah | Patient Monitor | Draeger/Vista120 | VGSRC0029 | Maintenance | Preventive Maintenance; General Check | Ready to use |  | Nanda
19000729 | 2024-07-03 | BKIM Provinsi Jawa Tengah | Patient Monitor | Draeger/Vista120 | VGSNH1039 | Maintenance | Preventive Maintenance; General Check | Ready to use |  | Nanda
19000729 | 2024-07-03 | BKIM Provinsi Jawa Tengah | Patient Monitor | Draeger/Vista120 | VGSNH1335 | Maintenance | Preventive Maintenance; General Check | Ready to use |  | Nanda
6000705 | 2024-07-02 | RSUD dr. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEL0081 | Survey | Tidak ada O2 dan udara masuk; Cek software dan hardware | Pending | Perlu penggantian kit check valve, O2 ring, proportional valve, flow sensor, filter, baterai pack | Afrizal
16000243 | 2024-07-02 | RSUD Kab. Bekasi | Bed Patient | SYAS/ Noble 4 Motor | 240579101260400106 | Service | Pengecekan unit karna terlalu panjang; Setting panjang bed | Ready to use |  | Ipan
7000570 | 2024-07-02 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Service | Error #11 ; Cek konektor power ke HVPS | Ready to use |  | Wahyu
7000571 | 2024-07-02 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Maintenance | PM Total Care, water filter cleaning dan refill air | Ready to use |  | Wahyu
11000347 | 2024-07-02 | RSUD Kota Kendari | ESWL | Medispec/E3000 | 0922003 | Maintenance | Preventinve Maintenance Warranty | Ready to use |  | Nasrul
11000348 | 2024-07-02 | RSUD Kota Kendari | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/127/C2 | Maintenance | Preventinve Maintenance Warranty | Ready to use |  | Nasrul
11000346 | 2024-07-01 | RSUD Bahteramas | Mesin Anasthesia | Draeger/Primus | ASNH0249 | Survey | Ventilator failed | Pending | Disarankan penggantian board cpu dan baterai internal | Nasrul
16000241 | 2024-07-01 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Survey | Burr handpiece tidak berputar; bearing handpiece bengkok pasca jatuh | Ready to use | Sudah ditawarkan penggantian bearing | Ipan
16000242 | 2024-07-01 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Service | Handpiece low speed macet; O-ring dan copping handpiece berminyak, sudah dibersihkan | Ready to use |  | Ipan
19000728 | 2024-07-01 | RSUP dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0303 | Service | Kabel flow sensor fail; Ganti service set 6 tahun | Ready to use |  | Nanda
14000316 | 2024-07-01 | RSUD Dr. Moh. Saleh | Kursi roda | SYAS/SW01-SS | 24023131300400342 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000316 | 2024-07-01 | RSUD Dr. Moh. Saleh | Kursi roda | SYAS/SW01-SS | 24043131300400785 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000316 | 2024-07-01 | RSUD Dr. Moh. Saleh | Kursi roda | SYAS/SW01-SS | 240431313000400685 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000315 | 2024-07-01 | RSUD Dr. Moh. Saleh | Nobie Bed | SYAS/SBW4 | 240679101260400127 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
| 2024-07-01 | RSUD Dr. Moh. Saleh | Nobie Bed | SYAS/SBW4 | 240679101260400128 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
2000183 | 2024-06-29 | RSUD Nyi Ageng Serang (KSO) | CT Scan | Siemens/Somatom | 120196 | Instalasi | Install unit dan UPS | Pending | Unit belum menyala karna listrik utama belum diinstall | Eddy
6000708 | 2024-06-28 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0013 | Survey | Error O2 sensor failure; PATO kadaluarsa | Pending | Perlu penggantian PATO sensor | Afrizal
18000527 | 2024-06-28 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASHD0033 | Survey | Peminjaman board untuk troubleshoot | Ready to use |  | Irfan R
8000569 | 2024-06-28 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0014 | Maintenance | General check, system check, running test | Ready to use |  | Amna
8000569 | 2024-06-28 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0011 | Maintenance | General check, system check, running test | Ready to use |  | Amna
3000313 | 2024-06-28 | RSUD Malingping | X-ray Digital | Apelem/Camargue DR | IMA703916 | Survey | Detektor tidak konek ke workstation; Baterai detektor rusak | Pending |  | Imam
11000313 | 2024-06-27 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Perseus A500 | ASNK0060 | Service | Leakage breathing system, flow sensor failed; Penggantian flow sensor | Ready to use |  | Nasrul
6000704 | 2024-06-27 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0082 | Survey | Gas sensor failure; Penggantian waterlock dan sample line | Ready to use |  | Afrizal
6000703 | 2024-06-27 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08777 | Service | Pemasangan kit door lacth dan hinge access panel right front | Ready to use |  | Afrizal
6000703 | 2024-06-27 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08779 | Service | Pemasangan kit door latch | Ready to use |  | Afrizal
6000703 | 2024-06-27 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08776 | Service | Pemasangan kit door latch | Ready to use |  | Afrizal
6000703 | 2024-06-27 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08778 | Service | Pemasangan kit door latch | Ready to use |  | Afrizal
6000703 | 2024-06-27 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | AF32541 | Service | Pemasangan kit door latch | Ready to use |  | Afrizal
6000703 | 2024-06-27 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | DSNF62 | Service | Pemasangan kit door latch | Ready to use |  | Afrizal
10000196 | 2024-06-27 | RSUD Sultan Suriansyah | X-ray Digital | Apelem/Camargue DR | CPD33483019 | Survey | Kunjungan unit | Ready to use | r | Heru
11000312 | 2024-06-26 | RSUP Makassar | Ventilator | Draeger/Evita V300 | ASHA0111 | Survey | GS5000 failure, baterai faild; Blower rusak, baterai internal dan eksteral rusak | Pending | Disarankan penggantian maintenance kit dan upgrade blower/ganti komporessor | Nasrul
19000735 | 2024-06-26 | RSUP dr. Kariadi | Dental unit | Kavo/Primus 1058 | 3000038 | Service | Handpiece rusak; Penggantian handpiece | Ready to use |  | Nanda
15000248 | 2024-06-26 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 9115101095 | Service | Error anode; Penggantian board S22 | Ready to use |  | Sutan
9000443 | 2024-06-26 | RSUD Toto Kabila | Pendant | SYAS/MOT | SY-PS2402001 | Instalasi | Alat berfungsi dengan baik. Sudah uji fungsi dan training ke user | Ready to use |  | Ruslan
9000444 | 2024-06-26 | RSUD Toto Kabila | Lampu Operasi | Draeger/Polaris 600 | ASML0274 | Test Function | Pemasangan unit lampu operasi | Ready to use |  | Ruslan
8000568 | 2024-06-26 | RS ISLAM PATI | Lampu Operasi | Draeger/Polaris 600 | ASRK0101 | Maintenance |  | Ready to use |  | Amna
8000567 | 2024-06-26 | RS ISLAM PATI | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0114 | Service | Pemasangan service set 2 tahun dan O2 sensor | Ready to use |  | Amna
18000528 | 2024-06-26 | RSUD Provinsi Sulawesi Utara | Mesin Anasthesia | Draeger/Atlan | ASRH0115 | Survey | O2 tidak terbaca; Tekanan O2 tidak sampai 4 bar | Ready to use |  | Irfan R
6000702 | 2024-06-26 | RSUD Brebes | Ventilator | Draeger/Babylog VN800 | ASTD0017 | Survey | Training dengan user | Ready to use |  | Afrizal
6000702 | 2024-06-26 | RSUD Brebes | Ventilator | Draeger/Evita V300 |  | Survey | Gas supply error; Setting ulang medical compressor | Ready to use |  | Afrizal
14000310 | 2024-06-26 | RSUD Ar Rozy | Baby Bed | SYAS/SBB2 | 240526211000400114 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000310 | 2024-06-26 | RSUD Ar Rozy | Baby Bed | SYAS/SBB2 | 240526211000400120 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
6000701 | 2024-06-25 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASJH0125 | Service | Pemasangan flow sensor, kit air filter, pressure regulator, partical filter, baterai pack, filter ipc dan table harness spirolog flow sensor | Ready to use |  | Afrizal
17000186 | 2024-06-25 | RSUD KRT. Setjonegoro Wonosobo | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/044/C0 | Survey | Pendampingan Ukes | Ready to use |  | Bahar
19000727 | 2024-06-25 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRH0033 | Service | Blower rusak; Dilakukan penggantian blowe | Ready to use |  | Nanda
6000724 | 2024-06-24 | RSUD Dr. Soeselo Slawi | Meja Operasi | SYAS/Epione OP850 | OP850XVZX00037 | Instalasi | Alat berfungsi dengan baik, aksesoris lengkap, pergerakan semua baik | Ready to use |  | Afrizal
14000317 | 2024-06-23 | RSUD Gambiran Kediri | C-arm | Apelem/Evo R+ | 55121001850 | Service | Camera ii bocor; Cek seal dan baut. Baut kendur | Ready to use |  | Purwono
2000273 | 2024-06-21 | RSUD Andi Djemma Masamba | Ventilator | Draeger/Savina 300 | ASHD0010 | Survey | Unit mati secara tiba-tiba; Power supply rusak | Pending | Perlu penggantian power supply dan baterai | Sujana
6000711 | 2024-06-21 | RSI Harapan Anda | Ventilator | Draeger/Savina 300 | ASFJ0001 | Service | Unit suka restart sendiri; Pemasangan oxyrace dan service set 6 tahun | Ready to use |  | Afrizal
6000711 | 2024-06-21 | RSI Harapan Anda | Ventilator | Draeger/Savina 300 | ASFJ0095 | Service | Unit suka restart sendiri; Pemasangan oxyrace dan service set 6 tahun | Ready to use |  | Afrizal
19000726 | 2024-06-20 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Service | Layar freezing; Cek koneksi kabel dan penggantian air | Ready to use |  | Nanda
6000710 | 2024-06-20 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0010 | Survey | NV RAM Fail; Pengecekan software dan hardware | Pending | Disarankan penggantian PBA control 3, Pressure sensor, Co2y service 6 tahun, O2 sensor | Afrizal
18000522 | 2024-06-20 | RSUD Dr. Iwan Boking | Ventilator | Draeger/Savina300 | ASJL0134 | Service | Penggantian baterai internal | Ready to use |  | Irfan R
18000523 | 2024-06-20 | RSUD Dr. Iwan Boking | Bed Patient | SYAS/Premium 3 crank | 240462404000400227 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000523 | 2024-06-20 | RSUD Dr. Iwan Boking | Bed Patient | SYAS/Premium 3 crank | 240462404000400228 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
8000566 | 2024-06-20 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius CE | ARYJ0111 | Survey | O2 flush button fail; Setting valve dan spring push button | Pending | Perlu penggantian O2 flush button | Amna
8000565 | 2024-06-19 | RSI Sultan Agung Semarang | Infant Warmer | Draeger/Isolette C2000 | DSSL0049 | Service | O2 module not installed; Install Oxy-control valve assy, O2 kit sensor | Ready to use |  | Amna
16000240 | 2024-06-19 | RSUD Kab. Bekasi | Bed Pediatric | SYAS | 240526302000400042 | Service | Kaki tidak bisa naik; Pengecekan pegas, sambungan pegas copot. Sudah dipasang kembali | Ready to use |  | Ipan
3000312 | 2024-06-19 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Service | Potter bucky tidak mengunci; Setting poter bucky lebih dalam | Ready to use |  | Imam
9000439 | 2024-06-15 | RSUD Weda | Dental Unit | Kavo/1058 Life | 10205459 | Service | Board dentist element rusak; Penggantian board dentist elemet | Ready to use |  | Ruslan
2000182 | 2024-06-15 | RSUD Sumbawa (KSO) | CT Scan | Siemens | 120196 | Uninstall |  | Ready to use |  | Eddy
11000311 | 2024-06-14 | RSUD Ainun Habibie | Ventilator | Draeger/Babylog VN800 | ASSM0042 | Survey | Device check failed; Pengecekan O2 supply | Ready to use |  | Nasrul
12000111 | 2024-06-14 | RSUD Tambulilato | X-ray Digital | Apelem/Camargue DR | CPD36814320 | Service | Penggantian ups komputer dan kabel LAN | Ready to use |  | Marshall
18000530 | 2024-06-14 | RSUD Toto Kabila | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150173 | Survey | Alat tidak mengeluarkan laser; Alat masih diperiksa lebih lanjut | Pending |  | Irfan R
2000272 | 2024-06-14 | RSUD Luwuk | Baby Incubator | Draeger/Caleo | ASHM0040 | Survey | Pengecekan unit | Pending | Perlu dilakukan penggantian sensor skin dan casa | Sujana
1600239 | 2024-06-13 | Signature Dental Care | Dental Unit | Kavo/1058 | 10207194 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
7000600 | 2024-06-13 | RS Duta Mulya | Ventilator | Draeger/Savina 300 | ASTB0088 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
18000532 | 2024-06-13 | RSUD Ainun Habibie | Infant Warmer | Draeger/Resuscitaire | OY03435 | Survey | On off oxygen bermasalah | Pending |  | Irfan R
18000532 | 2024-06-13 | RSUD Ainun Habibie | Infant Warmer | Draeger/Resuscitaire | IG22652 | Survey | On off oxygen bermasalah | Pending |  | Irfan R
2000271 | 2024-06-13 | RSUD Trikora Salakan | Mesin Anasthesia | Draeger/Primus | ASHD0036 | Survey | Ventilator fail, pengecekan unit | Pending | Perlu penggantian service set dan baterai | Sujana
7000600 | 2024-06-13 | RS Duta Mulya | Ventilator | Draeger/Savina 300 | ASTB0091 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
8000564 | 2024-06-13 | RS Bening Kasih | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ0082 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
3000311 | 2024-06-13 | RSUD Balaraja | C-arm | Apelem/Evo R+ | 55118034657 | Service | Slide brake rusak; Penggantian slide break | Ready to use |  | Imam
6000723 | 2024-06-12 | RSUD Bumiayu | Incubator Transport | Draeger/TI500 | AY11232 | Survey | Baterai drop | Pending | Disarankan penggnatian baterai | Afrizal
6000723 | 2024-06-12 | RSUD Bumiayu | Baby Incubator | Draeger/Babylog VN500 | ASLJ0005 | Survey | Engsel patah | Pending | Disarankan penggantian engsel | Afrizal
19000725 | 2024-06-12 | RSUD Dr. H. Soewondo Kendal | Incubator Transport | Draeger/TI500 | KF10742 | Survey | Baterai tidak mengisi; Baterai drop | Pending | Disarankan penggantian baterai internal | Nanda
19000724 | 2024-06-12 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0149 | Maintenance | General check, system check, running test | Ready to use | Baterai drop | Nanda
19000724 | 2024-06-12 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0150 | Maintenance | General check, system check, running test | Ready to use | Baterai drop | Nanda
6000800 | 2024-06-11 | RSUD dr. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEL0081 | Survey | Error selftest, gas supply sensor calibration; Cek software dan hardware. Alat kanibal dengan stock gudang, diambil part board 1Pb M13 | Ready to use | Disarankan penggantian board 1Pb M13, baterai internal, kit check valve, air filter, particle filter, dan filter | Afrizal
18000531 | 2024-06-11 | RSUD Toto Kabila | C-arm | Apelem/Evo R+ | 55121004853 | Survey | Alat mati total; Travo unit bermasalah | Pending | Travo coba diperbaiki | Irfan R
14000309 | 2024-06-11 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 0620007 | Service | Monitor tidak menyala; Baterai cmos tidak ada tegangan, penggantian baterai cmos | Ready to use |  | Purwono
16000238 | 2024-06-11 | Kesumah Dental Clinic | Dental Unit | Kavo/E30 TM | 10111157 | Service | Blank pada display dan alarm; Penggantian board dentist element | Ready to use |  | Ipan
9000436 | 2024-06-11 | RSD Tidore | Baby Incubator | Draeger/Isolette 8000 | HY06998 | Service | Hinge access panel right front rusak; Penggantian ginge access panelr | Ready to use |  | Ruslan
9000437 | 2024-06-11 | RSD Tidore | Patient Monitor | Draeger/Vista 120 | VTSLB0086 | Survey | NIBP fail | Pending | Disarankan penggantian manset | Ruslan
6000721 | 2024-06-11 | RSUD dr. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEL0081 | Survey | O2 supply merah, air supply merah, pressure sensor merah; Pengecekan software dan hardware | Ready to use | Diback up dengan spare part babylog yang digudang, unit bisa digunakan dalam pengawasan | Afrizal
8000574 | 2024-06-11 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Service | Ventilation fail; Setting diapram piston | Ready to use |  | Amna
3000310 | 2024-06-11 | RSUP Tadjudin Chalid | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/133/C2 | Service | Tomo tidak aktif/ Install usb dongle 3D tomo | Ready to use |  | Imam
10000187 | 2024-06-10 | RSUD Ulin | Mamografi | Metaltronica/Helianthus BYM | HELI/0048/CB1 | Survey | Pemaparan fungsi pada anak PKL | Ready to use |  | Heru
8000563 | 2024-06-10 | RSUD Kab. Temanggung | Bone Mineral Densitometer | Apelem/Medix DR | H237H1093 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Amna
19000723 | 2024-06-10 | RSUP dr. Kariadi | Baby Incubator | Draeger/Caleo | ARUC0047 | Service | Humidity fail; Pasang conv. Heater | Ready to use |  | Nanda
3000309 | 2024-06-10 | RSUD Kab. Bekasi | C-arm | Apelem/Evo R+ | 55118026649 | Survey | Kv error; Indikasi kerusakan pada monoblock dan inverter | Pending |  | Imam
4000432 | 2024-06-08 | Kesumah Dental Clinic | Dental Unit | Kavo/E30 TM | 10111157 | Survey | Laya blank, update software, dentist cup not activated; Perlu penggantian board dentist element | Pending |  | Satria
8000578 | 2024-06-07 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina 300 | ASPJ0383 | Maintenance | General Check | Ready to use |  | Amna
8000578 | 2024-06-07 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina 300 | ASPJ0382 | Maintenance | General Check | Ready to use |  | Amna
8000578 | 2024-06-07 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Babylog VN600 | ASSK0028 | Maintenance | General Check | Ready to use |  | Amna
8000578 | 2024-06-07 | RS PKU Muhammadiyah Mayong | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ0081 | Maintenance | General Check | Ready to use |  | Amna
9000435 | 2024-06-07 | RSUD Sanana | Dental Unit | Kavo/E30TM | 10111150 | Service | Penggantian board dentist elemet, filter air, board scaller | Ready to use |  | Ruslan
8000554 | 2024-06-07 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Babylog VN600 | ASSK0028 | Maintenance | General check, system check, running test | Ready to use |  | Wahyu
8000554 | 2024-06-07 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina300 | ASPJ0382 | Maintenance | General check, system check, running test | Ready to use |  | Wahyu
8000554 | 2024-06-07 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina300 | ASPJ0383 | Maintenance | General check, system check, running test | Ready to use |  | Wahyu
8000554 | 2024-06-07 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Fabius Plus | ASPH4822 | Maintenance | General check, system check, running test | Ready to use |  | Wahyu
19000721 | 2024-06-06 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Isolette 8000 | MC09026 | Survey | Humidity low | Pending | Unit ditarik untuk di cek lebih lanjut | Nanda
19000722 | 2024-06-06 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Service | Air low, membran kendor; Setting adjustment dan calibration | Ready to use |  | Nanda
8000552 | 2024-06-06 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL003 | Survey | Alarm O2 delivery system failure; Sensor P Tank dan P DMIK rusak | Pending | Perlu penggantian sensor Ptank dan PDMIK | Wahyu
7000561 | 2024-06-06 | RSUP dr. Soeradji Tirtonegoro |  | Draeger/Fabius | ASHA0109 | Service | Layar bercak hitam; Penggantian part TFT LCD 5:6 | Ready to use |  | Wahyu
7000558 | 2024-06-05 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASLL0064 | Service | Laya blank; Penggantian LC Display M14.5 | Ready to use |  | Wahyu
7000559 | 2024-06-05 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASNL0106 | Maintenance | Total care Maintenance | Ready to use |  | Wahyu
7000560 | 2024-06-05 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0091 | Maintenance | Total care Maintenance | Ready to use |  | Wahyu
8000562 | 2024-06-05 | RS Bhina Bakti Husada | Fototerapi | Draeger/PT4000 | ASJL0006 | Service | Penggantian fluorescent lamp 4 unit | Ready to use |  | Amna
11000309 | 2024-06-04 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Penggantian EMH | Ready to use |  | Nasrul
8000561 | 2024-06-04 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0037 | Survey | Gas analyser dan valve check fail; Flow power, ILCA, motor pump, O2 sensor rusak | Pending | Perlu penggantian manitenance kit CO2y | Amna
7000557 | 2024-06-03 | RSUD Bagas Waras | Lampu Operasi | Draeger/Polaris 100 | ASJC0523 | Survey | Lampu mati total; Baterai drop | Pending | Perlu penggantian psu | Wahyu
2000180 | 2024-06-03 | RSUD Kab. Temanggung | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/011/C0 | Service | Tidak bisa print; Setting print dan koneksi LAN | Ready to use |  | Eddy
8000560 | 2024-06-03 | RSI Sultan Agung Semarang | Infant Warmer | Draeger/Isolette C2000 | CR35553 | Survey | Hinged access panel kanan failure | Pending | Perlu penggantian hinged access panel | Amna
3000308 | 2024-06-03 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Service | DICOM tidak konek ke PACS dan Worklist; Pemindahan jalur kabel LAN | Ready to use |  | Imam
9000440 | 2024-05-31 | RSUD Dr. H. Chasan Boesoirie | ESWL | Medispec/E3000 | 0922002 | Service | coater electroda merosot dan muncul error; Pembersihan filter air dan penggantian air baru | Ready to use |  | Ruslan
6000797 | 2024-05-31 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASLD0042 | Survey | Gas module terkadang tidak membaca hasil dari panel; Ada kabel yang terjepit | Ready to use |  | Afrizal
6000797 | 2024-05-31 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0043 | Survey | Pengecekan software dan hardware | Ready to use |  | Afrizal
6000797 | 2024-05-31 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0042 | Survey | VT kadang tidak sesuai standard | Pending | Disarankan penggantian kabel harness flow sensor | Afrizal
14000308 | 2024-05-30 | RSUD Grati | Kursi Roda | SYAS/ SW01SS | 240531300040115 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
8000559 | 2024-05-30 | RSUD dr. R. Soetijono Blora | Ventilator | Draeger/Savina 300 | ASJM0216 | Maintenance | Preventive Maintenance - General Check | Ready to use |  | Amna
8000559 | 2024-05-30 | RSUD dr. R. Soetijono Blora | Ventilator | Draeger/Babylog VN800 | ASLL0066 | Maintenance | Preventive Maintenance - General Check | Ready to use |  | Amna
8000559 | 2024-05-30 | RSUD dr. R. Soetijono Blora | Ventilator | Draeger/Babylog VN800 | ASLF0054 | Maintenance | Preventive Maintenance - General Check | Ready to use |  | Amna
8000559 | 2024-05-30 | RSUD dr. R. Soetijono Blora | Ventilator | Draeger/Babylog VN800 | ASKJ0004 | Maintenance | Preventive Maintenance - General Check | Ready to use |  | Amna
8000558 | 2024-05-29 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | ASML0039 | Service | Alarm O2 sensor fail; Penggantian O2 sensor dan housing arm | Ready to use |  | Amna
7000556 | 2024-05-29 | RSUD KOTA YOGYAKARTA | Jaudience Meter | Draeger/JM105 | B3503416 | Survey | Unit dibawa ke draeger untuk dikalibrasi | Ready to use |  | Wahyu
19000719 | 2024-05-29 | RSUD Dr. M. Ashari | Mesin Anasthesia | Draeger/Primus | ASLD0025 | Service | Gas measurement inop; Ada sumbatan di jalur gas measure, jalur dibersihkan | Ready to use |  | Nanda
19000719 | 2024-05-29 | RSUD Dr. M. Ashari | Mesin Anasthesia | Draeger/Primus | ASMF0114 | Service | Gas measurement inop; Ada sumbatan di jalur gas measure, jalur dibersihkan | Ready to use |  | Nanda
9000446 | 2024-05-29 | RSUD Tobelo | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/088/C1 | Service | Unit tidak konek ke console, muncul error sama daily check; Perbaikan LAN Card pada console dan sensor screen | Ready to use |  | Ruslan
6000796 | 2024-05-28 | RSUD Brebes | Ventilator | Draeger/Babylog VN800 | ASTD0017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
8000557 | 2024-05-28 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0148 | Maintenance | Preventive Maintenance - General Check | Ready to use |  | Amna
8000557 | 2024-05-28 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0115 | Maintenance | Preventive Maintenance - General Check | Ready to use |  | Amna
12000105 | 2024-05-28 | RSUD Noongan | UPS | Yoshiga | 56201500578200027 | Service | UPS abnormal saat digunakan | Pending | Disarankan memperbaiki kabel input dari PLN | Marshall
6000795 | 2024-05-27 | RS Duta Mulya | Ventilator | Draeger/Savina300 | ASKK0024 | Survey | Pengecekan unit; Unit berfungsi dengan baik | Ready to use |  | Afrizal
6000795 | 2024-05-27 | RS Duta Mulya | Ventilator | Draeger/Savina300 | ASRJ0121 | Survey | Pengecekan unit | Ready to use | Disarankan penggantian O2 sensor dan service set 2 tahun | Afrizal
6000795 | 2024-05-27 | RS Duta Mulya | Ventilator | Draeger/Savina300 | ASSK0023 | Survey | Pengecekan unit; Unit berfungsi dengan baik | Ready to use |  | Afrizal
9000434 | 2024-05-27 | RSUD Weda | Dental Unit | Kavo/Primus 1058 | 10205459 | Survey | LCD Mati, muncur bunyi alarm error; Pengecekan pada board dentist element | Pending | Disarankan penggantian board dentist element | Ruslan
12000107 | 2024-05-27 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina300 | ASPJ0085 | Service | Oxytrace rusak; Penggantian oxytrace atau O2 sensor | Ready to use |  | Marshall
2000179 | 2024-05-27 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Error table, meja tidak bisa bergerak; Reinstall parameter, kalibrawsi pergerakan meja | Ready to use | Disarankan penambahan UPS 3phase 120kVa untuk kestabilan arus listrik | Eddy
3000306 | 2024-05-27 | RSUD Dr. Drajat Prawiranegara | X-ray Digital | Apelem/ Camargue DR | 2510220562 | Survey | Generator error; Fuse 3phase putus, inverter board terbakar | Pending |  | Imam
3000307 | 2024-05-27 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Survey | Paddle tidak terdeteksi, kolimasi undefined; Cek paddle dan compression | Pending | Paddle compression perlu diganti | Imam
8000556 | 2024-05-22 | RSI Aisyah Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASMJ0048 | Maintenance | Preventive Maintenance - General Check | Ready to use |  | Amna
8000556 | 2024-05-22 | RSI Aisyah Kudus | Ventilator | Draeger/Babylog VN800 | ASPF0017 | Maintenance | Preventive Maintenance - General Check | Ready to use |  | Amna
8000556 | 2024-05-22 | RSI Aisyah Kudus | Bubble CPAP | Draeger/Seatle CPAP | QM2015174 | Maintenance | Preventive Maintenance - General Check | Ready to use |  | Amna
16000237 | 2024-05-22 | RS Dr. Suyoto | Dental Unit | Kavo/1058 | 10207193 | Survey | Survey alarm pada dental unit dan stuck logo kavo; Pengecekan mainboard | Pending | Perlu penggantian mainboard | Ipan
19000718 | 2024-05-22 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0044 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
14000317 | 2024-05-21 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 0720015 | Survey | Uji Kesesuaian | Ready to use |  | Purwono, Bahar, Satria
6000794 | 2024-05-21 | RSUD Prof. Dr. Margono Soekarjo | Jaudience Meter | Draeger/JM105 | B3503802 | Service | Alat selesai di kalibrasi | Ready to use |  | Afrizal
7000555 | 2024-05-21 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0038 | Survey | Vaporizer tidak mau terbuka; Karet vaporizer kendor | Ready to use | Disarankan penggantian service set 3 tahun | Wahyu
19000716 | 2024-05-21 | RS Indriati Solo | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0063 | Service | O2 sensor fail; Penggantian O2 sensor | Ready to use |  | Nanda
19000717 | 2024-05-21 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Draeger/Fabius Plus | ARZM0051 | Survey | O2 sensor fail; O2 sensor rusak | Pending |  | Nanda
19000717 | 2024-05-21 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0145 | Survey | O2 sensor fail; O2 sensor rusak | Pending |  | Nanda
19000717 | 2024-05-21 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Draeger/Fabius Plus | ASHD0007 | Survey | O2 sensor fail; O2 sensor rusak | Pending |  | Nanda
12000104 | 2024-05-21 | RSUD Kotamobagu | Digital Radiografi Fluoroskopi | Apelem/Platinum | D1860270 | Service | Komputer mati, psu bermasalah; Sementara menggunakan psu backup | Ready to use | Disarankan penggantian PSU komputer DRF | Marshall
6000792 | 2024-05-20 | RSUD Banyumas | Trolley Medicine | SYAS | 24056940200400018 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000792 | 2024-05-20 | RSUD Banyumas | Trolley Medicine | SYAS | 24056940200400017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000792 | 2024-05-20 | RSUD Banyumas | Trolley Medicine | SYAS | 24056940200400019 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
19000714 | 2024-05-20 | RS Ken Saras | Mesin Anasthesia | Draeger/Fabius GS Premium | ASBC0054 | Survey | O2 fail, baterai drop; O2 sensor dan baterai bermasalah | Pending |  | Nanda
19000713 | 2024-05-20 | RS Ken Saras | Mesin Anasthesia | Draeger/Fabius GS Premium | ASCL0014 | Service | Gas measurement fail dan error; o2 sensor expired, flow sensor empty | Pending | Perlu penggantian service set 3 tahun dan O2 sensor & flow sensor | Nanda
7000554 | 2024-05-20 | RSUD KOTA YOGYAKARTA | Jaudience Meter | Draeger/JM105 | B3503376 | Service | Unit dibawa ke draeger untuk dikalibrasi | Ready to use |  | Wahyu
19000715 | 2024-05-20 | RS Ken Saras | Mesin Anasthesia | Draeger/Fabius GS Premium | ASBC0055 | Service | O2 sensor fail, baterai fail; Baterai dan O2 sensor rusak | Pending | Perlu penggantian O2 sensor dan baterai | Nanda
7000554 | 2024-05-20 | RSUD KOTA YOGYAKARTA | Jaudience Meter | Draeger/JM105 | B3501236 | Service | Unit dibawa ke draeger untuk dikalibrasi | Ready to use |  | Wahyu
6000793 | 2024-05-18 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0013 | Survey | VT tidak sesuai; Deep cleaning dan maintenance | Ready to use |  | Afrizal
6000791 | 2024-05-18 | RS Bhakti Asih Brebes | Ventilator | Draeger/Savina300 | ASHD0011 | Service | Pemasangan service set 6 tahun dan baterai internal set | Ready to use |  | Afrizal
18000502 | 2024-05-17 | RSUD Ainun Habibie | Patient Monitor | Draeger/Infinity | ASRM0926 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000517 | 2024-05-17 | RSUD Tani dan Nelayan | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/018/C0 | Survey | Pengecekan unit | Ready to use | Alat dalam kondisi baik | Irfan R
16000235 | 2024-05-17 | Gigi Dental House | Dental Unit | Kavo/E50 Life | 30211888 | Service | Pemasangan rest function handiece | Ready to use |  | Ipan
9000454 | 2024-05-17 | RSUD Sanana | X-ray Digital | Apelem/Camargue DR | x | Survey | Survey ruangan untuk pra install | Ready to use |  | Ruslan
15000245 | 2024-05-16 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Survey | CPU PU Box rusak; CPU PU Box akan dibawa ke Jakarta | Pending |  | Sutan
15000244 | 2024-05-16 | RSUD Raden Mattaher | X-ray Analog | Apelem/Camargue Analog | IMA703863 | Service | Tensator spring rusak; Penggantian tensator spring 1 set | Pending | Kurang 1 set lagi | Sutan
9000441 | 2024-05-15 | RSUD Dr. H. Chasan Boesoirie | ESWL | Medispec/E3000 | 0922002 | Survey | coater electroda merosot dan muncul error; Penggantian electroda baru | Ready to use | Unit masih muncul error namun bisa digunakan | Ruslan
6000649 | 2024-05-15 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASEF0030 | Survey | Error FiO2, error display, error power supply, error alarm | Pending |  | Afrizal
10000185 | 2024-05-15 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | Touchscreen console tidak berfungsi dan software crash; Perbaikan connector console | Pending | Disarankan upgrade dan reload software | Heru
6000790 | 2024-05-15 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHL0172 | Survey | Error LCD, baterai dan filter | Pending |  | Afrizal
6000790 | 2024-05-15 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHL0173 | Survey | Error power supply, baterai dan filter | Ready to use | Disarankan penggantian maintenance kit | Afrizal
9000453 | 2024-05-15 | RSD Tidore | Baby Incubator | Draeger/Isolette 8000 | KY07192 | Survey | Unit mati total dan motor fan tidak bergerak; Perbaikan tegangan pada board power supply | Ready to use |  | Ruslan
9000453 | 2024-05-15 | RSD Tidore | Baby Incubator | Draeger/Isolette 8000 | KY07193 | Survey | Unit mati total dan motor fan tidak bergerak; Perbaikan tegangan pada board power supply | Ready to use |  | Ruslan
18000511 | 2024-05-14 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASHD0037 | Survey | Pengecekan unit; Valve MV2, ILCA2, Baterai dan Elektro valve rusak | Pending |  | Irfan R
18000511 | 2024-05-14 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASED0171 | Survey | Pengecekan unit; Valve MV2, ILCA2, Baterai dan Elektro valve rusak | Pending |  | Irfan R
7000551 | 2024-05-14 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASLL0067 | Maintenance | Maintenance total care | Ready to use |  | Wahyu
7000552 | 2024-05-14 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASNL0107 | Maintenance | Maintenance total care | Ready to use |  | Wahyu
7000553 | 2024-05-14 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASJN0023 | Maintenance | Maintenance total care | Ready to use |  | Wahyu
19000712 | 2024-05-14 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0074 | Service | Penggantian blower | Ready to use |  | Nanda
6000789 | 2024-05-13 | RS Dadi Keluarga | Ventilator | Draeger/Babylog VN800 | ASPC0031 | Service | Pemasangan backup unit | Ready to use |  | Afrizal
6000789 | 2024-05-13 | RS Dadi Keluarga | Baby Incubator | Draeger/Isolette 8000 | DV38326 | Service | Pemasangan backup unit | Ready to use |  | Afrizal
19000711 | 2024-05-13 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0180 | Service | N2O low; Tarik error - setting mixer | Ready to use |  | Nanda
15000246 | 2024-05-13 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Service | Penggantian EMH | Ready to use |  | Sutan
8000553 | 2024-05-13 | RS ISLAM PATI | Lampu Operasi | Draeger/Polaris 600 | ASEK0100 | Survey | Lampu mati ketika digerakan; Cek koneksi lampu | Pending |  | Amna
3000305 | 2024-05-13 | RSUD Sumedang | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/132/C2 | Service | system star up not responding; Pengecekan pcb 190 pada emergency push button | Ready to use |  | Imam
18000504 | 2024-05-07 | RSUD Otohana | Ventilator | Draeger/Evita V800 | ASPL0095 | Service | Pressure regulator rusak, ada karat dan sudah dibersihkan | Ready to use | Unit sementara bisa digunakan namun disarankan penggantian pressure regulator | Irfan R
18000505 | 2024-05-07 | RSUD Toto Kabila | C-arm | Apelem/Evo R+ | 55121004853 | Service | Unit tiba tiba mati; Supply board hangus | Pending |  | Irfan R
6000709 | 2024-05-06 | RS Pertamina Cilacap | Central Monitor | Draeger/Vista 120 CMS | 1921684120 | Service | Re-install software | Ready to use |  | Afrizal
18000506 | 2024-05-06 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Fabius | ASKK0057 | Service | Kebocoran O2, selang lepas. Sudah dipasang kembali | Ready to use |  | Irfan R
2000178 | 2024-05-06 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Error table; Adjust table position | Ready to use |  | Eddy
19000710 | 2024-05-06 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Survey | Hasil gambar usg saat standby bergaris; Dicoba menggunakan coin masih bergaris | Pending |  | Nanda
6000798 | 2024-05-04 | RSUD dr. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | AH36596 | Service | Pemasangan engsel | Ready to use |  | Afrizal
6000798 | 2024-05-04 | RSUD dr. Soedirman - Kebumen | Incubator Transport | Draeger/TI500 | BH12346 | Service | Pemasangan cable asy | Ready to use |  | Afrizal
6000788 | 2024-05-03 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ0080 | Maintenance | Pengecekan software dan hardware | Ready to use |  | Afrizal
6000788 | 2024-05-03 | RSI Harapan Anda | Incubator Transport | Draeger/TI500 | HC13259 | Survey | Unit tidak mau charging, ada baut yang lepas. Unit sedang dilakukan pemantauan | Ready to use |  | Afrizal
6000787 | 2024-05-03 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0009 | Survey | O2 sensor failure | Ready to use | Disarankan penggantian O2 sensor dan flow sensor | Afrizal
6000787 | 2024-05-03 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASEK0064 | Survey | O2 sensor failure | Ready to use | Disarankan penggantian O2 sensor dan service set | Afrizal
6000787 | 2024-05-03 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASEK0065 | Survey | O2 sensor failure, ada kebocoran | Ready to use | Disarankan penggantian O2 sensor dan service set | Afrizal
6000787 | 2024-05-03 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0164 | Survey | Adanya kebocoran | Ready to use | Disarankan penggantian O2 sensor dan waterlock | Afrizal
6000787 | 2024-05-03 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ARZF0073 | Survey | O2 sensor failure | Ready to use | Disarankan penggantian O2 sensor dan waterlock | Afrizal
6000786 | 2024-05-02 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHK0151 | Survey | Sensor check merah; Baterai drop | Pending | Perlu dilakukan penggantian 2/2 port distribution valve dan primus set 6 tahun | Afrizal
6000786 | 2024-05-02 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASNE0151 | Survey | O2 sensor failure | Ready to use | Alat masih bisa digunakan namun pembacaan O2 tidak bisa | Afrizal
8000551 | 2024-05-02 | RSUD Kajen | Mesin Anasthesia | Draeger/Primus | ASLD0023 | Maintenance | Preventive Maintenance - General Check | Ready to use |  | Amna
8000551 | 2024-05-02 | RSUD Kajen | Mesin Anasthesia | Draeger/Primus | ASMA0191 | Maintenance | Preventive Maintenance - General Check | Ready to use |  | Amna
17000183 | 2024-05-02 | RSUD Malingping | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/131/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Bahar
10000184 | 2024-05-02 | RSUD Ulin | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150176 | Survey | Pendampingan penggunaan alat ke pasien | Ready to use |  | Heru
16000234 | 2024-04-30 | Dental Works | Dental Unit | Kavo/1058TM | 1003977 | Survey | Rembes air dari keran gelas dan bowl, air three way tidak keluar; Pengecekan selang pembuangan bowl dan pengecekan saluran air, saluran pembuangan mampet | Pending | Perlu penggantian selang pembuangan dan air three way | Ipan
19000709 | 2024-04-30 | RSUD Ibu Fatmawati Surakarta | Infant Warmer | Draeger/Resuscitaire | DSNL0048 | Maintenance | Preventive Maintenance | Ready to use | Ada engsel patah | Nanda
19000709 | 2024-04-30 | RSUD Ibu Fatmawati Surakarta | Infant Warmer | Draeger/Resuscitaire | DSNL0046 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000709 | 2024-04-30 | RSUD Ibu Fatmawati Surakarta | Infant Warmer | Draeger/Resuscitaire | DSNL0045 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000709 | 2024-04-30 | RSUD Ibu Fatmawati Surakarta | Infant Warmer | Draeger/Resuscitaire | DSNL0047 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
7000549 | 2024-04-30 | RSUP dr. Soeradji Tirtonegoro | Central Monitor | Draeger/Vista 120 CMS | CMSPC0056 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
18000512 | 2024-04-30 | RSUP Prof. Dr. R. D. Kandou Manado | Pendant | Draeger/Movita | ASHF0154 | Survey | Kerusakan pada board valve | Pending |  | Irfan R
18000512 | 2024-04-30 | RSUP Prof. Dr. R. D. Kandou Manado | Pendant | Draeger/Movita | ASHF0153 | Survey | Kerusakan pada board valve | Pending |  | Irfan R
18000512 | 2024-04-30 | RSUP Prof. Dr. R. D. Kandou Manado | Pendant | Draeger/Movita | ASJD0327 | Survey | Kerusakan pada board valve | Pending |  | Irfan R
18000512 | 2024-04-30 | RSUP Prof. Dr. R. D. Kandou Manado | Pendant | Draeger/Movita | ASJD0324 | Survey | Kerusakan pada board valve | Pending |  | Irfan R
18000521 | 2024-04-30 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0415 | Survey | Kerusakan pada nebulizer | Pending |  | Irfan R
18000520 | 2024-04-30 | RSUP Prof. Dr. R. D. Kandou Manado | Monitor Transport | Draeger/Infinity M540 | 5614876955 | Survey | Baterai dan aksesoris bermasalah | Pending |  | Irfan R
6000785 | 2024-04-29 | RS Dadi Keluarga | Ventilator | Draeger/Babylog VN800 | ASPF0030 | Survey | Display backlight mati, coba tukar dengan M48 dan reload software | Pending |  | Afrizal
16000233 | 2024-04-29 | Gigi Dental House | Dental Unit | Kavo/E50 | 30211888 | Service | Pemasangan monitor, kabel power dan geser breket monitor | Ready to use |  | Ipan, Bilal
10000183 | 2024-04-29 | RSUD Ulin | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150176 | Maintenance | Preventive Maintenance | Ready to use |  | Heru
20000270 | 2024-04-29 | RSUP Tadjudin Chalid | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150292 | Survey | Fiber optic rusak | Pending | Perlu dilakukan penggantian fiber optic | Sujana
3000304 | 2024-04-29 | RSUD Tarakan | C-arm | Apelem/Evo R+ | 55118030653 | Survey | Window blue screen, operating system window error); Pengecekan pc workstation c-arm | Pending | Akan diberikan penawaran upgrade SSD include software dan DDR 4 | Imam
6000780 | 2024-04-27 | RSUD Ajibarang | Patient Monitor | Draeger/Vista 120 | VGSNH1336 | Service | Pemasangan main pcb (claim warranty) | Ready to use |  | Afrizal
6000780 | 2024-04-27 | RSUD Ajibarang | Ventilator | Draeger/Babylog VN500 | ASMK0054 | Survey | Flow sensor failure, cable sensor rusak. Sementara di backup dengan kabel yang ada | Ready to use |  | Afrizal
9000427 | 2024-04-25 | RSUD Angkatan Laut Bitung | Mesin Anasthesia | Draeger/Fabius GS Premium | ASMH0171 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian O2 sensor | Irfan R
9000427 | 2024-04-25 | RSUD Angkatan Laut Bitung | Mesin Anasthesia | Draeger/Fabius GS Premium | ASMH0170 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian O2 sensor | Irfan R
9000431 | 2024-04-25 | RSUD Angkatan Laut Bitung | Ventilator | Draeger/Savina300 | ASMH0028 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian baterai internal | Irfan R
9000431 | 2024-04-25 | RSUD Angkatan Laut Bitung | Ventilator | Draeger/Savina300 | ASMH0027 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian baterai internal | Irfan R
9000431 | 2024-04-25 | RSUD Angkatan Laut Bitung | Ventilator | Draeger/Savina300 | ASMH0026 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian baterai internal | Irfan R
9000431 | 2024-04-25 | RSUD Angkatan Laut Bitung | Ventilator | Draeger/Savina300 | ASMH0023 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian baterai internal | Irfan R
9000431 | 2024-04-25 | RSUD Angkatan Laut Bitung | Ventilator | Draeger/Savina300 | ASMH0025 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian baterai internal | Irfan R
9000431 | 2024-04-25 | RSUD Angkatan Laut Bitung | Ventilator | Draeger/Savina300 | ASMH0029 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian baterai internal | Irfan R
6000784 | 2024-04-24 | RS Dadi Keluarga | Ventilator | Draeger/Babylog VN800 | ASPF0030 | Survey | Backlight monitor mati; Inverter lcd rusak | Pending |  | Afrizal
17000182 | 2024-04-24 | RSUD Balaraja | C-arm | Apelem/Evo R+ | 55118034657 | Survey | Pergerakan column tidak bisa stop | Pending | Perlu penggantian stop brake | Bahar
4000431 | 2024-04-24 | RSUD Provinsi NTB | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/032/C0 | Service | Tidak bisa melakukan kompresi; Penggantian board 215 dan kabel CV0558 | Pending | Unit belum bisa kompresi, disarankan pengecekan board 240 pada unit lain | Satria
12000101 | 2024-04-23 | RSUD Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Survey | Console stuck di initiation; Penggantian processor komputer | Ready to use |  | Marshall
6000783 | 2024-04-22 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASRK0231 | Survey | Error 10. Baterai drop, software error | Pending | Perlu dilakukan penggantian baterai internal, baterai eksternal, filter air, filter O2, filter electronic, upgrade software | Afrizal
6000783 | 2024-04-22 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASRK0410 | Survey | Error 10. Baterai drop, software error | Pending | Perlu dilakukan penggantian baterai internal, baterai eksternal, filter air, filter O2, filter electronic, upgrade software | Afrizal
16000232 | 2024-04-22 | Gigi Dental House | Dental Unit | Kavo/E50 | 30211888 | Instalasi | Alat berfungsi dengan baik | Ready to use | Kekurangan rst function | Ipan
17000181 | 2024-04-22 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Service | Pemindahan unit printer ke ruang user | Ready to use |  | Bahar
19000707 | 2024-04-22 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0087 | Service | Breathing circuit bocor; Pengecekan dan pembersihan breathing circuit | Ready to use |  | Nanda
18000508 | 2024-04-21 | RSUD Samratulangi Tondano | Ventilator | Draeger/Savina300 | ASJM0392 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian baterai internal | Irfan R
18000508 | 2024-04-21 | RSUD Samratulangi Tondano | Ventilator | Draeger/Savina300 | ASNL0255 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian baterai internal | Irfan R
18000508 | 2024-04-21 | RSUD Samratulangi Tondano | Ventilator | Draeger/Savina300 | ASNL0256 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian baterai internal | Irfan R
8000350 | 2024-04-19 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0148 | Service | Pemasangan maintenance kit, O2 sensor dan diapragma piston | Ready to use |  | Amna
8000350 | 2024-04-19 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0115 | Service | Pemasangan maintenance kit, O2 sensor dan diapragma piston | Ready to use |  | Amna
8000350 | 2024-04-19 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0049 | Service | Pemasangan maintenance kit, O2 sensor dan diapragma piston | Ready to use |  | Amna
8000350 | 2024-04-19 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASSC0025 | Service | Pemasangan maintenance kit, O2 sensor dan diapragma piston | Ready to use |  | Amna
17000180 | 2024-04-19 | RSUD Dr. Drajat Prawiranegara | X-ray Digital | Apelem/Camargue DR | 2510220562 | Maintenance | Preventive Maintenance | Ready to use |  | Bahar
11000310 | 2024-04-19 | RSUP Dr. Wahidin Sudirohusodo | ESST | Medispec/ED1000 | 09111375 | Survey | Touchscreen tidak berfungsi; Kabel fleksible rusak | Pending | Disarnakan penggantian kabel fleksible | Nasrul
3000302 | 2024-04-19 | RSUD Kab. Temanggung | Mobile X-ray | Apelem/Rafale EV30 | 009/19/101306 | Survey | Filament fault, inverter fault; Instalasi listrik gedung bermasalah | Ready to use |  | Imam
3000303 | 2024-04-19 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Service | generator clock error; Pengeeckan cmos cpu generator | Ready to use |  | Imam
18000344 | 2024-04-18 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0004 | Survey | Error optimix pressure sensor, baterai, differensial sensor | Pending |  | Irfan R
19000706 | 2024-04-18 | RSIA Umi Barokah Boyolali | Baby Incubator | Draeger/Isolette C2000 | DSSK0008 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
17000179 | 2024-04-18 | RSUD Tarakan | C-arm | Apelem/Evo R+ | 551118030653 | Survey | Unit tidak bisa menyala; Pengecekan cpu dan cpu dibawa untuk dicek lebih lanjut | Pending |  | Bahar
6000779 | 2024-04-17 | RSI Harapan Anda | Ventilator | Draeger/Babylog VN500 | ASMK0055 | Service | Pemasangan part PATO sensor dan pressure regulator | Ready to use |  | Afrizal
7000547 | 2024-04-17 | RS Universitas Islam Indonesia | Ventilator | Draeger/Babylog VN600 | ASTB0025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
8000555 | 2024-04-15 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASMD0040 | Survey | Backlight monitor mati; LCD rusak | Pending | Perlu penggantian LCD M.14.5 | Amna
6000778 | 2024-04-05 | RS Dadi Keluarga | Ventilator | Draeger/Babylog VN800 | ASPF0080 | Survey | FiO2 measurement failed; Self test, OKE | Ready to use |  | Afrizal
4000430 | 2024-04-04 | RSPPN Soedirman | Dental Unit | Kavo/1058 | 10207193 | Service | Bowl tidak keluar air; Penggantian selenoid valve | Ready to use |  | Satria
9000429 | 2024-04-03 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASHJ0179 | Survey | Power supply, baterai dan mainboard rusak | Pending |  | Ruslan
9000429 | 2024-04-03 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASHJ0181 | Survey | Power supply, baterai dan mainboard rusak | Pending |  | Ruslan
9000429 | 2024-04-03 | RSD Tidore | Ventilator | Draeger/Savina 300 | ASHJ0178 | Survey | Power supply, baterai dan mainboard rusak | Pending |  | Ruslan
8000349 | 2024-04-03 | RSUD RAA Soewondo Pati | Mayo Stand | Syas/Mayo Stand | 240335113000400092 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
7000546 | 2024-04-03 | RSUD Bagas Waras | Ventilator | Draeger/Babylog VN500 | ASJD0020 | Survey | Software failure; Baterai lemah | Pending | Disarankan penggantian baterai | Wahyu
6000777 | 2024-04-03 | RSUD Dr. Soeselo Slawi | Ventilator | Draeger/Babylog VN500 | ASJJ0104 | Survey | Touchscreen kadang bermasalah; Kalibrasi internal | Ready to use | Disarankan penggantian touchscreen M145, baterai pack M7 | Afrizal
6000776 | 2024-04-02 | RSUD Dr. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEL0081 | Survey | gas supply sensor calibration merah | Pending |  | Afrizal
19000705 | 2024-04-02 | RSUD dr. Gunawan Mangunkusumo | Incubator Transport | Draeger/Isolette TI500 | KY11400 | Survey | Baterai tidak mengisi; Ada air keluar dari baterai, exchange dengan baterai unit lain | Ready to use |  | Nanda
4000429 | 2024-04-02 | RSPPN Soedirman | Dental Unit | Kavo/1058 | 10207193 | Survey | Air bowl tidak keluar; Pengecekan dan pengukuran tegangan pada valve | Pending | Perlu penggantian selenoid valve | Satria
11000308 | 2024-04-01 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASAM0045 | Survey | Ventilator fail, battery fail; Baterai internal dan breathing system bermasalah | Pending | Disarankan penggantian service set 3 tahun, O2 sensor, housing O2 sensor, PBA Control 3, PBA lucy 2 dan elbow fitting | Nasrul
11000308 | 2024-04-01 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASAM0044 | Survey | Ventilator fail, battery fail; Baterai internal dan breathing system bermasalah | Pending | Disarankan penggantian service set 3 tahun, O2 sensor, housing O2 sensor, PBA Control 3, PBA lucy 2 dan elbow fitting | Nasrul
11000308 | 2024-04-01 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0115 | Survey | Ventilator fail, battery fail; Baterai internal dan breathing system bermasalah | Pending | Disarankan penggantian service set 3 tahun, O2 sensor, housing O2 sensor, PBA Control 3, PBA lucy 2 dan elbow fitting | Nasrul
6000775 | 2024-03-28 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASNB0071 | Service | Pemasangan waterlock | Ready to use |  | Afrizal
6000775 | 2024-03-28 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASJF0034 | Service | Pemasangan waterlock | Ready to use |  | Afrizal
6000775 | 2024-03-28 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0039 | Service | Pemasangan O2 sensor | Ready to use |  | Afrizal
6000775 | 2024-03-28 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASMH0115 | Service | Pemasangan waterlock | Ready to use |  | Afrizal
18000510 | 2024-03-28 | RSUP Prof. Dr. R. D. Kandou Manado | Infant Warmer | Draeger/Resuscitaire | CY11217 | Survey | Heater rusak | Pending |  | Irfan R
7000545 | 2024-03-27 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Maintenance | Preventive Maintenance Total Care | Ready to use |  | Wahyu
18000518 | 2024-03-26 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | CU03934 | Survey | Touchscreen di layar bermasalah, Control assy rusak | Pending |  | Irfan R
18000518 | 2024-03-26 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | KM04800 | Survey | Touchscreen di layar bermasalah, Control assy rusak | Pending |  | Irfan R
9000425 | 2024-03-26 | RSUD Weda | Dental Unit | Kavo/1058 TM | 10205459 | Survey | Foot control dan display tidak berfungsi; Dentist element board rusak | Pending | Disarankan penggantian board dentist element | Ruslan
18000516 | 2024-03-25 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | BY33330 | Survey | Alat tidak bisa menyala, switch dan power entry rusak | Pending |  | Irfan R
8000348 | 2024-03-25 | RSI Sultan Hadlirin | Mesin Anasthesia | Draeger/Fabius Plus | ASND0059 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000348 | 2024-03-25 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 | ASMN0162 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
18000503 | 2024-03-25 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | ASJJ0105 | Survey | Kerusakan pada O2 pato | Pending |  | Irfan R
18000509 | 2024-03-25 | RSUP Prof. Dr. R. D. Kandou Manado | Incubator Transport | Draeger/TI500 | JV13717 | Survey | Kerusakan pada board supply | Pending |  | Irfan R
18000509 | 2024-03-25 | RSUP Prof. Dr. R. D. Kandou Manado | Incubator Transport | Draeger/TI500 | DSNL0022 | Survey | Kerusakan pada board supply | Pending |  | Irfan R
3000301 | 2024-03-25 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Service | Setting PACS dan Worklist | Ready to use |  | Imam
7000544 | 2024-03-22 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Maintenance | Preventive Maintenance Total Care | Ready to use |  | Wahyu
13000170 | 2024-03-21 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Service | Tembakan tidak kuat; Penggantian EMH | Ready to use |  | Sofian
4000427 | 2024-03-21 | RSUD Kab. Temanggung | ESWL | Medispec/EM1000 | 1115056 | Survey | Pengecekan fungsi unit C-arm ESWL | Ready to use | Disarankan penambahan RAM | Satria
4000428 | 2024-03-21 | RSUD Kab. Temanggung | C-arm | Apelem/Evo R+ | 55115011409 | Survey | Pengecekan fungsi unit C-arm | Ready to use | Disarankan upgrade SSD | Satria
4000425 | 2024-03-21 | RSUD Kab. Temanggung | Mobile X-ray | Apelem/Rafale EV30 | 009/19/01306 | Survey | Error filament fault; Plug and play board S22 | Ready to use |  | Satria
4000426 | 2024-03-21 | RSUD Kab. Temanggung | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/011/C0 | Survey | Pengecekan fungsi unit | Ready to use |  | Satria
8000347 | 2024-03-21 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Babylog VN600 | ASSK0028 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000347 | 2024-03-21 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina300 | ASPJ0383 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000347 | 2024-03-21 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina300 | ASPJ0382 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000347 | 2024-03-21 | RS PKU Muhammadiyah Mayong | Mesin Anasthesia | Draeger/Fabius | ASPH0032 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000346 | 2024-03-20 | RSUD Limpung | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0146 | Maintenance | Preventive Maintenance | Ready to use | Perlu penggantian O2 sensir dan service set 3 thn | Amna
18000514 | 2024-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0173 | Survey | Kerusakan pada main board, power supply board, pressure regulator | Pending |  | Irfan R
18000513 | 2024-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0181 | Survey | Kerusakan pada switch on off | Pending |  | Irfan R
18000513 | 2024-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0183 | Survey | Kerusakan pada switch on off | Pending |  | Irfan R
3000300 | 2024-03-20 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | Software 3D error; Reload software | Ready to use |  | Imam
18000515 | 2024-03-18 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0165 | Survey | Kerusakan pada pressure regulator, baterai dan switch on-off | Pending |  | Irfan R
18000515 | 2024-03-18 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0167 | Survey | Kerusakan pada pressure regulator, baterai dan switch on-off | Pending |  | Irfan R
9000417 | 2024-03-18 | RSUD Sanana | Dental Unit | Kavo/E30TM | 10111150 | Survey | Scaller tidak berfungsi; Pengecekan dengan board scaller | Ready to use |  | Ruslan
8000345 | 2024-03-15 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASPH0121 | Maintenance | Preventive Maintenance | Ready to use | Perlu penggantian service set 6thn, flow sensor, water lock dan sample line | Amna
8000345 | 2024-03-15 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Maintenance | Preventive Maintenance | Ready to use | Perlu penggantian service set 6thn, flow sensor, water lock dan sample line | Amna
8000345 | 2024-03-15 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Maintenance | Preventive Maintenance | Ready to use | Perlu penggantian service set 6thn, flow sensor, water lock dan sample line | Amna
8000345 | 2024-03-15 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Perseus | ASJF0033 | Maintenance | Preventive Maintenance | Ready to use | Perlu penggantian service set 6thn, flow sensor, water lock dan sample line | Amna
14000304 | 2024-03-14 | RSUD Sidoarjo Barat | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0755 | Service | Table xray tidak bisa turun; Penggantian resistor Break | Ready to use |  | Purwono
13000168 | 2024-03-13 | RSUD Dr. Djasamen Saragih | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/137/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Sofian
19000704 | 2024-03-13 | RSUP dr. Kariadi | Infant Warmer | Draeger/Resuscitaire | MC13338 | Service | Error 05, heater rusak | Pending | Perlu penggantian heater | Nanda
3000299 | 2024-03-13 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Survey | Generator error; Ada indikasi board IDB bermasalah | Pending | Board IDB dan rectifier board belum bisa dibawa untuk di cek lebih lanjut | Imam
19000703 | 2024-03-08 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Service | Air mompa terus sehingga balon mengembang tidak normal;Kalibrasi spark gap | Ready to use |  | Nanda
100002 | 2024-03-08 | RSUP Dr. M. Djamil Padang | Bone Mineral Densitometer | Medilink/Medix 90 | J13015M270 | Service | Scan hanya sampai 40% | Pending | Perlu penggantian monoblock, inverter dan flexible cable | Syamsudin
8000344 | 2024-03-08 | RS Fastabiq Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASRC0070 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Amna
7000543 | 2024-03-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0026 | Survey | Gas supply fail; Inlet gas supply rs bermasalah, sudah diperbaiki | Ready to use |  | Wahyu
14000306 | 2024-03-08 | RSUD Kab. Kediri | Emergency trolly | SYAS/SCE2-SS | 240169501000400018 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000305 | 2024-03-08 | RSUD Kab. Kediri | Medicine Trolley | SYAS/SCM2-40 | 240369402000400003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000302 | 2024-03-08 | RSUD Kab. Kediri | Instrument Trolley | SYAS/5C1 | 240169101000400106 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000303 | 2024-03-08 | RSUD Kab. Kediri | Dressing trolley | SYASY/SCE1 | 24016920100400054 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000303 | 2024-03-08 | RSUD Kab. Kediri | Dressing trolley | SYASY/SCE1 | 24016920100400055 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000303 | 2024-03-08 | RSUD Kab. Kediri | Dressing trolley | SYASY/SCE1 | 24016920100400056 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000303 | 2024-03-08 | RSUD Kab. Kediri | Dressing trolley | SYASY/SCE1 | 24016920100400057 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000303 | 2024-03-08 | RSUD Kab. Kediri | Dressing trolley | SYASY/SCE1 | 24016920100400058 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000303 | 2024-03-08 | RSUD Kab. Kediri | Dressing trolley | SYASY/SCE1 | 24016920100400059 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
14000303 | 2024-03-08 | RSUD Kab. Kediri | Dressing trolley | SYASY/SCE1 | 24016920100400060 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono
3000298 | 2024-03-08 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Imam
7000542 | 2024-03-07 | RSUD Wonosari | ESWL | Medispec/E3000 | 0719139 | Maintenance | Preventive Maintenance | Ready to use |  | Wahyu
13000167 | 2024-03-07 | RSUD Dr. Fauziah | ESWL | Medispec/EM1000 | 0716249 | Service | Monitor C-arm mati, tembakan EMH berkurang tenaganya; Pengecekan konektor dan RAM | Ready to use | Disarankan ganti RAM monitor C-arm dan ganti EMH | Sofian
100003 | 2024-03-07 | RS Otak Moh. Hatta | ESWL | Medispec/EM1000 | 0922163 | Instalasi | Pemasangan USG | Ready to use |  | Syamsudin
3000297 | 2024-03-07 | RSUD Sumedang | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/132/C2 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Imam
6000650 | 2024-03-06 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08779
6000650 | 2024-03-06 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08776
6000650 | 2024-03-06 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | DSNF0062
6000650 | 2024-03-06 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | AF32541
6000650 | 2024-03-06 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08718
6000647 | 2024-03-06 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Tiro | ASBM0248 | Survey | Fresh gas low alarm; Selft test dan kalibrasi pressure. | Ready to use | Disarankan penggantian service set 3 tahun dan O2 sensor | Afrizal
6000647 | 2024-03-06 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0027 | Survey | Pengecekan berkala; O2 sensor calibration failed, O2 sensor expired | Ready to use | Disarankan penggantian O2 sensor | Afrizal
6000647 | 2024-03-06 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0039 | Survey | Pengecekan berkala; O2 sensor expired | Ready to use | Disarankan penggantian O2 sensor | Afrizal
7000541 | 2024-03-05 | RSPAU dr. S. Hardjolukito | Infant Warmer | Draeger/Babytherm 8004 | ASDL0042 | Survey | Error 67; Panel keypad rusak | Pending |  | Wahyu
8000343 | 2024-03-05 | RS PKU Aisyah Jepar | Mesin Anasthesia | Draeger/Fabius Plus | ASND0058 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
16000231 | 2024-03-05 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Service | Penggantian selang scaller | Ready to use |  | Ipan
7000538 | 2024-03-01 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Maintenance | Preventive Maintenance Total care | Ready to use |  | Wahyu
7000539 | 2024-03-01 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Maintenance | Preventive Maintenance Total care | Ready to use |  | Wahyu
7000540 | 2024-02-28 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Survey | Unit restart; Board console unit bermasalah | Pending |  | Wahyu
9000415 | 2024-02-28 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | J166F0184 | Service | console controller tidak hidup otomatis dan switch contras tidak berfungsi | Pending | Disarankan penggantian ups dan baterai cmos | Ruslan
7000536 | 2024-02-27 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Maintenance | Preventive Maintenance Total care | Ready to use |  | Wahyu
7000537 | 2024-02-27 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0118 | Maintenance | Preventive Maintenance Total care | Ready to use |  | Wahyu
11000307 | 2024-02-27 | RSUD Hajjah Andi | Mesin Anasthesia | Draeger/Perseus A500 | ASNJ0104 | Survey | Gas sensor failure, O2 sensor failure, ventilator failed; Cek gas sensor, hit software, reload software. Baterai internal dan flow power rusak | Pending | Disarankan penggantian flow power, baterai internal dan kontrak service | Nasrul
9000424 | 2024-02-27 | RSUD Kotamobagu | Ventilator | Draeger/Savina 300 | ASNL0998 | Survey | Baterai internal fail | Pending | Perlu penggantian baterai internal | Ruslan
9000424 | 2024-02-27 | RSUD Kotamobagu | Ventilator | Draeger/Savina 300 | ASNL0999 | Survey | Baterai internal fail | Pending | Perlu penggantian baterai internal | Ruslan
3000296 | 2024-02-27 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Survey | Komunikasi detector lost; ups console dan pu box rusak, ada masalah pada software workstation | Pending |  | Imam
16000230 | 2024-02-24 | Gigi Dental House | Dental Unit | Kavo/1058 | 2019-09-10207196 | Survey | Error water block press dan leakage water; Pengecekan jalur air dan angin dan water tank | Ready to use |  | Ipan
7000535 | 2024-02-23 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0007 | Maintenance | Maintenance total care | Ready to use |  | Wahyu
7000534 | 2024-02-23 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Maintenance | Maintenance total care | Ready to use |  | Wahyu
7000533 | 2024-02-23 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Maintenance | Maintenance total care | Ready to use |  | Wahyu
6000773 | 2024-02-21 | RSI Harapan Anda | Incubator Transport | Draeger/TI500 | HH12697 | Service | Pemasangan battery | Ready to use |  | Afrizal
16000228 | 2024-02-21 | RSGM Unsyiah | Dental Unit | Kavo/1058 | 2019-04-10206378 | Survey | Dental unit bunyi alarm dan kursi tidak berfungsi | Pending | Perlu penggantian mainboard | Ipan
16000228 | 2024-02-21 | RSGM Unsyiah | Dental Unit | Kavo/1058 | 2019-04-10206378 | Survey | Dental unit bunyi alarm dan kursi tidak berfungsi | Pending | Perlu penggantian mainboard | Ipan
16000229 | 2024-02-21 | RSGM Unsyiah | Dental Unit | Kavo/1058 | 2019-04-10206377 | Survey | Pengecekan dental unit | Ready to use | Unit normal tanpa pengecekan dengan air | Ipan
16000229 | 2024-02-21 | RSGM Unsyiah | Dental Unit | Kavo/1058 | 2019-04-10206379 | Survey | Pengecekan dental unit | Ready to use | Unit normal tanpa pengecekan dengan air | Ipan
1000001 | 2024-02-20 | Gigi Dental House | Dental Unit | Kavo/1058 | 2019-09-10207196 | Service | Support arm menabrak tembok; Pemasangan stopper pada arm | Ready to use |  | Ipan
19000701 | 2024-02-19 | RSUP dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0303 | Service | Flow sensor fail; Kerusakan pada kabel flow sensor | Pending | Disarankan penggantian kabel flow sensor | Nanda
16000227 | 2024-02-19 | Gigi Dental House | Dental Unit | Kavo/1058 | 10207194 | Service | Sistem motor tidak berfungsi | Ready to use |  | Ipan
16000227 | 2024-02-19 | Gigi Dental House | Dental Unit | Kavo/1058 | 10207196 | Service | Pengecekan suction | Ready to use |  | Ipan
7000530 | 2024-02-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Maintenance | Maintenance total care | Ready to use |  | Wahyu
7000529 | 2024-02-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Maintenance | Maintenance total care | Ready to use |  | Wahyu
7000531 | 2024-02-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHL0059 | Maintenance | Maintenance total care | Ready to use |  | Wahyu
7000527 | 2024-02-16 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina 300 Advance | ASTA0022 | Instalasi | Alat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
7000527 | 2024-02-16 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina 300 | ASSL0091 | Instalasi | Alat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
7000527 | 2024-02-16 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina 300 | ASSL0093 | Instalasi | Alat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
13000164 | 2024-02-15 | RSU Haji Medan | Dental Unit | Kavo/1058TM | 20008319 | Service | Air tidak masuk ke unit. Waterblock press; Membersihkan filter | Ready to use | Disarankan penggantian filter | Sofian
6000772 | 2024-02-15 | RSKBD Panti Nugroho | Infant Warmer | Draeger/Resuscitaire | BV14104 | Service | Pemasangan part heater assy | Ready to use |  | Afrizal
8000342 | 2024-02-14 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0093 | Service | Battery fault; Penggantian baterai | Ready to use |  | Amna
19000879 | 2024-02-13 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C2000 | CY34224 | Survey | Pengunci engsel patah | Pending |  | Nanda
8000341 | 2024-02-13 | RSUD Kab. Temanggung | Bone Mineral Densitometer | Apelem/Medix DR | H237H1094 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Amna
6000771 | 2024-02-13 | RSUD Brebes | Ventilator | Draeger/Evita V300 | JH0124 | Service | Pemasangan flow sensor, cable spirolog flow sensor, safety valve, expiratory valve | Ready to use |  | Afrizal
14000301 | 2024-02-12 | RSUD Sidoarjo Barat | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0755 | Service | Error ET117; Resistor break rusak | Pending |  | Purwono
6000769 | 2024-02-12 | RSI Harapan Anda | Ventilator | Draeger/Babylog VN500 | ASMK0055 | Survey | Air supply [utih, O2 supply putih, O2 sensor kalibrasi merah; Pengecekan software dan hardware, terdapat banyak kotoran serbuk di pnematic | Pending | Dianjurkan penggantian PATO O2 sensor dan pressure regulator | Afrizal
19000698 | 2024-02-12 | RSUD Dr. Soediran Mangun Sumarso | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/134/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
11000306 | 2024-02-06 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN300 | ASND0141 | Survey | Gas sensor cal. Fail; Service CasAir test | Pending | Dianjurkan kit proportional valve air faulty dan Valve O2 | Nasrul
8000340 | 2024-02-06 | RS Panti Wilasa | Ventilator | Draeger/Savina 300 | ASMM0026 | Service | Membran valve hilang; Penggantian service set 2 tahun | Ready to use |  | Amna
16000226 | 2024-02-06 | Signature Dental Care | Dental Unit | Planmeca/Compact i5 | UPCV574532 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
10000182 | 2024-02-06 | RSUD Ulin | Bone Mineral Densitometer | Medilink/Medix 90 | J15016M269 | Survey | Generator fault; Perlu penggantian monoblock, inverter dan detector | Pending |  | Heru
8000339 | 2024-02-02 | RSI Aisyah Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0048 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000339 | 2024-02-02 | RSI Aisyah Kudus | Ventilator | Draeger/Babylog VN500 | ASPF0017 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000339 | 2024-02-02 | RSI Aisyah Kudus | CPAP | Draeger/Seatle CPAP | ASKE0070 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
6000770 | 2024-02-02 | RS Pertamina Cilacap | Ventilator | Draeger/Savina 300 | ASNN0613 | Survey | Pengukuran tekanan tidak akurat, merah pada katup ekspirasi; Cek software dan hardware, pressure sensor ekspirasi tidak terbaca | Ready to use |  | Afrizal
9000423 | 2024-02-02 | RSUD Sultan H. Mudaffar Sjah | Mobile X-ray | Apelem/Rafale | 1872342200002 | Service | Unit tidak menghasilkan X-ray; Perbaikan konektor dan kabel tombol expose, penggantian cairan fikser dan developer pada APE | Ready to use |  | Ruslan
19000692 | 2024-02-01 | RSUD Dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0149 | Service | Layar tidak bisa disentuk dan batterai check; Kalibrasi layar | Ready to use |  | Nanda
6000751 | 2024-02-01 | RSUD dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Atlan A300 | ASRK0304 | Survey | O2 measurement inop; Pengecekan software dan hardware | Ready to use |  | Afrizal
10000181 | 2024-02-01 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Perbaikan kabel LAN | Ready to use |  | Heru
7000523 | 2024-01-31 | RSA UGM | Mesin Anasthesia | Draeger/Fabius Plus | USEJ0057 | Survey | Gas bocor dan tombol tidak responsif; Pengecekan koneksi, koneksi baik | Pending | Perlu penggantian keypad front panel | Wahyu
19000691 | 2024-01-31 | RSUD Pandan Arang | Baby Incubator | Draeger/Isolette C2000 | CF32780 | Service | Humidity tidak naik; Pembersihan chamber himidity | Ready to use |  | Nanda
7000524 | 2024-01-31 | RS JIH | Mesin Anasthesia | Draeger/ATLAN A350 | ASRJ0460 | Instalasi | Alat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
19000689 | 2024-01-30 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Service | Prisma pecah; Penggantian prisma | Ready to use |  | Nanda
5000191 | 2024-01-30 | RSUD Benda | Gynecology Bed | Syas/SGT1 | 230964112000400747 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000191 | 2024-01-30 | RSUD Benda | Gynecology Bed | Syas/SGT1 | 230964112000400749 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000191 | 2024-01-30 | RSUD Benda | Gynecology Bed | Syas/SGT1 | 230964112000400746 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000191 | 2024-01-30 | RSUD Benda | Gynecology Bed | Syas/SGT1 | 230964112000400748 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000191 | 2024-01-30 | RSUD Benda | Gynecology Bed | Syas/SGT1 | 230964112000400752 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
20000264 | 2024-01-29 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120S | S2SNM2639 | Survey | Respirasi low dan heart rate tidak terbaca | Ready to use |  | Sujana
20000264 | 2024-01-29 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120S | S2SNM2651 | Survey | Respirasi low dan heart rate tidak terbaca | Ready to use |  | Sujana
11000305 | 2024-01-29 | RSUD Tanduale | Ventilator | Draeger/Babylog VN800 | ASPH0008 | Service | Internal gas sensor fail; Service hit software | Ready to use | Disarankan mengganti maintenance kit atau kontrak service | Nasrul
16000225 | 2024-01-29 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Service | Coupling low speed mampet; Penggantian o ring | Ready to use |  | Ipan
10000180 | 2024-01-27 | RSI Sultan Agung Banjarbaru | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0409 | Service | Pembersihan breathing circuit | Ready to use |  | Heru
10000179 | 2024-01-27 | RSI Sultan Agung Banjarbaru | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0126 | Service | Piston bocor; Pembersihan valve, lempengena besi Co2y | Ready to use | Disarankan penggantian testlang | Heru
6000752 | 2024-01-26 | RSUD Ajibarang | Infant Warmer | Draeger/Resuscitaire | DSNJ0049 | Service | Error 05, heater putus; Pemasangan heater baru | Ready to use |  | Afrizal
19000702 | 2024-01-26 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0179 | Survey | Failure saat device check; Tarik error - Valve Ptank dan PSY5 rusak | Pending | Disarankan penggantian Valve Ptank dan PSY5 | Nanda
9000421 | 2024-01-25 | RSUD Dr. H. Chasan Boesoirie | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/089/C1 | Survey | Uji kesesuaian dan uji paparan | Ready to use |  | Ruslan
19000687 | 2024-01-25 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0044 | Service | Inspiration dan expriration flow failure; Penggantian flow sensor | Ready to use |  | Nanda
4000424 | 2024-01-24 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0919012 | Service | Penggantian EMH | Ready to use |  | Bayu
19000685 | 2024-01-24 | RS JIH | Mesin Anasthesia | Draeger/ATLAN A350 | ASMM0125 | Maintenance | Kontrak service maintenance | Ready to use |  | Nanda
19000686 | 2024-01-24 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Maintenance | Kontrak service maintenance | Ready to use |  | Nanda
7000529 | 2024-01-23 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0026 | Maintenance | Maintenance total care | Ready to use |  | Wahyu
8000338 | 2024-01-23 | RSUD RA Kartini | Ventilator | Draeger/Savina 300 | ASEL0163 | Survey | Error 42.000; General and system check, check pulse on motor control pcb | Pending | Perlu penggantian blower dan driver control | Amna
19000684 | 2024-01-23 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Service | Setting adjusment xray centering | Ready to use |  | Nanda
3000294 | 2024-01-23 | RSUD Alimuddin Umar | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0402 | Survey | Generator tidak bisa auto ON, tekanan sering drop; Ada kerusakan pada board mover | Pending | Sudah diberikan penawaran | Imam
11000304 | 2024-01-22 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASLA0082 | Survey | Unit mati total, motor knop tidak berfungsi, alarm speaker; PBA P1 power, PBA connector, Motor, dan speaker rusak | Pending | Disarankan penggantian PBA P1 power, PBA connector, motor knop, speaker set, pressure regulator dan membran expiration | Nasrul
9000463 | 2024-01-21 | RSUD Dr. H. Chasan Boesoirie | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/089/C1 | Service | Penggantian detector | Ready to use |  | Ruslan
11000303 | 2024-01-19 | RSKD Dadi | C-arm | Apelem/Evo R+ | 55120063849 | Service | Blue screen, tidak bisa booting; Reload software | Ready to use |  | Nasrul
19000683 | 2024-01-18 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASHM0015 | Service | Inspiration dan expriration flow failure; Tarik error - tidak ada masalah | Ready to use |  | Nanda
19000683 | 2024-01-18 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASJE0005 | Service | Inspiration dan expriration flow failure; Tarik error - tidak ada masalah | Ready to use |  | Nanda
9000418 | 2024-01-18 | RSUD Tobelo | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/088/C1 | Service | Error 6005; Perbaikan koneksi DICOM | Ready to use |  | Ruslan
15000243 | 2024-01-17 | RSUD Hasanuddin Damrah | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/138/C2 | Survey | Training alat | Ready to use |  | Sutan
8000337 | 2024-01-17 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0136 | Maintenance | Preventive Maintenance | Ready to use | Perlu service set 3years | Amna
8000337 | 2024-01-17 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASHM0012 | Maintenance | Preventive Maintenance | Ready to use | Perlu service set 3years | Amna
8000337 | 2024-01-17 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Primus | ASEH0118 | Maintenance | Preventive Maintenance | Ready to use | Perlu service set 3years | Amna
7000521 | 2024-01-16 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | NF31848 | Survey | Penggantian engsel | Ready to use |  | Wahyu
7000521 | 2024-01-16 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | FH36368 | Survey | Penggantian engsel | Ready to use |  | Wahyu
7000521 | 2024-01-16 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | FH36366 | Survey | Penggantian engsel | Ready to use |  | Wahyu
6000753 | 2024-01-16 | RSUD Dr. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEL0081 | Survey | O2 supply warna merah, air supply merah; Pengecekan software dan hardware | Pending | Disarankan penggantian flow sensor, proportional valve, housing set, pressure regulator, kit air filter, particle filter, filter electronic | Afrizal
19000682 | 2024-01-15 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | QC remake terus; Setting detector dan monoblok | Ready to use |  | Nanda
9000414 | 2024-01-13 | RSUD Labuha | X-ray Digital | Apelem/Camargue DR | IMA704030 | Service | Tabung tube miring; Setting tabung tube | Ready to use |  | Ruslan
6000754 | 2024-01-12 | RSUD Brebes | Ventilator | Draeger/Evita 300 | ASJH0125 | Survey | VTE tidak stabil, nebulizer merah; Perlu penggantian cable harness spirolog flow sensor dan baterai pack | Pending |  | Afrizal
6000754 | 2024-01-12 | RSUD Brebes | Ventilator | Draeger/Evita 300 | ASHA0095 | Survey | Pengecekan unit; Unit berfungsi dengan baik | Ready to use |  | Afrizal
19000699 | 2024-01-12 | Rumah Sakit St. Elisabeth | Lampu Operasi | Draeger/Polaris 600 | ASRK1018 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
8000336 | 2024-01-12 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASJF0033 | Survey | Ventilator not constant; system check, check diaphragm cup & piston | Pending | Perlu penggantian service set 3 thn | Amna
19000681 | 2024-01-11 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN800 | ASRL0013 | Survey | Pressure sensor calibration failed; Tarik error - kerusakan pada selenoid valve 12V | Pending |  | Nanda
17000178 | 2024-01-11 | RSUD Dr. Drajat Prawiranegara | ESWL | Medispec/EM1000 | 0617048 | Service | Penggantian EMH. EMH shock habis | Ready to use |  | Bahar
6000755 | 2024-01-11 | RS Bhakti Asih Brebes | Ventilator | Draeger/Savina 300 | ASHD0011 | Survey | Error 30.000, error 40.2000; Pengecekan software dan hardware | Pending | Perlu penggantian service set 6thn dan 2 kit baterai internal | Afrizal
6000756 | 2024-01-10 | RSIA Bunda Arif | Jaudience Meter | Draeger/JM105 | B3501432 | Survey | Error 05 | Pending | Perlu penggantian baterai internal | Afrizal
11000301 | 2024-01-09 | RSUD Kota Kendari | ESWL | Medispec/E3000 | 0922003 | Maintenance | Preventive Maintenance | Ready to use |  | Nasrul
17000177 | 2024-01-09 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0917012 | Survey | Bunyi tembakan keras, plat EMH pecah | Pending | Disarankan mengganti EMH | Bahar
18000341 | 2024-01-08 | RSUD Kalooran Minahasa Selatan | Ventilator | Draeger/Savina300 | ASMM0158 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
11000302 | 2024-01-08 | RSUD Kota Kendari | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/127/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Nasrul
3000293 | 2024-01-08 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Penggantian joystick | Ready to use |  | Eddy
18000343 | 2024-01-05 | RS Siloam Sonder | Ventilator | Draeger/Savina300 | ASSM0159 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
17000176 | 2024-01-05 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0488 | Service | baterai ups rusak, kamera mati, gambar terpotong. Adjust colimator tube | Ready to use | Disarankan penggantian baterai ups. Camera sudah berfungsi | Bahar
18000342 | 2024-01-04 | RSU GMIM Tonsea | Ventilator | Draeger/Savina300 | ASSN0157 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
17000175 | 2024-01-04 | RSUD Berkah Pandeglang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0497 | Service | Pada saat meja diturunkan, pergerakan berhenti | Pending | Disarankan penggantian resistor break | Bahar
19000678 | 2024-01-03 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Survey | Selenoid T3 rusak (Valve actuator DC); Sementara diganti dengan selenoid backup | Ready to use | Disarankan dilakukan penggantian valve actuator DC | Nanda
8000335 | 2024-01-02 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Service | Penggantian service set 3 tahun | Ready to use |  | Amna
8000335 | 2024-01-02 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | Service | Penggantian service set 3 tahun | Ready to use |  | Amna
7000525 | 2023-12-28 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASHL0031 | Survey | Unit restrat; Cek M16 therapy board (Rusak) | Pending | Perlu penggantian board M16 | Wahyu
15000242 | 2023-12-28 | RSUD Hasanuddin Damrah | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/138/C2 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Sutan
17000174 | 2023-12-27 | RSPAD Gatot Subroto | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150377 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar
10000178 | 2023-12-27 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Survey | Pengecekan unit | Ready to use |  | Heru
20000269 | 2023-12-26 | RSUP Tadjudin Chalid | Digital Radiografi Fluoroskopi | Apelem/Platinum | J236F0882 | Instalasi | Proses installasi; | Pending |  | Sujana
19000676 | 2023-12-22 | RSUP dr. Kariadi | Jaudience Meter | Draeger/JM105 | H3505909 | Service | Lampu tidak menyala; Unit dikirim ke principle, dan sudah diperbaiki. Pemasangan kembali unit | Ready to use | Unit diberikan kepada tim biomedis rumah sakit | Nanda
19000675 | 2023-12-22 | RSUP dr. Kariadi | Dental Unit | Kavo/1058TM | 2004002722 | Survey | Bearing goyang - handpiece jalur angin patah | Pending | Disarankan penggantian bearing dan head handpiece | Nanda
3000291 | 2023-12-22 | RSAL Dr. Mintohardjo | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150378 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam
19000673 | 2023-12-21 | RSUP dr. Kariadi | Infant Warmer | Draeger/Resuscitaire | AF10367 | Service | Error 05, heater putus dan kabel sensor suhu rusak | Ready to use |  | Nanda
19000673 | 2023-12-21 | RSUP dr. Kariadi | Infant Warmer | Draeger/Resuscitaire | AF10368 | Service | Error 05, heater putus dan kabel sensor suhu rusak | Ready to use |  | Nanda
19000672 | 2023-12-21 | RSUP dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0305 | Service | VT low; Penggantian flow sensor | Ready to use |  | Nanda
17000173 | 2023-12-21 | RSUD Labuha | X-ray Digital | Apelem/DR Camargue | IMA704030 | Service | Hasil gambar terpotong. Adjust colimator | Ready to use |  | Bahar
1000177 | 2023-12-21 | RSUD Ulin | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150176 | Service | Monitor tidak bisa dirubah posisi ke atas. Pengencangan baut | Ready to use |  | Heru
8000334 | 2023-12-21 | RSI Sultan Agung Semarang | Baby Incubator | Draeger/Isolette C2000 | DSSL0049 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
3000292 | 2023-12-21 | RS Polri Said Sukanto | Laser Surgical | Jena Surgical/MultiPulse HoPlus | T191150376 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam
5000193 | 2023-12-20 | RSUD Benda | Emergency trolly | SyasTech/Emergency Trolly | 231169501551 | Instalasi | alat berfungsi dengan baik | Ready to use |  | Bilal
10000176 | 2023-12-20 | RSUD Ulin | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150176 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Heru
20000267 | 2023-12-19 | RSUD Bahteramas | Incubator Transport | Draeger/TI500 | MA12470 | Service | Penggantian baterai dan air filter | Ready to use |  | Sujana
17000172 | 2023-12-18 | RSD Tidore | X-ray Digital | Apelem/DR Camargue | 2510220561 | Service | Hasil thoracal dan lumbal kurang baik. Melakukan kalibrasi tube | Ready to use |  | Bahar
7000519 | 2023-12-15 | RSUD Merah Putih | Baby Incubator | Draeger/Isolette C2000 | DSRL0026 | Service | O2 calibration fail;Pemasangan oxygen sensor kit and plate, oxygen cell mounting | Ready to use |  | Wahyu
3000290 | 2023-12-14 | RSUD Provinsi NTB | Digital Radiografi Fluoroskopi | Apelem/Platinum | B196F0329 | Instalasi | Alat dapat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
15000240 | 2023-12-14 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Service | Penggantian applicator | Ready to use |  | Sutan
15000241 | 2023-12-14 | RS Otak Moh. Hatta | ESWL | Medispec/EM1000 | 0922163 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
19000670 | 2023-12-12 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0075 | Service | gas unit fail; Penggantian blower | Ready to use |  | Nanda
9000410 | 2023-12-12 | RSD Tidore | Baby Incubator | Draeger/Isolette 8000 | DH08425 | Survey | Suhu tidak stabil; Pengecekan motor fan pada unit | Ready to use |  | Ruslan
9000410 | 2023-12-12 | RSD Tidore | Baby Incubator | Draeger/Isolette 8000 | DH08424 | Survey | Suhu tidak stabil; Pengecekan motor fan pada unit | Ready to use |  | Ruslan
7000518 | 2023-12-11 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Babyleo | ASMJ0026 | Survey | Humidity tidak sesuai; Pengecekan heater-humidity. Board PBA AC rusak | Pending |  | Wahyu
19000669 | 2023-12-11 | RSUP dr. Kariadi | Ventilator | Draeger/Savina300 | ASMK0322 | Service | Tekanan O2 rendah; Adjusment ambients pressure | Ready to use |  | Nanda
8000329 | 2023-12-11 | RSI Sunan Kudus | Mesin Anasthesia | Draeger/Atlan300 XL | ASNL0042 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
19000667 | 2023-12-08 | RS Panti Wilasa | Ventilator | Draeger/Savina 300 | ASMM0026 | Service | Blower berdengung; Expiration valve membran hilang | Pending | Disarankan penggantians service set 3 tahun | Nanda
19000666 | 2023-12-07 | RSUD Dr. Moewardi | Infant Warmer | Draeger/Resuscitaire | HR12025 | Service | Resusitasi tidak keluar; Pemasangan circuit sudak dibenarkan | Ready to use |  | Nanda
19000665 | 2023-12-07 | RSUD Dr. Moewardi | Ventilator | Draeger/Babylog VN500 | ASEJ0053 | Service | Layar restart sendiri; Update software | Ready to use |  | Nanda
7000517 | 2023-12-07 | RS Panti Rapih | Ventilator | Draeger/Oxylog 2000 Plus | ASBN0047 | Survey | VT tidak sesuai; Pengecekan flow sensor. Flow sensor rusak | Pending | Perlu penggantian flow sensor | Wahyu
6000757 | 2023-12-06 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF007 | Service | Penggantian O2 sensor, spirolog flow sensor, pressure regulator, fabius GS service set and cozy service set | Ready to use |  | Afrizal
6000757 | 2023-12-06 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0071 | Service | Penggantian O2 sensor, spirolog flow sensor, pressure regulator, fabius GS service set and cozy service set | Ready to use |  | Afrizal
18000340 | 2023-12-05 | RSUD Ainun Habibie | Incubator Transport | Draeger/TI500 | DSSL0006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
2000177 | 2023-12-04 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Error table top, bergerak sendiri; Penggantian joystick set | Ready to use |  | Eddy
9000408 | 2023-12-01 | RSUD Dr. H. Chasan Boesoirie | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/089/C1 | Service | Board 204 tidak ada; Pemasangan board 204 | Pending | Hasil gambar masih bergaris | Ruslan
6000758 | 2023-12-01 | RSUD Brebes | Baby Incubator | Draeger/Caleo | ASJM0056 | Survey | Humidity gak naik naik; Pembersihan pelampung dan container humidity | Ready to use |  | Afrizal
8000328 | 2023-11-30 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Babylog VN600 | ASSK0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
7000516 | 2023-11-30 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001195 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Wahyu
19000662 | 2023-11-30 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0075 | Survey | Gas supply dan air supply error; Tarik error - Setting adjustment motor. Kerusakan pada blower | Pending | Akan diajukan claim warranty | Nanda
8000327 | 2023-11-28 | RSJD Dr. Amino Gondohutomo | Jaudience Meter | Draeger/JM105 | H3510377 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
7000514 | 2023-11-27 | RSA UGM | Baby Incubator | Draeger/Isolette C8000 | DSSK0012 | Instalasi | Alat dapat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
7000513 | 2023-11-27 | RSA UGM | Incubator Transport | Draeger/TI500 | DSSL0007 | Instalasi | Alat dapat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
6000759 | 2023-11-27 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0072 | Service | Pemasangan O2 sensor, spirolog flow power, pressure regulator, fabius GS service | Ready to use |  | Afrizal
3000289 | 2023-11-25 | RSUD Balaraja | C-arm | Apelem/Evo R+ | 55118034657 | Service | Software error. Reload software, penggantian SSD dan upgrade RAM | Ready to use |  | Imam
19000661 | 2023-11-24 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0039 | Service | Mixer sensor check fail; Penggantian sensor pdiff high | Ready to use |  | Nanda
16000224 | 2023-11-24 | Klinik Dermatoven | Meja Operasi | Syas/DS-4J | 001654 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
7000512 | 2023-11-24 | RS PKU Muhammadiyah Bantul | Mesin Anasthesia | Draeger/Fabius GS Premium | ASEH0119 | Service | Gas flow bocor & O2 sensor fail. Kalibrasi gas flow | Ready to use | Disarankan penggantian service set 3 tahun dan O2 capsule | Wahyu
19000660 | 2023-11-23 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | Monoblock dan inverter QC fail; Cek kerusakan. Monoblock dan inverter rusak | Pending |  | Nanda
8000326 | 2023-11-23 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0120 | Service | O2 fail; Penggantian O2 sensor | Ready to use |  | Amna
7000510 | 2023-11-22 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASJH0128 | Service | Gas supply calibatrion failed, baterai failure. Pengencangan valve | Ready to use | Disarankan penggantian baterai | Wahyu
6000761 | 2023-11-22 | RSU An Ni'mah | Ventilator | Draeger/Babylog VN600 | ASSK0026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
7000515 | 2023-11-21 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0026 | Maintenance | Preventive Maintenance total care. Penggantian baterai pack | Ready to use |  | Wahyu
7000509 | 2023-11-21 | RS PKU Muhammadiyah Yogyakarta | Mesin Anasthesia | Draeger/Atlan | ASSL0253 | Instalasi | Alat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
3000288 | 2023-11-21 | RSUD Dr. Adjidarmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Survey | Status grid error (table ngeblock) | Pending | Disarankan penggantian motor grid dan ups untuk komputer WS | Imam
19000657 | 2023-11-20 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Atlan A380 | ASPH0141 | Service | Kebocoran karna piston seret; Pelumasan pada motor piston | Ready to use |  | Nanda
7000511 | 2023-11-20 | RSUD Merah Putih | Jaudience Meter | Draeger/JM105 | H3510375 | Instalasi | Alat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
6000763 | 2023-11-20 | RSU Wiradadi Husada | Ventilator | Draeger/Savina300 | ASSL0087 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000763 | 2023-11-20 | RSU Wiradadi Husada | Ventilator | Draeger/Savina300 | ASSL0086 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000763 | 2023-11-20 | RSU Wiradadi Husada | Ventilator | Draeger/Savina300 | ASSL0100 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000764 | 2023-11-18 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASJH0124 | Survey | VT tidak stabil; Pengecekan software & hardware | Pending | Perlu penggantian flow sensor dan cable harness spirolog flow sensor | Afrizal
7000501 | 2023-11-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Maintenance | Preventive Maintenance total care. Penggantian baterai pack | Ready to use |  | Wahyu
7000503 | 2023-11-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0007 | Maintenance | Preventive Maintenance total care. Penggantian baterai pack | Ready to use |  | Wahyu
7000502 | 2023-11-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Maintenance | Preventive Maintenance total care. Penggantian baterai pack | Ready to use |  | Wahyu
9000402 | 2023-11-17 | RSUD Dr. M. M. Dunda | Mesin Anasthesia | Draeger/Primus | ASDI0027 | Survey | Bar analyser fail, O2 sensor fail, flow sensor fail; Flow power rusak | Pending | Perlu penggantian flow power, flow sensor dan Ilca3 | Ruslan
9000403 | 2023-11-16 | RSUD Tani dan Nelayan | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0080 | Service | Gas analyser fail; Flow power rusak | Pending | Disarankan penggantian flow power dan internal baterai | Ruslan
9000404 | 2023-11-16 | RSUD Tani dan Nelayan | Ventilator | Draeger/Evita V800 | ASPH0002 | Survey | Internal gas fail, pressure sensor fail; Board M1.3 rusak | Pending | Disarankan penggantian board M1.3 | Ruslan
6000767 | 2023-11-16 | RSI Harapan Anda | Ventilator | Draeger/Savina 300 | ARZF2008 | Survey | VT tidak sesuai, baterai drop | Pending |  | Afrizal
6000766 | 2023-11-16 | RSI Harapan Anda | Ventilator | Draeger/Savina 300 | ASJF0001 | Survey | VT kadang tidak stabil, baterai drop; Cek software dan hardware | Pending | Perlu penggantian service set 6thn, O2 sensor, kit 2 baterai eksternal | Afrizal
9000405 | 2023-11-16 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0481 | Maintenance | Preventive Maintenance - Pemasangan baterai, filter baterai, membran | Ready to use |  | Ruslan
9000405 | 2023-11-16 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0135 | Maintenance | Preventive Maintenance - Pemasangan baterai, filter baterai, membran | Ready to use |  | Ruslan
9000405 | 2023-11-16 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0480 | Maintenance | Preventive Maintenance - Pemasangan baterai, filter baterai, membran | Ready to use |  | Ruslan
9000405 | 2023-11-16 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | SNL0476 | Maintenance | Preventive Maintenance - Pemasangan baterai, filter baterai, membran | Ready to use |  | Ruslan
9000406 | 2023-11-15 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0061 | Survey | Gas analyser fail; Flow power rusak | Pending | Perlu penggantian flow power dan ILCA 3 | Ruslan
19000656 | 2023-11-15 | RS Amal Sehat | Ventilator | Draeger/Savina 300 | ASRA0340 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
7000507 | 2023-11-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus | ASKF0069 | Maintenance | Preventive Maintenance total care. Penggantian baterai pack | Ready to use |  | Wahyu
7000506 | 2023-11-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0118 | Maintenance | Preventive Maintenance total care. Penggantian baterai pack | Ready to use |  | Wahyu
7000508 | 2023-11-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Isolette C8000 | DSSF0003 | Service | Controller failure 14. Penggantian board PBA Assy Interface dan heating element | Ready to use |  | Wahyu
8000325 | 2023-11-15 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Service | Gas analyzer failed; Penggantian flow power, motor pump, dan ILCA 3 | Ready to use |  | Amna
6000768 | 2023-11-15 | RSUD Dr. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEL0081 | Survey | O2 supply merah, air supply merah, pressure sensor merah; Pengecekan software dan hardware | Pending | Perlu penggantian 8415081 1baM13 | Afrizal
7000505 | 2023-11-14 | RSUP dr. Soeradji Tirtonegoro | Infant Warmer | Draeger/Caleo | ASKL0038 | Service | Sensor suhu tidak terbaca. Penggantian sensor thermotruce central dan peripheral | Ready to use |  | Wahyu
7000504 | 2023-11-14 | RSUP dr. Soeradji Tirtonegoro | C-arm | Draeger/Babylog VN500 | ASKL0091 | Service | O2 calibration failed, O2 supply fail, air supply fail. Penggantian PATo, Kit air filter dan battery pack | Ready to use |  | Wahyu
100000175 | 2023-11-14 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Error anode; Restart generator manual | Ready to use |  | Heru
3000287 | 2023-11-14 | RSUD Balajara | Baby Incubator | Apelem/Evo R+ | 55118034657 | Survey | Komputer C-arm blue screen, OS tidak booting | Ready to use | Akan dibuatkan penawaran perbaikan | imam
7000484 | 2023-11-13 | RSA UGM | Ventilator | Draeger/Resuscitaire | DSSJ0002 | Instalasi | Alat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
7000484 | 2023-11-13 | RSA UGM | Infant Warmer | Draeger/Resuscitaire | DSSJ0001 | Instalasi | Alat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
7000491 | 2023-11-13 | RSA UGM | Infant Warmer | Draeger/JM105 | H3508916 | Instalasi | Alat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
7000483 | 2023-11-13 | RSA UGM | Jaudience Meter | Draeger/Oxymixer | QM20220205 | Instalasi | Alat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
7000483 | 2023-11-13 | RSA UGM | Bubble CPAP | Draeger/Oxymixer | QM20220208 | Instalasi | Alat berfungsi dengan baik. Aksesoris lengkap | Ready to use |  | Wahyu
8000324 | 2023-11-13 | RSUD Dr. Loekmono Hadi | Infant Warmer | Draeger/Isolette8000 | CU03951 | Service | Penggantian service set 3 tahun, kit access door latch, filter dan fan | Ready to use |  | Amna
7000497 | 2023-11-11 | RSUP Dr. Sardjito | Bubble CPAP | Draeger/Primus | ASHL0059 | Maintenance | Preventive Maintenance total care. Penggantian baterai pack | Ready to use |  | Wahyu
7000496 | 2023-11-11 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Maintenance | Preventive Maintenance total care. Penggantian baterai pack | Ready to use |  | Wahyu
7000497 | 2023-11-11 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Maintenance | Preventive Maintenance total care. Penggantian baterai pack | Ready to use |  | Wahyu
7000500 | 2023-11-11 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Maintenance | Preventive Maintenance total care. Penggantian baterai pack | Ready to use |  | Wahyu
7000494 | 2023-11-10 | RSUD Panembahan Senopati Bantul | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/027/C0 | Service | RIS & Unscheduled unselectable. Cek koneksi LAN dan reconnect sistem | Ready to use |  | Wahyu
6000648 | 2023-11-10 | RSUD Kardinah Tegal | Patient Monitor | Draeger/Infinity Gamma XL | 5517381877 | Service | Layar kebalik; Sudah dilakukan troubleshoot | Ready to use | Unit dalam pemantauan | Afrizal
6000648 | 2023-11-10 | RSUD Kardinah Tegal | Patient Monitor | Draeger/Vista 120 | V3SHN0007 | Survey | NIBP Error; Module NIBP rusak | Pending |  | Afrizal
6000648 | 2023-11-10 | RSUD Kardinah Tegal | Patient Monitor | Draeger/Infinity Gamma XL | 5517266371 | Survey | Tensi error; Mobul NIBP dan Multimed rusak | Pending |  | Afrizal
6000648 | 2023-11-10 | RSUD Kardinah Tegal | Patient Monitor | Draeger/Infinity Gamma XL | 5517262580 | Survey | Unit error; Perlu penggantian NIBP dan Multimed | Pending |  | Afrizal
6000646 | 2023-11-09 | RSI Harapan Anda | Ventilator | Draeger/Savina300 | ASFJ0001 | Survey | Tidak lulus kalibrasi, settingan VT tidak sesuai; Pengecekan software dan hardware | Pending | Akan dijadwalkan ukur ulang dengan mesin kalibrasi, kemungkinan error pada mainboard | Afrizal
6000646 | 2023-11-09 | RSI Harapan Anda | Ventilator | Draeger/Savina300 | ARZF0049 | Survey | Tidak lulus kalibrasi, hasil VT tidak sesuai, baterai failed; Pengecekan software dan hardware | Pending | Disarankan ganti mainboard, service set 3thn, dan O2 cell | Afrizal
6000645 | 2023-11-09 | RSI Harapan Anda | Ventilator | Draerger/Savina300 | ASMM0051 | Service | Pemasangan service set 2 tahun | Ready to use |  | Afrizal
15000239 | 2023-11-09 | RSUP Dr. M. Djamil Padang | Laser Surgical | Jena Surgical/MultiPulse TM+1470 | J190100147 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Sutan
7000495 | 2023-11-09 | RS JIH | Ventilator | Draeger/Babylog VN500 | ASPA0051 | Service | Penggantian baterai pack | Ready to use |  | Wahyu
19000655 | 2023-11-08 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Service | Alat mati total; Penggantian PSU dan baterai | Ready to use |  | Nanda
7000493 | 2023-11-08 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Wahyu
3000285 | 2023-11-08 | RSUD Pringsewu | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0756 | Service | Selalu muncul notif pada console setelah exposed; Upgrade software, update filmware generator dan software platinum view | Ready to use |  | Imam
3000286 | 2023-11-08 | RSUD Pringsewu | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0756 | Maintenance | Preventinve Maintenance Warranty | Ready to use |  | Imam
6000644 | 2023-11-07 | RSUD Madjenang | Ventilator | Draeger/Savina 300 | ASNL1000 | Survey | Pengecekan unit | Ready to use | Perlu penggantian service set dan O2 sensor | Afrizal
15000238 | 2023-11-07 | RS dr. Reksodiwiryo | Mayo Stand | Syas/Mayo Stand | 230735113000400000 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000238 | 2023-11-07 | RS dr. Reksodiwiryo | Mayo Stand | Syas/Mayo Stand | 230735113000400383 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000238 | 2023-11-07 | RS dr. Reksodiwiryo | Mayo Stand | Syas/Mayo Stand | 231035113000400487 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000238 | 2023-11-07 | RS dr. Reksodiwiryo | Mayo Stand | Syas/Mayo Stand | 231035113000400488 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000237 | 2023-11-07 | RS dr. Reksodiwiryo | Stretcher | Syas/Hydraulic | 231031214000400182 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000237 | 2023-11-07 | RS dr. Reksodiwiryo | Stretcher | Syas/Hydraulic | 231031214000400181 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000237 | 2023-11-07 | RS dr. Reksodiwiryo | Stretcher | Syas/Hydraulic | 231031214000400183 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000236 | 2023-11-07 | RS dr. Reksodiwiryo | Emergency trolly | SYAS/SCE2-SS | 231069501000400530 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000234 | 2023-11-07 | RS dr. Reksodiwiryo | Foot step | SYAS | 2310613001000 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000235 | 2023-11-07 | RS dr. Reksodiwiryo | Infused Stand | SYAS | 23103616 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
12000109 | 2023-11-06 | RSUD Toto Kabila | C-arm | Apelem/Evo R+ | 55121004853 | Survey | Power supply terbakar, travo pverload, inverter overload | Pending | Travo dibawa untuk dicek lebih lanjut | Marshall
12000110 | 2023-11-06 | RSUD Toto Kabila | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150173 | Survey | Laser tidak expose; | Pending | Menunggu pemeriksaan lebih lanjut | Marshall
19000653 | 2023-11-06 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDL0003 | Service | Gas anasthesi tidak terbaca; | Pending | Disarankan penggantian scio | Nanda
7000499 | 2023-11-06 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus | ASKF0077 | Maintenance | Preventive Maintenance total care. Penggantian baterai pack | Ready to use |  | Wahyu
8000323 | 2023-11-06 | RS Keluarga Sehat Tayu | UPS | Delta/10000 | UP10202111100022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
2000265 | 2023-11-04 | RSUD Morowali | Bubble CPAP | Draeger/Oxymixer | QM20230004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
3000284 | 2023-11-03 | RSUD Malingping | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/131/C2 | Maintenance | Preventinve Maintenance Warranty | Ready to use |  | Imam
20000260 | 2023-11-02 | RSUD Kota Kendari | ESWL | Medispec/E3000 | 0922003 | Service | Pemasangan instalasi listrik, refresh training | Ready to use |  | Sujana
19000652 | 2023-11-02 | RS Indriati Solo | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0063 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000680 | 2023-11-02 | Rumah Sakit St. Elisabeth | Lampu operasi | Draeger/Polaris 600 | ASRK1018 | Maintenance | General check unit | Ready to use |  | Nanda
18000337 | 2023-11-02 | RSUD Sanana | Dental unit | Kavo/E30TM | 10111150 | Survey | Suction tidak berfungsi saat running pasien | Pending |  | Ruslan
8000322 | 2023-11-01 | RS ISLAM PATI | UPS | Yoshiga/HPT2 | 50C1101193281J2300009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
3000276 | 2023-11-01 | RSUD SUmedang | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/132/C2 | Maintenance | Preventinve Maintenance Warranty | Ready to use |  | Imam
3000283 | 2023-11-01 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Maintenance | Preventinve Maintenance Warranty | Ready to use |  | Imam
4000421 | 2023-10-31 | RSUP dr. Ben Mboi | ESWL | Medispec/EM1000 | 0922162 | Survey | Training alat | Ready to use |  | Bayu
2000176 | 2023-10-31 | RSUD Kab. Temanggung | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/011/C0 | Service | Koneksi unit dengan komputer workstation lost; Bridging connection RJ5 | Ready to use |  | Eddy
10000174 | 2023-10-30 | RSUD Ulin | Mamografi | Metaltronica/Helianthus BYM | HELIDBT/0048/CB1 | Maintenance | General check, alat oke dan hasil normal | Ready to use |  | Heru
3000282 | 2023-10-28 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Service | Dicom error, kabel usb extension rusak; Penggantian kabel usb extension | Ready to use |  | Imam
10000173 | 2023-10-27 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Komputer workstation tidak konek. Konek komputer ws secara manual dengan mode BNC | Ready to use |  | Heru
20000259 | 2023-10-26 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina300 | ASJL0393 | Survey | Mati total ketika listrik padam; Baterai rusak | Pending | Perlu penggantian baterai | Sujana
20000259 | 2023-10-26 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina300 | ASJL0394 | Survey | Mati total ketika listrik padam; Baterai rusak | Pending | Perlu penggantian baterai | Sujana
20000259 | 2023-10-26 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina300 | ASJL0395 | Survey | Mati total ketika listrik padam; Baterai rusak | Pending | Perlu penggantian baterai | Sujana
20000259 | 2023-10-26 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina300 | ASJL0396 | Survey | Mati total ketika listrik padam; Baterai rusak | Pending | Perlu penggantian baterai | Sujana
20000259 | 2023-10-26 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina300 | ASJL0392 | Survey | Mati total ketika listrik padam; Baterai rusak | Pending | Perlu penggantian baterai | Sujana
20000259 | 2023-10-26 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina300 | ASJL0030 | Survey | Mati total ketika listrik padam; Baterai rusak | Pending | Perlu penggantian baterai | Sujana
20000259 | 2023-10-26 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina300 | ASJL0138 | Survey | Mati total ketika listrik padam; Baterai rusak | Pending | Perlu penggantian baterai | Sujana
7000492 | 2023-10-26 | RSUD Wonosari | C-arm | Apelem/Evo R+ | 55121060911 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Wahyu
6000643 | 2023-10-26 | RSU Wiradadi Husada | Mesin Anasthesia | Draeger/Fabius Plus | ASSK0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
2000258 | 2023-10-25 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina300 | ASHD0007 | Survey | Unit restart; Power supply dan baterai bermasalah | Pending | Perlu penggantian power supply dan baterai | Sujana
2000258 | 2023-10-25 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina300 | ASJJ0405 | Survey | Unit restart; Baterai rusak | Pending | Perlu penggantian baterai | Sujana
2000258 | 2023-10-25 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina300 | ASJJ0409 | Survey | Unit restart; Baterai rusak | Pending | Perlu penggantian baterai | Sujana
10000172 | 2023-10-25 | RSI Sultan Agung Banjarbaru | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0409 | Service | Breathing circuit bocor; Pengecekan watertrap, sample line, absorber dan testlang | Ready to use |  | Heru
6000642 | 2023-10-25 | RSUD Brebes | Ventilator | Draeger/Evita V800 | ASPF0153 | Service | Pemasangan part claim warranty "PI mainboard" dan M48 board | Ready to use |  | Afrizal
20000257 | 2023-10-24 | RSUD Sawerigading | Mesin Anasthesia | Draeger/Fabius Plus | ASDH0113 | Service | Pressure regulator berdengung; Penggantian pressure regulator | Ready to use |  | Sujana
4000420 | 2023-10-24 | RSAB Harapan Kita | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/129/C2 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu
2000175 | 2023-10-24 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Service | Hasil x-ray tidak sesuai; Perbaikan posisi tube dengan adjust position antara tube dengan detector | Ready to use |  | Eddy
7000490 | 2023-10-24 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKE0070 | Survey | Unit restart terus. | Pending | Disarankan penggantian baterai internal | Wahyu
20000256 | 2023-10-23 | RSUD I Lagaligo | Ventilator | Draeger/Babylog VN500 | ASKJ0005 | Service | Penggantian kit proportional valve air | Ready to use |  | Sujana
8000330 | 2023-10-21 | RSI Sultan Agung Semarang | Mesin Anasthesia | Draeger/Primus | ASLD0028 | Survey | Gas analyzer fail; System check | Pending | Perlu penggantian flow power, ILCA 3, minimodule motor pump | Amna
7000488 | 2023-10-20 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
7000489 | 2023-10-20 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
6000641 | 2023-10-19 | RS JIH | Jaudience Meter | Draeger/JM 105 | H3507946 | Survey | Alarm time periodic calibration | Pending | Disarankan di adjust dan kalibrasi | Afrizal
18000439 | 2023-10-18 | Puskesmas Mabapura | Bed patient | Syastech/Bed 3 cranck | C-0165 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
7000485 | 2023-10-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus | ASKK0077 | Service | Layar kedip; Mengencangkan kmebali semua konektor bakclight | Ready to use |  | Wahyu
7000482 | 2023-10-17 | RSUD Merah Putih | X-Ray Analog | Apelem/Camargue Analog | 703979 | Survey | Table top tidak berfungsi; Penggantian foot switch | Ready to use |  | Wahyu
7000481 | 2023-10-17 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001182 | Service | Lampu tindakan mati; Penggantian power supply | Ready to use |  | Wahyu
8000321 | 2023-10-17 | RS ISLAM PATI | UPS | Yoshiga/HPT2 | 501101236060KA100047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
3000281 | 2023-10-17 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | Software critical error;Reload software duet | Ready to use |  | Imam
6000640 | 2023-10-13 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0082 | Survey | Bunyi dengung | Pending | Penggantian O2 sensor | Afrizal
6000640 | 2023-10-13 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0025 | Survey | gas sensor failure | Pending | Disarankan penggantian service set 5 tahun, baterai dan breathing circuit | Afrizal
6000640 | 2023-10-13 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Tiro | ASBH0248 | Survey | O2 sensor kalibrasi gagal | Pending | Disarankan penggantian O2 sensor dan layar fabius 6,5 inch | Afrizal
7000480 | 2023-10-12 | RSUD Wonosari | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASLE0013 | Survey | Nilai Vt rendah | Pending | Disarankan penggantian flowsensor | Wahyu
6000639 | 2023-10-12 | RSUD Brebes | Jaudience Meter | Draeger/JM 105 | B3503457 | Survey | Pengecekan unit setelah di kalibrasi dan di adjust | Ready to use |  | Afrizal
13000163 | 2023-10-11 | RSUD Dr. Djasamen Saragih | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/137/C2 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Sofian
3000279 | 2023-10-11 | RSUP dr. Ben Mboi | UPS | K-Star/HPR-10 | 501106300002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam
19000647 | 2023-10-11 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | Air bocor dan tidak bisa pompa; Pembersihan selenoid valve | Ready to use |  | Nanda
18000435 | 2023-10-11 | RSUD Dr. H. Chasan Boesoirie | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/089/C1 | Pra Installasi | Board 204 tidak ada | Pending | Menunggu board dikirim pihak Pabrikan | Ruslan
3000280 | 2023-10-11 | RSUP dr. Ben Mboi | ESWL | Medispec/EM1000 | 0922162 | Service | Pemindahan unit ke ruang baru | Ready to use |  | Imam
6000635 | 2023-10-10 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0042 | Service | Pemasangan pBA CPU Primus | Ready to use |  | Afrizal
6000634 | 2023-10-09 | RS PKU Muhammadiyah Gombong | Mesin Anasthesia | Draeger/Fabius Tiro | ARZM0024 | Survey | Layar blank putih; Pemasangan display board | Ready to use |  | Afrizal
18000434 | 2023-10-07 | RSUD Sanana | Dental Unit | Kavo/E30TM | 10111150 | Service | Scaller dan micromotor tidak berfungsi; Penggantian control valve, selang scaller dan ESB Spray | Ready to use |  | Ruslan
6000633 | 2023-10-07 | RS Duta Mulya | Ventilator | Draeger/Savina 300 | ASSK0023 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000633 | 2023-10-07 | RS Duta Mulya | Ventilator | Draeger/Savina 300 | ASSK0024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
8000320 | 2023-10-06 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0011 | Service | Ventilator low respons; Penggantian diafragma cup pision, filter dan bersihin filter | Ready to use |  | Amna
4000419 | 2023-10-05 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Survey | Uji kesesuaian C-arm ESWL | Ready to use |  | Bayu
17000170 | 2023-10-05 | RSUD Labuha | X-ray Digital | Apelem/Camargue DR | IMA704030 | Service | Koneksi detector terkadang susah menyambung; Penggantian kabel LAN | Ready to use |  | Bahar
6000632 | 2023-10-05 | RSI Banjarnegara | Ventilator | Draeger/Savina 300 | ASSE0141 | Survey | Kalibrasi | Ready to use |  | Afrizal
6000632 | 2023-10-05 | RSI Banjarnegara | Ventilator | Draeger/Savina 300 | ASSE0142 | Survey | Kalibrasi | Ready to use |  | Afrizal
6000632 | 2023-10-05 | RSI Banjarnegara | Ventilator | Draeger/Savina 300 | ASSE0144 | Survey | Kalibrasi | Ready to use |  | Afrizal
6000632 | 2023-10-05 | RSI Banjarnegara | Ventilator | Draeger/Savina 300 | ASSE0143 | Survey | Kalibrasi | Ready to use |  | Afrizal
19000645 | 2023-10-04 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Service | air tidak lancar dan sensor pressure rusak; Penggantian sensor pressure dan water valve | Ready to use |  | Nanda
8000319 | 2023-10-04 | RS Mitra Bangsa Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0041 | Maintenance | System check, running test | Ready to use |  | Amna
7000479 | 2023-10-02 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette C8000 | DSSF0003 | Survey | Controller failure 14 | Pending | Disarankan penggantian heater | Wahyu
6000636 | 2023-09-29 | RSI Bunda Arif | Jaudience Meter | Draeger/JM 105 | B3501432 | Survey | Hasil diluar batas toleransi | Pending | Dianjurkan adjustment dan kalibrasi | Afrizal
19000644 | 2023-09-29 | RS Telogorejo | Mesin Anasthesia | Draeger/Primus | ASHJ0179 | Service | Mixer fail; Pembersihan mixer | Ready to use |  | Nanda
19000643 | 2023-09-27 | RSUD Bung Karno | Mesin Anasthesia | Draeger/Primus | ASMA0192 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000643 | 2023-09-27 | RSUD Bung Karno | Mesin Anasthesia | Draeger/Primus | ASMA0195 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000643 | 2023-09-27 | RSUD Bung Karno | Mesin Anasthesia | Draeger/Primus | ASMA0194 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
7000478 | 2023-09-27 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus | ASKF0069 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
7000471 | 2023-09-26 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0118 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
7000477 | 2023-09-26 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0026 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
2000254 | 2023-09-26 | RSUD Morowali | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/126/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Sujana
9000401 | 2023-09-25 | RSUD Datoe Binangkang | Bed patient | Syastech/Bed 3 cranck | 230968303000400271 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ruslan
7000472 | 2023-09-25 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASPL0087 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
7000476 | 2023-09-25 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus | ASKF0077 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
6000637 | 2023-09-25 | RSI Harapan Anda | Ventilator | Draeger/Evita 4 edition | ASCC0210 | Service | VT tidak konstan | Pending | Disarankan penggantian pressure measurement, valve O2, valve air, PEEP valve, PCB Pneumatic controller dan O2 sensor | Afrizal
8000318 | 2023-09-25 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0037 | Service | Penggantian baterai | Ready to use |  | Amna
8000318 | 2023-09-25 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0033 | Service | Penggantian baterai | Ready to use |  | Amna
7000468 | 2023-09-23 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHL0059 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
7000469 | 2023-09-23 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
7000467 | 2023-09-23 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
7000470 | 2023-09-23 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
7000465 | 2023-09-22 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Delta XL | 6008809569 | Survey | Layar blank hitam | Pending | Disarankan penggantian layar LCD | Wahyu
2000174 | 2023-09-22 | RSUD Kab. Temanggung | Bone Mineral Densitometer | Apelem/Medix DR | H237H1095 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2025-09-22
19000642 | 2023-09-21 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Service | Jumlah air sedikit dan tidak mengembang | Pending | Sudah diganti sensor pressure | Nanda
7000473 | 2023-09-21 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001168 | Service | Lampu tindakan tidak menyala -Power supply mati; Penggantian AC Adapter (Power Supply) | Ready to use |  | Wahyu
7000461 | 2023-09-20 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASPL0088 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
7000466 | 2023-09-20 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASPL0087 | Service | O2 sensor fail dan Gas sensor fail | Pending |  | Wahyu
7000462 | 2023-09-20 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
19000640 | 2023-09-19 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Service | Shock lemah, batu tidak pecah; Penggantian EMH dan plat | Ready to use |  | Nanda
7000474 | 2023-09-19 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0007 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
7000463 | 2023-09-19 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Maintenance | Preventive Care Total Care | Ready to use |  | Wahyu
10000171 | 2023-09-19 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | Tidak ada hasil gambar pembacaaan; Melakukan kalibrasi detector cilling | Ready to use |  | Heru
19000639 | 2023-09-18 | RSUP dr. Kariadi | Dental Unit | Kavo/1058 | 1002722 | Service | membran pecah; Penggantian membran | Ready to use |  | Nanda
6000602 | 2023-09-18 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASKK0232 | Survey | Pengecekan unit | Ready to use |  | Afrizal
6000602 | 2023-09-18 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASKK0231 | Survey | Pengecekan unit | Ready to use |  | Afrizal
6000602 | 2023-09-18 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASKK0230 | Survey | Pengecekan unit | Ready to use |  | Afrizal
7000464 | 2023-09-18 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASHL0023 | Survey | O2 measurement fail, auxilary alarm failed; O2 sensor dan auxilary alarm error perlu diganti | Pending |  | Wahyu
7000460 | 2023-09-18 | RSUD Merah Putih | Infant Warmer | Draeger/Isolette C2000 | DSRL0026 | Survey | Oksigen calibration fail;Oksigen sensor tidak terpasang | Pending | Pengajuan claim warranty | Wahyu
8000317 | 2023-09-18 | RS Keluarga Sehat Tayu | UPS | Yoshiga/HPT1 | 501101213200K6200034 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
6000601 | 2023-09-15 | RSUD Dr. Soeselo Slawi | Jaudience Meter | Draeger/JM 105 | H3508915 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
7000459 | 2023-09-15 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Service | Cover pecah; Penggantian lower snap cap tray | Ready to use |  | Wahyu
7000458 | 2023-09-14 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001168 | Survey | Lampu tindakan tidak menyala; Power supply mati | Pending | Melakukan ajukan claim warranty power supply | Wahyu
4000418 | 2023-09-14 | Dinas Kesehatan Murotai | Mobile X-ray | Apelem/Rafale | 1872342200002 | Survey | Pelaksanaan uji kesesuaian dan Uji Paparan | Ready to use |  | Bayu
10000170 | 2023-09-14 | RSUD Ulin | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J19150176 | Maintenance | Preventive Maintenance | Ready to use | Fiber optic 272 patah ditengah | Heru
6000603 | 2023-09-14 | RSUD Brebes | Baby Incubator | Draeger/Isolette C2000 | HM31390 | Service | Pemasangan evaporator | Ready to use |  | Afrizal
19000637 | 2023-09-13 | RSUD Dr. Moewardi | Mesin Anasthesia | Draeger/Primus | ASNE0153 | Service | Setting dan adjustment | Ready to use |  | Nanda
17000169 | 2023-09-13 | RS Immanuel | Table Examination | SYASTECH/PREMIUM EXAMINATION TABLE | 2.3083e+17 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2024-09-13
2000173 | 2023-09-13 | RSK Dharmais | Digital Radiografi Fluoroskopi | Apelem/Platinum | L216F0664 | Service | Penggantian display kit console | Ready to use |  | Eddy
8000316 | 2023-09-13 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0138 | Maintenance | System check, running test, check valve, check O2 sensor | Ready to use |  | Amna
8000316 | 2023-09-13 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0136 | Maintenance | System check, running test, check valve, check O2 sensor | Ready to use |  | Amna
8000316 | 2023-09-13 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0118 | Maintenance | System check, running test, check valve, check O2 sensor | Ready to use |  | Amna
10000169 | 2023-09-12 | RSUD Sultan Suriansyah | Dental Unit | Kavo/1058TM | 10205321 | Service | Penggantian mainboard control | Ready to use |  | Heru
6000604 | 2023-09-12 | RSUD Dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Fabius Plus | ASLH0077 | Survey | cek kalibrasi vaporizer | Ready to use |  | Afrizal
7000457 | 2023-09-11 | RSKB Sinduadi | Jaudience Meter | Draeger/JM105 | MU20105 | Service | Penggantian baterai | Ready to use |  | Wahyu
19000636 | 2023-09-11 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0025 | Service | Layar bergaris | Pending | Perlu penggantian layar lcd board | Nanda
8000315 | 2023-09-11 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0011 | Service | O2 sensor fail; Pengecekan O2 sensor | Ready to use | Perlu penggantian O2 sensor | Amna
8000315 | 2023-09-11 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0120 | Service | O2 sensor fail; Pengecekan O2 sensor | Ready to use | Perlu penggantian O2 sensor | Amna
8000314 | 2023-09-11 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0138 | Maintenance | Penggantian service set 3 tahun dan O2 sensor | Ready to use |  | Amna
15000233 | 2023-09-11 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Fixed merah; Melakukan pembersihan cpu workstation dan pu box. Penggantian hardisk komputer pu box | Ready to use |  | Sutan
19000635 | 2023-09-08 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0042 | Service | Dima & coma fail SW/HW; Reload software | Ready to use |  | Nanda
8000313 | 2023-09-08 | RS Fastabiq Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASRC0070 | Maintenance | System check, cleaning valve | Ready to use | Perlu penggantian O2 sensor | Amna
6000605 | 2023-09-07 | RSUD Dr. Soeselo Slawi | Mobile X-ray | Apelem/Rafale | 190614220001 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Afrizal
7000455 | 2023-09-07 | RSUD Merah Putih | Baby Incubator | Draeger/Isolette C2000 | DSRL0028 | Instalasi | Alat berfungsi dengan baik; Aksesoris lengkap | Ready to use |  | Wahyu
7000456 | 2023-09-07 | RSUD Merah Putih | Baby Incubator | Draeger/Isolette C2000 | DSRL0026 | Instalasi | Alat berfungsi dengan baik; Aksesoris lengkap | Ready to use |  | Wahyu
16000222 | 2023-09-07 | RSUD Sanana | Dental Unit | Kavo/E30TM | 10111150 | Service | Air handpiece dan scaler tidak keluar, pengecekan air dan sensor katup air | Ready to use | Diperlukan penggantian block katup air | Ipan
13000160 | 2023-09-06 | RSUD Arifin Achmad | Dental Unit | Kavo/E30TM | 10111156 | Survey | Troubleshoot dengan board back up | Pending | Sudah diajukan penawaran untuk penggantian mainboard | Sofian
5000192 | 2023-09-06 | RSUD Malingping | Dental Unit | Kavo/1058TM | 10205319 | Service | Penggantian main board | Ready to use |  | Bilal
15000232 | 2023-09-05 | RSUD Sekayu | X-ray Digital | Apelem/DR Camargue | IMA703957 | Service | Hasil gambar pada pengaturan tertentu hasilnya putih; Kalibrasi tube, pemberian oli pada anoda dan katoda | Ready to use |  | Sutan
7000475 | 2023-09-05 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001197 | Service | Tombol AP1 tidak berfungsi; Mengganti display kit DSE Expert | Ready to use |  | Wahyu
19000633 | 2023-09-04 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | Scanner error; Reconnect dan setting ulang | Ready to use |  | Nanda
10000168 | 2023-09-04 | RSUD Sultan Suriansyah | Dental Unit | Kavo/1058 TM | 10205321 | Service | Troubleshoot dengan board back up | Ready to use |  | Heru
19000632 | 2023-09-01 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0090 | Service | GS500 tidak tersambung | Pending | Akan dijadwalkan dengan teknisi Draeger | Nanda
6000606 | 2023-09-01 | RSUD Brebes | Ventilator | Draeger/Babylog VN500 | ASLF0053 | Service | Upgrade babylog ke HFO | Ready to use |  | Afrizal
6000609 | 2023-08-31 | RSI Harapan Anda | Ventilator | Draeger/Savina 300 | ASMM0051 | Survey | Unit sering mati bootlooping | Pending | Disarankan penggantian service set 2 tahun | Afrizal
6000608 | 2023-08-31 | RSI Harapan Anda | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASLE0048 | Survey | Device check incomplete, test lung tidak sesuai, pressure regulator tabung rusak | Pending |  | Afrizal
6000607 | 2023-08-31 | RSUD Dr. Soeselo Slawi | Ventilator | Draeger/Savina 300 | ASHD0014 | Survey | Baterai internal dan eksternal drop | Pending | Disarankan penggantian service set 6 tahun, kit 2 baterai external dan oxytrace | Afrizal
4000416 | 2023-08-31 | RSUD Panembahan Senopati Bantul | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/027/C0 | Survey | Nilai akurasi dan kolimasi tidak sesuai, kalibrasi detector namun masih muncul artefak; Disarankan penggantian grid dan ukes ulang | Pending | Sudah diajukan penawaran grid | Bayu
8000312 | 2023-08-31 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina 300 | ASRE0304 | Service | O2 measurement fail; kalibrasi O2 manual | Ready to use |  | Amna
8000312 | 2023-08-31 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina 300 | ASPJ0184 | Service | O2 measurement fail; kalibrasi O2 manual | Ready to use |  | Amna
19000630 | 2023-08-30 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Service | Fan failure; Penggantian fan | Ready to use |  | Nanda
6000610 | 2023-08-30 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHM0013 | Service | Pemasangan Pba VGC power board, prmus breathsy dan service set battery | Ready to use |  | Afrizal
4000417 | 2023-08-30 | RSUP dr. Ben Mboi | Mobile X-ray | Apelem/Rafale BDR 40 | 97123004385 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2024-08-30
19000629 | 2023-08-29 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDK0300 | Survey | Layar mati backlight nyala; Kemungkinan ada kerusakan pada board control LCD | Pending | Perlu penggantian board control LCD | Nanda
7000454 | 2023-08-29 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Service | O2 sensor failed; Penggantian part PIA 2 | Ready to use |  | Wahyu
7000453 | 2023-08-28 | RSUD Sleman | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0128 | Service | Flow sensor calibration system failure; Penggantian part selenoid | Ready to use |  | Wahyu
10000167 | 2023-08-28 | RSUD Ulin | Laser Surgical | Jena Surgical/MultiPulse HoPlus | 119150176 | Maintenance | Preventive Maintenance | Ready to use | Fiber optic 272 patah ditengah | Heru
6000611 | 2023-08-28 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASHL0011 | Service | Pemasangan breathsy MX08890 dan MX09999 | Ready to use |  | Afrizal
6000612 | 2023-08-28 | RSUD Banyumas | C-arm | Apelem/Evo R+ | 55016055585 | Service | Penggantian baterai CMOS pada board | Ready to use |  | Afrizal
17000168 | 2023-08-26 | RSUD Toto Kabila | X-ray Digital | Apelem/Camargue DR | IMA703943 | Service | Kolimasi tidak sesuai; Adjust kolimasi sesuai dengan ketentuan | Ready to use |  | Bahar
3000277 | 2023-08-26 | RSUD Kab. Bekasi | C-arm | Apelem/Evo R+ | 55118026649 | Service | Unit berbahasa Perancis; Setting ulang bahasa | Ready to use |  | Imam
5000191 | 2023-08-25 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Service | Bocor angin pada sequence condensator; Perbaikan pada sequence condensator separator | Ready to use |  | Bilal
17000167 | 2023-08-25 | RSUD Toto Kabila | Mobile X-ray | Apelem/Rafale EV30 | 009/19/01305 | Service | Battery fault; Pengecekan tegangan pada board | Ready to use |  | Bahar
7000487 | 2023-08-25 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Maintenance | Preventive Maintenance | Ready to use |  | Faugi
6000613 | 2023-08-24 | RSUD Banyumas | C-arm | Apelem/Evo R+ | 55016055585 | Service | Bootloop pada unit | Ready to use | Sementara board dibawa | Afrizal
5000190 | 2023-08-24 | Signature Dental Care | Dental Unit | Kavo/1058TM | 10207194 | Service | Air pada bowl tidak keluar. Penggantian valve selenoid | Ready to use |  | Bilal
8000310 | 2023-08-23 | RSI Sultan Agung Semarang | Infant Warmer | Draeger/Resuscitaire | DSSE0043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000309 | 2023-08-22 | RS Keluarga Sehat Tayu | UPS | Yoshiga/HPT1 | 501101213200KB200010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000309 | 2023-08-22 | RS Keluarga Sehat Tayu | UPS | Yoshiga/HPT1 | 501101213200KB200009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000309 | 2023-08-22 | RS Keluarga Sehat Tayu | UPS | Yoshiga/HPT1 | 501101213200KB200008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000309 | 2023-08-22 | RS Keluarga Sehat Tayu | UPS | Yoshiga/HPT1 | 501101213200KB200003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
19000627 | 2023-08-22 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Service | water level tidak bisa kembali sehingga bocor | Pending | Sudah diganti water level sensor namun masih error | Nanda
5000189 | 2023-08-22 | RS Dr. Suyoto | Dental Unit | Kavo/1058TM | 10207193 | Maintenance | Preventive Maintenance | Ready to use |  | Bilal
5000188 | 2023-08-22 | RS Dr. Suyoto | Dental Unit | Kavo/1058TM | 10207194 | Maintenance | Preventive Maintenance | Ready to use |  | Bilal
19000626 | 2023-08-21 | RSUD Pandan Arang | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASKL0036 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000626 | 2023-08-21 | RSUD Pandan Arang | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASKL0040 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
8000308 | 2023-08-16 | RSUD Ki Ageng Getas Pondowo Grobogan | Infant Warmer | Draeger/Resuscitaire | AC13515 | Service | Error 21; Penggantian PCB1 Display | Ready to use |  | Amna
7000486 | 2023-08-16 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Maintenance | Preventive Maintenance | Ready to use |  | Wahyu
7000451 | 2023-08-16 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASJE0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000452 | 2023-08-16 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette C8000 | DSSF0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
2000172 | 2023-08-16 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Service | Error ET61, penggantian kabel RJ45 untuk motor driver | Ready to use |  | Eddy
2000170 | 2023-08-15 | RSUD Kab. Temanggung | C-arm | Apelem/Evo R+ | 55115011409 | Service | Break arm tidak bisa dikunci, adjust break arm handle | Ready to use |  | Eddy
16000221 | 2023-08-15 | RSUD Karawang | Dental Unit | Kavo/Primus 1058 | 10205314 | Service | Waterblock sudah dibersihkan dan sudah dipasang kembali | Ready to use |  | Ipan
8000307 | 2023-08-14 | RS ISLAM PATI | Lampu Operasi | Draeger/Polaris 600 | ASRK0101 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
7000450 | 2023-08-14 | RS Panti Rapih | Mesin Anasthesia | Draeger/Fabius Plus | ASDA008 | Maintenance | Pemasangan service set 3 tahun | Ready to use |  | Wahyu
7000450 | 2023-08-14 | RS Panti Rapih | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0193 | Maintenance | Pemasangan service set 3 tahun | Ready to use |  | Wahyu
6000616 | 2023-08-12 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Evita V300 | ASHL0169 | Service | Pemasangan baterai internal, filter O2 | Ready to use |  | Afrizal
7000449 | 2023-08-11 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Survey | Ventilator inaporation; Penggantian flow sensor, breathing circuit leak, sample line dan water lock | Ready to use |  | Wahyu
7000448 | 2023-08-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Survey | O2 sensor fail; Terdapat kerusakan pada sensor PIA | Pending |  | Wahyu
7000447 | 2023-08-10 | RSI Klaten | Mesin Anasthesia | Draeger/Fabius | ARWC0166 | Survey | Battery fault; Kerusakan pada Trafo step down, internal battery, motor piston error, insiratory valve, oksigen sensor | Pending |  | Wahyu
15000231 | 2023-08-10 | RSUD Dr. M. Yunus | ESWL | Medispec/EM1000 | 0917011 | Service | Membran tidak otomatis berhenti; Reminder user untuk menekan tombol membran yang ada di console | Ready to use |  | Sutan
2000162 | 2023-08-10 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | J166F0184 | Service | Penggantian kolimator | Ready to use |  | Eddy
8000305 | 2023-08-10 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASNE0151 | Survey | O2 sensor fail | Pending | Perlu penggantian PIA2 | Amna
3000275 | 2023-08-10 | RSUD Provinsi NTB | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/032/C0 | Service | Compressing fault. PCB 215 trouble. | Ready to use | Disarankan penggantian pcb 215 dan konektor kabel | Imam
19000679 | 2023-08-09 | Rumah Sakit St. Elisabeth | Lampu Operasi | Draeger/Polaris 600 | ASRK1018 | Maintenance | Pengecekan general unit | Ready to use |  | Nanda
19000625 | 2023-08-09 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Service | Breathing circuit bocor, venti bocor | Pending | Disarankan penggantian peep valve, distributing valve dan service set | Nanda
15000230 | 2023-08-09 | RSUD Hasanuddin Damrah | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/138/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Sutan
7000446 | 2023-08-09 | RSUP dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0109 | Survey | Layar bercak; LCD rusak | Pending | Disarankan penggantian layar LCD | Wahyu
8000304 | 2023-08-09 | RSUD Kayen Pati | Baby Incubator | Draeger/Babytherm 8004 | ASEJ0008 | Survey | Heater off dan error 05 | Ready to use |  | Amna
8000304 | 2023-08-09 | RSUD Kayen Pati | Baby Incubator | Draeger/Resuscitaire | BV14106 | Survey | Heater off dan error 05 | Pending | Perlu penggantian heater assy | Amna
19000624 | 2023-08-08 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | Tidak bisa QC/kalibrasi; Cleaning shutter | Ready to use |  | Nanda
8000303 | 2023-08-08 | RSUD dr. R. Soetrasno | Ventilator | Draeger/Evita 4 edition | ASDC0183 | Survey | O2 fail alarm | Pending | Perlu penggantian O2 sensor, housing arm dan door latch | Amna
8000303 | 2023-08-08 | RSUD dr. R. Soetrasno | Ventilator | Draeger/Evita 4 edition | ASDC0184 | Survey | O2 fail alarm | Pending | Perlu penggantian O2 sensor, housing arm dan door latch | Amna
5000187 | 2023-08-08 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Service | Penggantian applicator | Ready to use |  | Bilal
17000166 | 2023-08-08 | RSUD Kembangan | Bed Gynaecology Table | Syastech | 1000400099 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2024-08-08
17000166 | 2023-08-08 | RSUD Kembangan | Bed Gynaecology Table | Syastech | 1000400098 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2024-08-08
19000623 | 2023-08-07 | RS Nasional Diponegoro | Infant Warmer | Draeger/Resuscitaire | AF10306 | Service | Error 05; ganti heater | Ready to use |  | Nanda
8000302 | 2023-08-06 | RSI Sultan Agung Semarang | Mesin Anasthesia | Draeger/Primus | ASLD0028 | Maintenance | Penggantians service set 3 tahun | Ready to use |  | Amna
6000618 | 2023-08-06 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0043 | Service | Pemasangan service set 6 tahun | Ready to use |  | Afrizal
6000617 | 2023-08-06 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0042 | Service | Pemasangan service set 6 tahun + battery pack | Ready to use |  | Afrizal
19000621 | 2023-08-05 | RSUP dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0154 | Service | Venti piston failed; Penggantian venti asm | Ready to use |  | Nanda
19000622 | 2023-08-05 | RSUP dr. Kariadi | Baby Incubator | Draeger/Babyleo TN500 | ASRL0127 | Service | Sensor error; Setting dan adjusment | Ready to use |  | Nanda
15000229 | 2023-08-04 | RS Siti Aisyah | Trolley Emergency | Syastech | 2307400296 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000228 | 2023-08-04 | RS Siti Aisyah | Bed Gynaecology Table | Syastech | 23074000464 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000227 | 2023-08-04 | RS Siti Aisyah | Bed Patient | Syas/Noble Bed 4 Motor | 23077910134 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
15000226 | 2023-08-04 | RS Siti Aisyah | Bed Patient | Syas/Noble Bed | 2307790132 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan
17000165 | 2023-08-04 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Service | Power lemah, hvps trouble, hvg trouble; Perbaikan HVG, penggantian HVPS | Ready to use |  | Bahar
16000220 | 2023-08-04 | RSUD Karawang | Dental Unit | Kavo/1058 Life | 10205314 | Service | Waterblock mamper. Pembersihan waterblock | Ready to use |  | Ipan
8000301 | 2023-08-03 | RSUD Dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASSE0030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000301 | 2023-08-03 | RSUD Dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASSE6521 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000306 | 2023-08-03 | RS ISLAM PATI | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0114 | Service | Penggantian O2 sensor | Ready to use |  | Amna
5000186 | 2023-08-02 | RS Abdi Waluyo | EECP | Vasomedical/Lumenair | 5TC1034 | Service | Layar monitor error touchscreennya | Pending | Disarankan penggantian touchscreen | Bilal
7000445 | 2023-08-01 | RS PKU Muhammadiyah Gombong | Mesin Anasthesia | Draeger/Fabius Tiro | ARZM0024 | Survey | Layar blank putih; Board PCBA Lucy 2 | Pending | Disarankan penggantian board pcba lucy 2 | Wahyu
19000620 | 2023-07-31 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Service | Water level low; Mencoba exchange control modul, masih belum berfungsi | Pending |  | Nanda
19000619 | 2023-07-31 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Service | Ventilator leak; Ada kebocoran pada VGC | Pending | Disarankan penggantian peep valve | Nanda
4000413 | 2023-07-31 | RSUD Kab. Bekasi | C-arm | Apelem/Evo R+ | 55118026649 | Service | Pemasangan ccd camera | Ready to use |  | Bayu
18000334 | 2023-07-31 | RSUD Kotamobagu | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0043 | Maintenance | Preventive Maintenance | Ready to use |  | Faugi
19000654 | 2023-07-28 | RSUP dr. Kariadi | Infant Warmer | Draeger/Resuscitaire | AF10368 | Service | Heater putus; Penggantian heater | Ready to use |  | Nanda
7000444 | 2023-07-28 | RSUD Wonosari | C-arm | Apelem/Evo R+ | 55121060911 | Maintenance | Preventive Maintenance | Ready to use |  | Wahyu
7000443 | 2023-07-28 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Maintenance | Preventive Maintenance | Ready to use |  | Wahyu
7000442 | 2023-07-27 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Survey | Layar blank putih; Kemungkinan software bermasalah, board layar dan console dibawa untuk dicek lebih lanjut | Pending |  | Wahyu
18000328 | 2023-07-27 | RSUP Prof. Dr. R. D. Kandou Manado | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J190100103 | Maintenance | Preventive Maintenance | Ready to use |  | Faugi
19000618 | 2023-07-26 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0129 | Maintenance | Preventive Mainteanance Total Care | Ready to use |  | Nanda
19000617 | 2023-07-26 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Survey | Leak breathing system; Karet piston ada yang bocor | Pending |  | Nanda
15000225 | 2023-07-26 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Software error; Pembersihan CPU | Ready to use |  | Sutan
18000332 | 2023-07-26 | RSUD Kotamobagu | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0043 | Service | Pemasangan O2 sensor dan batery set | Ready to use |  | Faugi
18000333 | 2023-07-26 | RSUD Kotamobagu | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0043 | Service | Penggantian housing O2 sel | Ready to use |  | Faugi
3000274 | 2023-07-26 | RSUD KRT. Setjonegoro Wonosobo | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0302 | Survey | Cek supply tegangan, 3 unit fuse 60A 600V putus, board rectifier terbakar | Pending | Sudah diajukan penawaran harga | Imam
19000616 | 2023-07-25 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | Nilai QC tidak masuk arrange; Setting dan adjustment | Ready to use |  | Nanda
7000441 | 2023-07-25 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette C2000 | HM31340 | Survey | Nilai kelembapan tidak sesuai; Kerusakan pada part kit Evaporizer reservoir | Pending | Disarankan penggantian part kit vaporizer reservoir | Wahyu
16000219 | 2023-07-25 | RSUD Karawang | Dental Unit | Kavo/1058TM | 10205314 | Survey | Waterblock mampet, handpiece tidak keluar | Pending | Disarankan penggantian waterblock set dan selang handpiece. Sudah diajukan penawaran harga | Ipan
7000440 | 2023-07-21 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001205 | Maintenance | Preventive Maintenance | Ready to use |  | Wahyu
17000164 | 2023-07-21 | RSUD Berkah Pandeglang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0497 | Service | Tidak bisa mengirim gambar ke pacs; cek IP Address, kabel LAN dan server | Ready to use |  | Bahar
7000439 | 2023-07-21 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001197 | Service | AP1 button tidak berfungsi; Control panel trouble | Ready to use | Proses claim warranty control panel. Sudah dilakukan perbaikan claim warranty | Wahyu
10000166 | 2023-07-21 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | Kolimator tidak terbuka. Main board kolimator terbakar memutus jalur jalur isomotor | Pending | Disarankan penggantian kolimator | Heru
19000614 | 2023-07-21 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Caleo | ASKL0013 | Service | Air bocor; Pengecekan selenoid valve | Ready to use |  | Nanda
19000615 | 2023-07-21 | RSUD Dr. Moewardi | Ventilator | Draeger/Babylog VN500 | ASEF0142 | Survey | Pengecekan unit | Ready to use | Alat bisa digunakan namun aksesoris tidak lengkap (breathing circuit dispo dan reusable, flow sensor habis) | Nanda
19000613 | 2023-07-20 | RSUD Tidar Magelang | Incubator Transport | Draeger/Isolette TI500 | DSNF0010 | Service | Baterai tidak bisa di charge; Tegangan pada baterai drop | Ready to use | Sementara bisa digunakan namun disarankan mengganti baterai internal | Nanda
18000331 | 2023-07-19 | RSUD Kotamobagu | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0270 | Service | Unit tidak bisa expose untuk mode fluoroscopy; Upgrade ssd dan reload software | Ready to use |  | Faugi
17000163 | 2023-07-18 | RSUD Dr. Drajat Prawiranegara | Trolly instrument | Syastech/SCT1 | 1000400870 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2024-07-18
17000163 | 2023-07-18 | RSUD Dr. Drajat Prawiranegara | Emergency trolly | Syastech/SCE2 | 1000400295 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2024-07-18
17000163 | 2023-07-18 | RSUD Dr. Drajat Prawiranegara | Mayo Stand | Syastech/SCY1 | 3000400304 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2024-07-18
2000169 | 2023-07-18 | RSUP Fatmawati | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0493 | Service | Hasil gambar gelap, reboot detector | Ready to use |  | Eddy
8000546 | 2023-07-18 | RSI Aisyah Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0048 | Service | Penggantian O2 sensor | Ready to use |  | Amna
19000612 | 2023-07-17 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHH0122 | Survey | Unit mati; baterai drop | Pending | Disarankan service set include baterai | Nanda
5000185 | 2023-07-17 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Service | Air pada applicator kempes. Pengecekan udara dan pengisian air pada applicator | Ready to use |  | Bilal
19000611 | 2023-07-14 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0124 | Service | Penggantian valve a-cone | Ready to use |  | Nanda
19000610 | 2023-07-13 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Maintenance | Preventive Maintenance Warranty | Ready to use |  | Nanda
6000523 | 2023-07-13 | RSUD Dr. Soeselo Slawi | Mobile X-ray | Apelem/Rafale | 1906142200001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2025-07-13
16000218 | 2023-07-13 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Service | Penggantian board scaller | Ready to use |  | Ipan
7000438 | 2023-07-12 | RSUP dr. Soeradji Tirtonegoro | Infant Warmer | Draeger/Babytherm 8010 | ASMM0045 | Survey | Error 35. Heater rusak | Pending | Disarankan penggantian heater | Wahyu
16000217 | 2023-07-12 | RSUD Karawang | Dental Unit | Kavo/1058 TM | 10205314 | Survey | Pengecekan tekanan air, angin, dan valve sensor air | Pending | Disarankan penggantian valve | Ipan
2000168 | 2023-07-11 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Maintenance | Preventive Maintenance | Ready to use |  | Eddy
3000273 | 2023-07-11 | RSK Dharmais | Digital Radiografi Fluoroskopi | Apelem/Platinum | L216F0664 | Survey | Monitor console mati. Joystick board bermasalah | Ready to use | Pengajuan claim warranty. Sudah dilakukan perbaikan claim warranty | Imam
19000608 | 2023-07-10 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Service | Fan failure; Ada fan mati bagian box mixer | Ready to use |  | Nanda
7000437 | 2023-07-07 | RSUP Dr. Sardjito | Infant Warmer | Draeger/Resuscitaire | DSSE0017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
4000409 | 2023-07-07 | RSUD ODSK Provinsi Sulawesi Utara | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/130/C2 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2025-07-07
19000606 | 2023-07-07 | RSUP dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0154 | Service | Ventilator piston failed; Penggantian karet piston dan motor ventilator | Ready to use |  | Nanda
8000550 | 2023-07-06 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0033 | Service | Penggantian motor pump, ILCA 3, dan flow power | Ready to use |  | Amna
19000605 | 2023-07-05 | RSUP dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0305 | Survey | Battery internal low | Pending | Perlu penggantian service set 2 years include battery | Nanda
7000436 | 2023-07-05 | RSGM UGM | Mesin Anasthesia | Draeger/Primus | ASMA0129 | Survey | Nilai EtCO2 fluktuatif karna menggunakan breathing circuit pcd bukan neo | Ready to use |  | Wahyu
2000167 | 2023-07-05 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Service | Penggantian kolimator | Ready to use |  | Eddy
19000604 | 2023-07-04 | RSUD Dr. Soehadi Prijonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASKE0187 | Service | Penggantian board | Ready to use |  | Nanda
7000435 | 2023-07-04 | RSUD Sleman | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0128 | Survey | System failure; Selenoid P. Bag valve rusak | Pending | Disarankan penggantian part Selenoid P. Bag valve | Wahyu
18000489 | 2023-07-04 | RSUD Maria Walanda Maramis | Mesin Anasthesia | Draeger/Primus | ASMA0145 | Survey | Block anasthesi diduga overheat | Ready to use |  | Faugi
7000406 | 2023-06-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0026 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
15000233 | 2023-06-27 | RSUD Harapan dan Doa | X-ray Digital | Apelem/Camargue DR | CPD33564E19 | Service | Kolimator goyang, selongsong nyangkut, pemegang tombol goyang; Sudah dilakukan perbaikan | Ready to use |  | Sutan
16000216 | 2023-06-27 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Service | Scaller kadang nyala kadang mati. Penggantian board foot switch | Ready to use |  | Ipan
6000522 | 2023-06-27 | RSI Banjarnegara | Ventilator | Draeger/Savina 300 | ASSE0141 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000522 | 2023-06-27 | RSI Banjarnegara | Ventilator | Draeger/Savina 300 | ASSE0142 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000522 | 2023-06-27 | RSI Banjarnegara | Ventilator | Draeger/Savina 300 | ASSE0143 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000522 | 2023-06-27 | RSI Banjarnegara | Ventilator | Draeger/Savina 300 | ASSE0144 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
19000602 | 2023-06-27 | RS Nasional Diponegoro | Infant Warmer | Draeger/Resuscitaire | AF10307 | Service | Tombol on gas bocor; Penggantian valve control mini 5 way | Ready to use |  | Nanda
19000603 | 2023-06-27 | RSUP dr. Kariadi | Baby Incubator | Draeger/Isolette C8000 | MR07739 | Service | Pengecekan water supply chamber, ada kemungkinan rusak pada heater humidity | Pending | Disarankan penggantian heater humidity | Nanda
19000601 | 2023-06-26 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Survey | Pressure sensor rusak | Pending |  | Nanda
7000434 | 2023-06-26 | RSUD Bagas Waras | Infant Warmer | Draeger/Resuscitaire | JY03386 | Service | Error 05. Pemasangan heater element | Ready to use |  | Wahyu
18000484 | 2023-06-23 | RSUD Tobelo | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/088/C1 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi
7000430 | 2023-06-23 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
7000431 | 2023-06-23 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
8000549 | 2023-06-23 | Puskesmas Kedung Mundu | UPS | Yoshiga/HPT03 | 310023219CB73 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2024-06-23
3000272 | 2023-06-22 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | Counter shock loncat-loncat, penggantian board triger HVG | Ready to use |  | Imam
19000600 | 2023-06-21 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000599 | 2023-06-21 | RS JIH | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0125 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
8000548 | 2023-06-21 | RS PKU Muhammadiyah Jepara | Mesin Anasthesia | Draeger/Fabius Plus | ASND0058 | Maintenance | General check. Running test. O2 sensor fail | Ready to use | Disarankan penggantian O2 sensor | Amna
3000271 | 2023-06-21 | RSUD Kab. Temanggung | X-ray Digital | Apelem/Camargue DR | IMA703922 | Service | Penggantian tensator spring | Ready to use |  | Imam
18000326 | 2023-06-20 | RSUD Toto Kabila | Infant Warmer | Draeger/Resuscitaire | 2125850 | Service | Penggantian heater | Ready to use |  | Faugi
17000162 | 2023-06-20 | RSUD Dr. Adjidarmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Service | Grid macet, software crash dan lambat. Perbaikan pintu grid | Ready to use |  | Bahar
7000429 | 2023-06-20 | RSA UGM | Mesin Anasthesia | Draeger/Primus | ASMA0129 | Survey | Pendampingan alat dengan dokter | Ready to use |  | Wahyu
10000164 | 2023-06-20 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Komputer workstation tidak menyala. Cek kabel generator ke workstation | Ready to use |  | Heru
3000270 | 2023-06-20 | RSUD Tidar Magelang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0496 | Service | Penggantian board electronic BD | Ready to use |  | Imam
18000325 | 2023-06-19 | RSUD dr. Zainal Umar Sidiki | Ventilator | Draeger/Babylog VN800 | ASSD0008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi
19000596 | 2023-06-19 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0180 | Survey | Breathing system low. Cleaning dan adjustment | Ready to use | Disarankan penggantian waterlock dan sample line | Nanda
19000598 | 2023-06-19 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0016 | Survey | kalibrasi, cleaning dan adjustment flow sensor | Ready to use |  | Nanda
19000597 | 2023-06-19 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0129 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
7000428 | 2023-06-19 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Service | Nilai keluaran vaporizer tidak sesuai. Lakukan adjustment pada vaporizer | Ready to use |  | Wahyu
7000427 | 2023-06-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
15000224 | 2023-06-17 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Tombol joystick tidak sesuai arah; Penggantian push button joystick | Ready to use |  | Sutan
7000425 | 2023-06-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHL0659 | Service | Nilai keluaran vaporizer tidak sesuai. Lakukan adjustment pada vaporizer | Ready to use |  | Wahyu
7000426 | 2023-06-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Service | Nilai keluaran vaporizer tidak sesuai. Lakukan adjustment pada vaporizer | Ready to use |  | Wahyu
7000424 | 2023-06-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Service | Nilai keluaran vaporizer tidak sesuai. Lakukan adjustment pada vaporizer | Ready to use |  | Wahyu
19000595 | 2023-06-16 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0484 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000595 | 2023-06-16 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0483 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
10000162 | 2023-06-16 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Server NB RS232 communication. Reset generator | Ready to use |  | Heru
3000269 | 2023-06-16 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helianthus DBT | HELI/0023/CB0 | Service | Image artefak, melakukan kalibrasi detector | Ready to use |  | Imam
18000485 | 2023-06-15 | RS Bunda Gorontalo | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0042 | Service | Penggantian O2 sensor | Ready to use |  | Faugi
2000166 | 2023-06-15 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Service | Error table random. Update firmware, update axis controller, update firmware driver motor | Ready to use |  | Eddy
8000547 | 2023-06-15 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | Survey | General check | Ready to use | Disarankan penggantian O2 sensor dan service set 3 tahun | Amna
8000547 | 2023-06-15 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Survey | General check | Ready to use | Disarankan penggantian O2 sensor dan service set 3 tahun | Amna
7000423 | 2023-06-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
4000408 | 2023-06-14 | RSUD Selaparang | Mobile X-ray | Apelem/Rafale EV30 | 1876552200001 | Survey | Uji kesesuaian dan uji paparan | Ready to use |  | Bayu
5000184 | 2023-06-14 | RS Dr. Suyoto | Dental Unit | Kavo/1058TM | 10207194 | Service | Suction tidak menyedot maksimal. Pengecekan jalur selang suction | Ready to use |  | Bilal
19000594 | 2023-06-14 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Maintenance | Preventive maintenance | Ready to use |  | Nanda
18000327 | 2023-06-14 | RSUD Datoe Binangkang | Mesin Anasthesia | Draeger/Fabius | ASHA0165 | Service | Penggantian O2 sensor | Ready to use |  | Faugi
13000158 | 2023-06-13 | RSUD Djasamen Saragih | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/137/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Sofian
70000421 | 2023-06-13 | RSUD Bagas Waras | Infant Warmer | Draeger/Resuscitaire | JY03386 | Survey | Error 05. Heater rusak | Pending | Disarankan penggantian Heater | Wahyu
7000422 | 2023-06-13 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette C8000 | AC09111 | Survey | Skin sensor tidak respon dan engsel patah | Pending | Disarankan penggantian skin sensor, engsel bagian belakang kanan, depan kanan dan depan kiri | Wahyu
7000418 | 2023-06-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHL0059 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
7000419 | 2023-06-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
17000160 | 2023-06-12 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Service | Scaller tidak berfungsi. Pengencengan karet scaller | Ready to use |  | Bahar
8000545 | 2023-06-12 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 edition | ASEL0143 | Service | Neoflow and O2 sensor calibration fail. Penggantian neoflow dan O2 sensor | Ready to use |  | Amna
8000545 | 2023-06-12 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 edition | ASEL0142 | Service | Neoflow and O2 sensor calibration fail. Penggantian neoflow dan O2 sensor | Ready to use |  | Amna
7000420 | 2023-06-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Service | Nilai keluaran vaporizer tidak sesuai. Lakukan adjustment pada vaporizer | Ready to use |  | Wahyu
19000593 | 2023-06-10 | RSUP dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0154 | Survey | Motor piston fail | Pending | Disarankan penggantian piston motor dan service set 6 tahun | Nanda
6000519 | 2023-06-10 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASHM0015 | Service | Pemasangan part PBA CPU | Ready to use |  | Afrizal
18000483 | 2023-06-10 | RSUD Toto Kabila | X-ray Digital | Apelem/Camargue DR | IMA703943 | Service | Hasil gambar ekspose berwarna putih; Dilakukan resert diduga alat overheat | Ready to use |  | Faugi
7000414 | 2023-06-09 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
7000410 | 2023-06-09 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Babyleo TI500 | ASMJ0009 | Survey | Kelembapan terlalu rendah. Lakukan pembersihan humidifier | Ready to use |  | Wahyu
10000163 | 2023-06-09 | RSUD Sultan Suriansyah | X-ray Digital | Apelem/Camargue DR | CPD33483019 | Service | Perakitan ulang workstation apelem | Ready to use |  | Heru
3000268 | 2023-06-09 | RSUD Dr. Drajat Prawiranegara | X-ray Digital | Apelem/Camargue DR | 2510220562 | Service | Low MA fault. Cek kabel HTT generator, cek kabel HTT tube, kalibrasi tube, dan rubah kelistrikan kabel ceilling | Ready to use |  | Imam
18000480 | 2023-06-08 | RSUD Dr. Iwan Boking | Mesin Anasthesia | Draeger/Fabius Plus | ASKE0185 | Service | Penggantian selang oksigen | Ready to use |  | Faugi
7000415 | 2023-06-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0007 | Service | Kelembapan terlalu rendah. Lakukan pembersihan humidifier | Ready to use |  | Wahyu
7000417 | 2023-06-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Service | Nilai keluaran vaporizer tidak sesuai. Lakukan adjustment pada vaporizer | Ready to use |  | Wahyu
7000409 | 2023-06-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMH0118 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
7000416 | 2023-06-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Service | Nilai keluaran vaporizer tidak sesuai. Lakukan adjustment pada vaporizer | Ready to use |  | Wahyu
15000222 | 2023-06-07 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Survey | CPU workstation berbunyi; CPU dikirim ke kantor pusat untuk di cek | Pending | Sudah diberikan penawaran | Sutan
19000592 | 2023-06-07 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASJC0044 | Survey | O2 calibration failed | Pending | Disarankan penggantian sensor PATO | Nanda
7000408 | 2023-06-06 | RSUP dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0109 | Service | Pemasangan service set 3 tahunan | Ready to use |  | Wahyu
7000411 | 2023-06-05 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
7000413 | 2023-06-05 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
7000412 | 2023-06-05 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
19000591 | 2023-06-05 | RSUP dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0154 | Service | Piston fail, kemungkinan rusak pada motor dan pcb vgc power | Pending |  | Nanda
19000590 | 2023-06-05 | RSUD Tidar Magelang | Ventilator | Draeger/Evita V800 | ASRL0090 | Survey | gas unit supply failed, kemungkinan rusak pada interface cable | Pending |  | Nanda
5000183 | 2023-06-05 | RS Abdi Waluyo | EECP | Vasomedical/Lumenair | 5TC1034 | Service | selenoid tidak berfungsi, dilakukan penggantian selenoid valve | Ready to use |  | Bilal
6000518 | 2023-06-01 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHM0013 | Survey | Sensor check merah | Pending | Disarankan penggantian VGC power board, service set 6 tahun | Afrizal
8000544 | 2023-05-31 | Rumah Sakit St. Elisabeth | Lampu Operasi | Draeger/Polaris 600 | ASRK1018 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000543 | 2023-05-30 | RS Keluarga Sehat Tayu | UPS | Yoshiga/HPT-2 | 501101236060KA | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2024-05-30
19000589 | 2023-05-30 | RSUD RAA Soewondo Pati | Ventilator | Draeger/Savina 300 | ASEK0197 | Service | exp valve fail, blower off | Pending | Kemungkinan ada kerusakan pada blower | Nanda
5000182 | 2023-05-30 | RS Abdi Waluyo | EECP | Vasomedical/Lumenair | 5TC1034 | Service | Manset tidak mengembang, selenoid valve rusak | Pending |  | Bilal
7000407 | 2023-05-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
3000267 | 2023-05-30 | RSUD Sumedang | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/132/C2 | Survey | Training alat | Ready to use |  | Imam
18000323 | 2023-05-29 | RSUD Toto Kabila | Dental Unit | Kavo/1058TM | 205320 | Service | Penggantian main board | Ready to use |  | Faugi
16000215 | 2023-05-29 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1086410 | Service | Handpiece rusak | Pending | Disarankan penggantian bearing | Ipan
2000165 | 2023-05-29 | RSUP Fatmawati | Digital Radiografi Fluoroskopi | Apelem/Platinum | J202F0493 | Service | Upgrade software | Ready to use |  | Eddy
6000517 | 2023-05-26 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0011 | Service | Pemasangan part battery pack | Ready to use |  | Afrizal
3000266 | 2023-05-26 | RSUD Gunung Jati Cirebon | X-ray Digital | Apelem/Camargue DR | IMA704014 | Service | Cathoda MA Fault, kerusakan pada board inverter | Pending | Sudah diajukan penawaran harga | Imam
18000322 | 2023-05-25 | RSAL dr. Wahyu Bitung | Dental Unit | Kavo/1058TM | 20008879 | Service | Suction pump tidak berfungsi; Pengecekan motor suction dan penggantian fuse | Ready to use |  | Faugi
19000588 | 2023-05-25 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Survey | Video ultrasound blur | Ready to use | S-video card diambil untuk ditest. Sudah dilakukan penambahan monitor | Nanda
11000227 | 2023-05-25 | RSAU Dr. Dody Sarjoto | Baby Incubator | Draeger/Babyleo TI500 | ASLL0008 | Service | Skin temperature, humidity dan air temp. error | Pending | Disarankan penggantian board PBA sensor | Nasrul
3000265 | 2023-05-25 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Maintenance | Preventive Maintenance | Ready to use |  | Imam
18000321 | 2023-05-23 | RS Mata Sulawesi Utara | Mesin Anasthesia | Draeger/Perseus | ASJM0178 | Maintenance | Preventive Care Total Care | Ready to use |  | Faugi
11000226 | 2023-05-23 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0395 | Service | Battery internal fault | Pending | Disarankan penggantian battery internal | Nasrul
11000226 | 2023-05-23 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0392 | Service | Battery internal fault | Pending | Disarankan penggantian battery internal | Nasrul
11000226 | 2023-05-23 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0036 | Service | Battery internal fault | Pending | Disarankan penggantian battery internal | Nasrul
11000226 | 2023-05-23 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0138 | Service | Battery internal fault | Pending | Disarankan penggantian battery internal | Nasrul
11000226 | 2023-05-23 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0396 | Service | Battery internal fault | Pending | Disarankan penggantian battery internal | Nasrul
19000587 | 2023-05-23 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Isolette C8000 | MC09026 | Survey | Humidity low, water empty low. Cleaning system | Ready to use |  | Nanda
7000405 | 2023-05-23 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | NF31848 | Survey | Sensor Disconnected. Kerusakan pada sensor modul | Pending |  | Wahyu
19000586 | 2023-05-22 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASJC0044 | Survey | Device check incomplete. PATO rusak | Pending | Disarankan penggantian sensor PATO dan maintenance kit dan baterai | Nanda
18000475 | 2023-05-21 | RSUD Ainun Habibie | Ventilator | Draeger/Evita V300 | ASLD0172 | Survey | Edukasi user | Ready to use |  | Irfan R
11000225 | 2023-05-20 | RSUD Tenriawaru Bone | X-ray Digital | Apelem/Camargue DR | CPD26084L16 | Service | Komputer tidak booting. | Ready to use | Disarankan untuk upgrade ssd dan reload software. Sudah diajukan penawaran harga | Nasrul
19000585 | 2023-05-19 | RSUP dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0154 | Service | Piston fail, breathing system fail. Cleaning membran piston | Ready to use | Disarankan penggantian service set | Nanda
11000224 | 2023-05-19 | RSUD LA Temmamala | Ventilator | Draeger/Savina 300 | ASMA0116 | Service | O2 failed, internal batterai failed | Pending | Disarankan penggantian service set 3 tahun | Nasrul
11000223 | 2023-05-19 | RSUD LA Temmamala | Ventilator | Draeger/Babylog VN500 | ASML0095 | Service | Air supply failed | Pending | Disarankan maintenance kit dan flow sensor | Nasrul
18000319 | 2023-05-19 | RSAL dr. Wahyu Bitung | Dental Unit | Kavo/1058TM | 20008879 | Service | Update software | Ready to use | Alat bisa digunakan namun suction motor mati | Faugi
19000584 | 2023-05-17 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Maintenance | setting adjustemnt | Ready to use |  | Nanda
7000402 | 2023-05-17 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASHN0040 | Service | Piezo alarm. Power supply fail. Penggantian electrical emergency unit, battery pack, pressure regulator, dan kit proportional valve. | Ready to use |  | Wahyu
70000403 | 2023-05-17 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKE0070 | Survey | Check nebulizer failed | Pending | Disarankan penggantian pressure regulator | Wahyu
8000542 | 2023-05-16 | RS Aisyiyah Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0048 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian O2 sensor | Amna
3000264 | 2023-05-16 | RSUD Provinsi NTB | Digital Radiografi Fluoroskopi | Apelem/Platinum | B196F0329 | Service | Penggantian fuse 60A 600V | Ready to use | Terdapat kerusakan input distribution board, fan 3,4,5, dan kerusakan software workstation duet. Sudah diajukan penawaran harga | Imam
19000583 | 2023-05-15 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | Air pada applicator kosong. Setting water and adjusment, penggantian busi | Ready to use |  | Nanda
6000516 | 2023-05-14 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASNB0134 | Service | Pemasangan demo unit | Ready to use |  | Afrizal
6000515 | 2023-05-12 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0025 | Service | Pemasangan O2 sensor | Ready to use |  | Afrizal
6000515 | 2023-05-12 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASBM0248 | Survey | Perlu penggantian O2 sensor | Ready to use |  | Afrizal
6000515 | 2023-05-12 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0039 | Survey | Perlu penggantian O2 sensor | Ready to use |  | Afrizal
6000515 | 2023-05-12 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0047 | Survey | Perlu penggantian O2 sensor | Ready to use |  | Afrizal
6000515 | 2023-05-12 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0027 | Survey | Perlu penggantian O2 sensor | Ready to use |  | Afrizal
6000515 | 2023-05-12 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASKJ0082 | Survey | Perlu penggantian O2 sensor | Ready to use |  | Afrizal
6000514 | 2023-05-11 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | AF32541 | Service | Pemasangan MU13393, MU12695, MU12504 | Ready to use |  | Afrizal
6000514 | 2023-05-11 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C8000 | EH08777 | Service | Pemasangan MU21361, MU13393, dan MU12504 | Ready to use |  | Afrizal
6000514 | 2023-05-11 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C8000 | EH08779 | Service | Pemasangan MU21361 dan MU13393 | Ready to use |  | Afrizal
6000514 | 2023-05-11 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | DSNSF0062 | Service | Pemasangan MU13393 & MU12504 | Ready to use |  | Afrizal
6000514 | 2023-05-11 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C8000 | EH08778 | Survey | Perlu penggantian MU19804, MU21361 | Pending |  | Afrizal
6000514 | 2023-05-11 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | AF32541 | Survey | Perlu penggantian MU12601 | Pending |  | Afrizal
6000514 | 2023-05-11 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | AF32537 | Survey | Perlu penggantian MU12601, MU12600, MU12602 | Pending |  | Afrizal
8000541 | 2023-05-11 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASSC0025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
3000263 | 2023-05-11 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Service | Penggantian EMH | Ready to use |  | Imam
7000401 | 2023-05-10 | RSA UGM | Mesin Anasthesia | Draeger/Primus | ASMA0129 | Survey | Pembacaan EtCo2 fluktuatif. Cek mPGm module, set vacuum pump | Ready to use |  | Wahyu
11000300 | 2023-05-10 | RSUD Kota Makassar | Mesin Anasthesia | Draeger/Perseus A500 | ASJD0033 | Service | Gas monitoring failure | Pending | Disarankan penggantian ILCA | Nasrul
11000299 | 2023-05-10 | RSUD Kota Makassar | Baby Incubator | Draeger/Caleo | ASJF0032 | Service | Kondisi motor hidrolik rusak | Pending | Disarankan penggantian hidrolik | Nasrul
18000318 | 2023-05-05 | RSUD Datoe Binangkang | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0165 | Service | Penggantian selang oksigen | Ready to use |  | Faugi
17000161 | 2023-05-05 | Universitas Prof. Dr. Moestopo | Dental Simulator | Kavo/Compact | 0000085 | Survey | Handle lampu patah | Pending | Sudah diajukan penawaran untuk penggantian handle | Bahar
2000164 | 2023-05-05 | RSUD Labuang Baji | Digital Radiografi Fluoroskopi | Apelem/Platinum | J20F0494 | Maintenance | Preventive Maintenance | Ready to use |  | Eddy
13000157 | 2023-05-04 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Service | Setiap menyalakan unit harus atur ulang tanggal; Ganti baterai CMOS | Ready to use |  | Sofian
3000262 | 2023-05-03 | RSUD Kab. Bolaang Mongondow | X-ray Digital | Apelem/Camargue DR | CPD30503B18 | Service | Upgrade SSD 2TB, penggantian WAP Cisco, Reload software davinci | Ready to use |  | Imam
7000400 | 2023-05-02 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKE0069 | Service | Penggantian part battery internal dan pressure regulator | Ready to use |  | Wahyu
11000229 | 2023-05-02 | RSIA Sentosa | Ventilator | Draeger/Babylog VN600 | ASRN0085 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
7000399 | 2023-05-02 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Babylog VN500 | ASHJ0032 | Service | Penggantian battery pack internal dan pressure regulator | Ready to use |  | Wahyu
8000540 | 2023-04-28 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASHD0034 | Service | Leakage and VT low. General check, device check | Ready to use |  | Amna
6000513 | 2023-04-28 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0011 | Survey | Selftest gagal | Pending | Disarankan penggantian sensor PDMIX,PDMGS, Baterai internal | Afrizal
1000161 | 2023-04-27 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | console tidak berfungsi, dilakukan pembersihan board dan pemindahan kabel usb console | Ready to use |  | Heru
13000156 | 2023-04-26 | RSU Haji Medan | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/033/C0 | Service | Tidak konek dengan printer dan komputer dokter; Penggantian switch hub dan konektor LAN | Ready to use |  | Sofian
3000261 | 2023-04-26 | RSUD Dr. Soediran Mangun Sumarso | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/134/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Imam
18000459 | 2023-04-25 | RS Bunda Gorontalo | Mesin Anasthesia | Draeger/Fabius Plus | USNF0013 | Service | Penggantian O2 sensor | Ready to use |  | Faugi
6000512 | 2023-04-18 | RS PKU Purbalingga | Mesin Anasthesia | Draeger/Fabius | ARZM0025 | Service | Flow meter bocor. Penggantian selang | Ready to use |  | Afrizal
4000405 | 2023-04-18 | RSUD Kab. Bekasi | C-arm | Apelem/Evo R+ | 55116026649 | Survey | Board CCD Camera II dibawa ke principle untuk pengecekan lebih lanjut | Ready to use |  | Bayu
7000398 | 2023-04-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKK0069 | Survey | Pengecekan BHP | Ready to use | Disarankan penambahan stock flow sensor, sample line dan waterlock | Wahyu
7000397 | 2023-04-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0077 | Survey | Pengecekan BHP | Ready to use | Disarankan penambahan stock flow sensor, sample line dan waterlock | Wahyu
7000396 | 2023-04-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Survey | Pengecekan BHP | Ready to use | Disarankan penambahan stock flow sensor, sample line dan waterlock | Wahyu
7000395 | 2023-04-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMH0118 | Survey | Pengecekan BHP | Ready to use | Disarankan penambahan stock flow sensor, sample line dan waterlock | Wahyu
7000394 | 2023-04-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0007 | Survey | Pengecekan BHP | Ready to use | Disarankan penambahan stock flow sensor, sample line dan waterlock | Wahyu
7000393 | 2023-04-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Survey | Pengecekan BHP | Ready to use | Disarankan penambahan stock flow sensor, sample line dan waterlock | Wahyu
7000392 | 2023-04-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Survey | Pengecekan BHP | Ready to use | Disarankan penambahan stock flow sensor, sample line dan waterlock | Wahyu
19000580 | 2023-04-18 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHA01140 | Service | Valve Cozy dan jackson rease mampet. Setting ulang | Ready to use |  | Nanda
8000539 | 2023-04-18 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0039 | Service | Sensor check fail. Penggantian PDIFF | Ready to use |  | Amna
6000511 | 2023-04-17 | RSUD Dr. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette | AH36595 | Service | Pemasangan U12601 hinge access PNL (R+L) | Ready to use |  | Afrizal
6000511 | 2023-04-17 | RSUD Dr. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette | AH36597 | Service | Pemasangan MU12600 & MU12695 dan Filter isollete | Ready to use |  | Afrizal
19000579 | 2023-04-17 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Service | Shock loncat. Penggantian HVG dan power supply | Ready to use |  | Nanda
6000510 | 2023-04-14 | RSIA Harapan Ibu | Bubble CPAP | Draeger/Bubble CPAP | QM20221069 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000510 | 2023-04-14 | RSIA Harapan Ibu | Bubble CPAP | Draeger/Bubble CPAP | QM20220156 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000510 | 2023-04-14 | RSIA Harapan Ibu | Bubble CPAP | Draeger/Bubble CPAP | QM20220163 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
19000578 | 2023-04-14 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0094 | Service | Kalibrasi internal | Ready to use |  | Nanda
18000315 | 2023-04-13 | RSUD Tobelo | X-Ray Analog | Apelem/Camargue Analog | IMA703799 | Service | Penggantian lampu koliamator board | Ready to use |  | Faugi
4000404 | 2023-04-13 | RSUD R. Syamsudin | X-ray Digital | Samsung/ X-Geo | 502SM3A1500001R | Survey | Uji kesesuaian dan uji paparan | Ready to use |  | Bayu
19000577 | 2023-04-13 | RSUD dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0150 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000576 | 2023-04-12 | RS Nasional Diponegoro | Infant Warmer | Draeger/Resuscitaire | AF10306 | Service | Error 05. Heater sementara disambung | Ready to use | Alat masih bisa berfungsi namun tetap disarankan penggantian heater assy | Nanda
13000154 | 2023-04-12 | RSUD Simeulue | Bone Mineral Densitometer | Medilink/Medix DR | F207H0702 | Survey | Training alat | Ready to use |  | Sofian
11000298 | 2023-04-12 | RSUD I Lagaligo | Ventilator | Draeger/Babylog VN500 | ASKJ0005 | Maintenance | Preventive Maintenance (Total care) | ready to use |  | Nasrul
7000391 | 2023-04-11 | RSUD Dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Atlan A350 | ASRL0046 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000390 | 2023-04-11 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0026 | Survey | Pengecekan BHP | Ready to use | Disarankan penambahan stock flow sensor, sample line dan waterlock | Wahyu
3000260 | 2023-04-11 | RSUD Tanjung Priuk | UPS | Yoshiga/HPR 6 | 501101310220M930004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2024-04-11
3000259 | 2023-04-11 | RSUD Koja | UPS | Yoshiga/HPR 2 | 501101236060KA100044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2024-04-11
18000314 | 2023-04-10 | RSD Tidore | X-ray Digital | Apelem/Camargue DR | 2510220501 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi | 2025-04-10
4000403 | 2023-04-10 | RSUD Kab. Bekasi | C-arm | Apelem/Evo R+ | 55116026649 | Service | X-ray tidak keluar, terdapat kerusakan pada board CCD Caera dan Gige camera | Ready to use |  | Bayu
7000389 | 2023-04-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Survey | Pengecekan BHP | Ready to use | Disarankan penambahan stock flow sensor, sample line dan waterlock | Wahyu
7000387 | 2023-04-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Survey | Pengecekan BHP | Ready to use | Disarankan penambahan stock flow sensor, sample line dan waterlock | Wahyu
7000388 | 2023-04-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHL0059 | Survey | Pengecekan BHP | Ready to use | Disarankan penambahan stock flow sensor, sample line dan waterlock | Wahyu
2000161 | 2023-04-06 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Service | Perbaikan joystick colum | Ready to use |  | Eddy
8000538 | 2023-04-06 | RSUP dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0190 | Service | Preventive Maintenance. Penggantian service set 2 tahun | Ready to use |  | Amna
6000534 | 2023-04-05 | RSUD Bumiayu | Jaundice Meter | Draeger/JM 105 | B3503805 | Survey | Error 05 dan muncul time periodic calibration | Pending | Disarankan adjustment dan kalibrasi dan penggantian baterai | Afrizal
7000386 | 2023-04-05 | RSJD Dr. RM. Soedjarwadi | Baby Incubator | Draeger/Caleo | ASJM0054 | Service | Cover display pecah. Penggantian case monitor controlled | Ready to use |  | Wahyu
6000534 | 2023-04-05 | RSUD Bumiayu | Infant Warmer | Draeger/Resuscitaire | CY11224 | Survey | Error 05 | Pending | Disarankan penggantian heater | Afrizal
18000317 | 2023-04-04 | RS Mata Sulawesi Utara | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0021 | Service | Penggantian battery | Ready to use |  | Faugi
19000582 | 2023-04-04 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0044 | Service | EtCo2 belum bisa keluar | Pending | Disarankan penggantian waterlock dan sensor ILCA | Nanda
15000221 | 2023-04-04 | RSUD Hasanuddin Damrah | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/138/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Sutan
14000220 | 2023-04-04 | RSUD Syarifah Ambami Rato Ebu | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/055/C1 | Maintenance | Preventive Maintenance | Ready to use |  | Purwono
8000537 | 2023-04-04 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina 300 | ASPJ0380 | Service | Alarm FIO2 measurement fail. Kalibrasi O2 | Ready to use |  | Amna
8000537 | 2023-04-04 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina 300 | ASPJ0382 | Service | Alarm FIO2 measurement fail. Kalibrasi O2 | Ready to use |  | Amna
8000536 | 2023-04-03 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina 300 | ASSC0038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
6000503 | 2023-04-01 | RSUD Prof. Dr. Margono Soekarjo | Jaundice Meter | Draeger/JM 105 | B3503802 | Service | Penggantian baterai dan adjusment calibration | Ready to use |  | Afrizal
6000504 | 2023-03-31 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASRJ0002 | Service | Pemasangan pressure sensor 140mbar (Claim warranty) | Ready to use |  | Afrizal
19000572 | 2023-03-31 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0044 | Survey | Pembacaan inspiration dan expiration tidak sesuai. | Pending | Disarankan penggantian waterlock dan sample line | Nanda
6000505 | 2023-03-30 | RSUD Majenang | Ventilator | Draeger/Babylog VN800 | ASRL0024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000505 | 2023-03-30 | RSUD Majenang | Ventilator | Draeger/Babylog VN800 | ASRK0010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
19000574 | 2023-03-30 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Survey | Shock kembali lompat-lompat. Kalibrasi KV | Ready to use |  | Nanda
4000402 | 2023-03-29 | RSUD Kota Kendari | ESWL | Medispec/E3000 | 0922003 | Survey | Pelaksanaan uji paparan dan uji kesesuaian | Ready to use |  | Bayu
4000402 | 2023-03-29 | RSUD Kota Kendari | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/127/C2 | Service | Kalibrasi detector, setting kolimasi dan pendampingan uji kesesuaian | Ready to use |  | Bayu
o | 2023-03-29 | RSUP Dr. M. Djamil Padang | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J190100147 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan | 2025-03-29
6000506 | 2023-03-29 | RS Siti Aisyah | Ventilator | Draeger/Savina 300 | ASCC0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
19000571 | 2023-03-28 | RSUD Tidar Magelang | Bone Mineral Densitometer | Medilink/Medix DR | M15016M258 | Service | QC not ready. Setting adjusment | Ready to use | Disarankan upgrade ssd dan reload software | Nanda
6000507 | 2023-03-28 | RSI Harapan Anda | Ventilator | Draeger/Babylog VN800 | ASRL0029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
14000219 | 2023-03-27 | RSUD Dr. R. Sosodoro Djatikoesoemo Bojonegoro | Dental Unit | Kavo/E30 | 10111154 | Service | Unit tidak bisa nyala, sementara main board dan dentist element board dibawa ke HO untuk dicek lebih lanjut | Ready to use |  | Purwono
7000384 | 2023-03-27 | RSA UGM | Mesin Anasthesia | Draeger/Primus | ASMA0129 | Survey | CO2 tidak akurat pembacaannya. Kalibrasi dan ganti water trap | Ready to use |  | Wahyu
19000570 | 2023-03-27 | RSUP dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0094 | Service | Daily check tidak lolos. O2 calibration, penggantian sensor PATO, O2 filter | Ready to use |  | Nanda
19000569 | 2023-03-24 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Service | Energy dan shock tidak sesuai. Penggantian HVG | Ready to use |  | Nanda
8000535 | 2023-03-21 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 edition | ASEL0143 | Service | O2 sensor fail | Pending | Disarankan penggantian O2 sensor | Amna
8000535 | 2023-03-21 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 edition | ASEL0142 | Service | O2 sensor fail | Pending | Disarankan penggantian O2 sensor | Amna
8000535 | 2023-03-21 | RSUD Dr. Loekmono Hadi | Baby Incubator | Draeger/Isolette C8000 | CU09351 | Service | controller fail 14 | Pending | Disarankan penggantian heater assy dan fan impeller | Amna
5000181 | 2023-03-21 | RS Abdi Waluyo | EECP | Vasomedical/Lumenair | STC1034 | Service | Penggantian CPU board dan kabel power | Ready to use |  | Bilal
18000423 | 2023-03-21 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Evita V300 | ASNM0050 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000423 | 2023-03-21 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Evita VN800 | ASRK0009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
17000159 | 2023-03-21 | RSUD Selaparang | Mobile X-ray | Apelem/Rafale EV30 | 1876552200001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2025-03-21
700383 | 2023-03-21 | RSIA Aisyiyah Klaten | Ventilator | Draeger/Babylog VN600 | ASRB0114 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
3000258 | 2023-03-21 | RSUD Alimuddin Umar | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0402 | Service | Unit mati, cek panel listrik terdapat masalah pada kabel 3 phase | Ready to use | Disarankan untuk penarikan kabel baru. Unit berfungsi dengan disambungkan ke panel lain | Imam
18000406 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASHD0006 | Survey | Pressure regulator dan O2 sensor sudah waktunya diganti | Pending | Disarankan penggantian pressure regulator dan oksigen sensor | Irfan R
18000406 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0416 | Survey | Pressure regulator dan O2 sensor sudah waktunya diganti | Pending | Disarankan penggantian pressure regulator dan oksigen sensor | Irfan R
18000403 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0168 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000403 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0413 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000403 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0170 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000403 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0181 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000403 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0411 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000403 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0417 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000403 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0410 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000403 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0172 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000403 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0419 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000403 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0178 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000403 | 2023-03-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0414 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
11000297 | 2023-03-20 | RSUD Hajjah Andi | Mesin Anasthesia | Draeger/Perseus A500 | ASNJ0104 | Service | O2 sensor failure, flow sensor failure. Kalibrasi air flow sensor | Ready to use |  | Nasrul
19000568 | 2023-03-20 | RSUP dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Survey | Shock dan counter tidak sesuai/ lompat-lompat | Ready to use | Akan dijadwalkan tekni pusat untuk dicek lebih lanjut | Nanda
8000534 | 2023-03-20 | RSUD Limpung | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0146 | Survey | Pengecekan unit. | Pending | Disarankan penggantian O2 sensor dan service set 3 tahun | Amna
6000509 | 2023-03-20 | RSUD Majenang | Ventilator | Draeger/Babylog VN500 | ASRK0018 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
3000257 | 2023-03-20 | RSUD Pringsewu | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0756 | Maintenance | Preventive Maintenance | Ready to use |  | Imam
18000401 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0183 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000401 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0177 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000401 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0180 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000401 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0171 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000401 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0182 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000401 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0174 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000401 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0166 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000401 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0169 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000401 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0185 | Survey | pressure regulator, oksigen sensor, dan baterai rusak | Pending | Disarankan penggantian pressure regulator, oksigen sensor, dan baterai | Irfan R
18000404 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0173 | Survey | Baterai dan oksigen sensor bermasalah | Pending | Disarankan penggantian baterai dan oksigen sensor serta service set 1 tahun | Irfan R
18000404 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0415 | Survey | Baterai dan oksigen sensor bermasalah | Pending | Disarankan penggantian baterai dan oksigen sensor serta service set 1 tahun | Irfan R
18000405 | 2023-03-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0167 | Survey | Engine drive, valve mixing dan baterai error | Pending | Disarankan penggantian engine drive, valve mixing, baterai, dan servoce set 1 tahun | Irfan R
3000256 | 2023-03-17 | RSUD Malingping | X-ray Digital | Apalem/Camargue DR | IMA703916 | Service | Upgrade ssd, ddr, penggantian monitor, dan kalibrasi detector | Ready to use |  | Imam
18000412 | 2023-03-16 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | BY33325 | Survey | Unit mati total | Pending | Disarankan penggantian front panel | Irfan R
18000413 | 2023-03-16 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | CM04709 | Survey | Motor dan engsel rusak | Pending | Disarankan service set 3 tahun dan penggantian engsel depan, serta filter | Irfan R
18000408 | 2023-03-16 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | BY3329 | Survey | Sensor modular dan engsel belakang patah | Pending | Disarankan penggantian sensor modular dan engsel. Serta penggantian filter | Irfan R
18000416 | 2023-03-16 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | CR35573 | Survey | Masalah pada hinge hood dan repl kit Assy (patah) | Pending | Disarankan mengganti hinge hood dan repl kit door serta filter | Irfan R
18000416 | 2023-03-16 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | AY34825 | Survey | Masalah pada hinge hood dan repl kit Assy (patah) | Pending | Disarankan mengganti hinge hood dan repl kit door serta filter | Irfan R
18000416 | 2023-03-16 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | AR35153 | Survey | Masalah pada hinge hood dan repl kit Assy (patah) | Pending | Disarankan mengganti hinge hood dan repl kit door serta filter | Irfan R
17000157 | 2023-03-16 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Service | Console dan komputer medecom terendam air hujan, pengecekan general | Ready to use | Alat dapat berfungsi dengan baik kecuali keyboard medecom (rusak) | Bahar
17000158 | 2023-03-16 | RSUD Dr. Drajat Prawiranegara | X-ray Digital | Apelem/DR Camargue | 2510220562 | Maintenance | Preventive Maintenance | Ready to use |  | Bahar
7000382 | 2023-03-16 | RSUD Wonosari | C-arm | Apelem/Evo R+ | 55121060911 | Maintenance | Preventive Maintenance | Ready to use |  | Wahyu
7000385 | 2023-03-16 | RSA UGM | Baby Incubator | Kavo/DSE Expert | 10001204 | Maintenance | Preventive Maintenance | Ready to use |  | Wahyu
19000567 | 2023-03-16 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDK0300 | Survey | Flow meter error. Re-calibration | Ready to use |  | Nanda
18000419 | 2023-03-15 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Fabius Plus | ASFM0035 | Survey | Pressure regulator rusak | Pending | Disarankan penggantian pressure regulator | Irfan R
18000419 | 2023-03-15 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Fabius Plus | ASFM0036 | Survey | Pressure regulator rusak | Pending | Disarankan penggantian pressure regulator | Irfan R
18000419 | 2023-03-15 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Fabius Plus | ASJK0016 | Survey | Pressure regulator rusak | Pending | Disarankan penggantian pressure regulator | Irfan R
3000255 | 2023-03-15 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Survey | Training alat | Ready to use |  | Irfan R
18000407 | 2023-03-15 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0412 | Survey | Regulator, battery, dan O2 sensor rusak | Pending | Disarankan penggantian O2 sensor, pressure regulator dan baterai serta melakukan service set 1 tahun | Irfan R
8000533 | 2023-03-15 | RSUD Kraton Pekalongan | Lampu Operasi | Draeger/Polaris 500 | ASDK0401 | Service | Lampu mati total arm 1. Power supply failure | Pending | Disarankan penggantian power supply unit | Amna
8000533 | 2023-03-15 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASMA0130 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000533 | 2023-03-15 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0014 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
8000533 | 2023-03-15 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0011 | Maintenance | Preventive Maintenance | Ready to use |  | Amna
6000631 | 2023-03-15 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0013 | Survey | Air compressor kemasukan air dan kondensasi gas sentral | Pending | Disarankan penggantian PDMIX dan PDMGS | Afrizal
18000414 | 2023-03-14 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | BY33327 | Survey | Motor dan lcd error. Front panel dan motor rusak | Pending | Disarankan penggantian motor (Service set 3 tahun) dan front panel serta filter | Irfan R
18000415 | 2023-03-14 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | BY33340 | Service | Hinga hood patah | Pending | Disarankan penggantian hinge hood dan filter | Irfan R
18000411 | 2023-03-14 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | KM04800 | Survey | Kerusakan pada sensor modular | Pending | Disarankan penggantian sensor modular dan filter | Irfan R
18000409 | 2023-03-14 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | BY33331 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000409 | 2023-03-14 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | BY333323 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000410 | 2023-03-14 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | CU03934 | Survey | Motor rusak | Pending | Disarankan penggantian service set 3 tahun dan filter | Irfan R
18000410 | 2023-03-14 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | KM0481 | Survey | Motor rusak | Pending | Disarankan penggantian service set 3 tahun dan filter | Irfan R
18000410 | 2023-03-14 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | AR35152 | Survey | Motor rusak | Pending | Disarankan penggantian service set 3 tahun dan filter | Irfan R
18000417 | 2023-03-14 | RSUP Prof. Dr. R. D. Kandou Manado | Infant Warmer | Draeger/Resuscitaire | AY0355 | Survey | Error 05. Heater hangus | Pending | Disarankan penggantian heater | Irfan R
18000417 | 2023-03-14 | RSUP Prof. Dr. R. D. Kandou Manado | Infant Warmer | Draeger/Resuscitaire | JY0967 | Survey | Error 05. Heater hangus | Pending | Disarankan penggantian heater | Irfan R
6000630 | 2023-03-14 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0013 | Survey | Pressure tidak keluar/ macet intermitten | Pending | Disarankan penggantian breathing Vbar 2 tahun, membran ABV, Check valve, NGV Valve, ABV valve dan battery pack | Afrizal
3000254 | 2023-03-14 | RSUD Sumedang | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/132/C2 | Maintenance | Preventive Maintenance | Ready to use |  | Imam
8000532 | 2023-03-11 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina 300 | ASRE0304 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
7000381 | 2023-03-10 | RSUD Bagas Waras | Infant Warmer | Draeger/Resuscitaire | JY03386 | Service | Penggantian part mini valave 5 way | Ready to use |  | Wahyu
18000312 | 2023-03-10 | RSUD Dr. M. M. Dunda | Infant Warmer | Draeger/Babytherm 8010 | ASMJ0091 | Survey | Refresh training | Ready to use |  | Faugi
6000629 | 2023-03-09 | RS Bhakti Asih Brebes | Ventilator | Draeger/Savina 300 | ASHD0011 | Survey | Device check. | Ready to use | Dianjurkan penggantian service set 6 tahun, kit battery external, O2 sensor | Afrizal
8000531 | 2023-03-09 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | ASML0039 | Service | Penggantian baterai | Ready to use |  | Amna
8000531 | 2023-03-09 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | ASML0040 | Service | Penggantian baterai | Ready to use |  | Amna
19000573 | 2023-03-09 | RSUP dr. Kariadi | Baby Incubator | Draeger/Caleo | ASJD0002 | Service | Penggantian spare part converter heater water kit | Ready to use |  | Nanda
7000380 | 2023-03-09 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Maintenance | Preventive Maintenance | Ready to use |  | Wahyu
2000159 | 2023-03-09 | RSUP Prof. Dr. R. D. Kandou Manado | C-arm | Apelem/Evo R+ | 55112079338 | Survey | Software crash, error filament, realtime clock error | Pending | Disarankan penggantian board CSC, filament board, dan HDD+upgrade software | Eddy
2000160 | 2023-03-09 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | J1G6F0184 | Survey | Kolimator tidak bisa buka tutup kolimasi, image detector bergaris | Ready to use | Disarankan penggantian kolimator | Eddy
2000157 | 2023-03-08 | RSUP Fatmawati | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0493 | Maintenance | Preventive Maintenance | Ready to use |  | Eddy
7000377 | 2023-03-08 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASRK0017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000378 | 2023-03-08 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASRK0006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
7000379 | 2023-03-08 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASRK0016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
8000530 | 2023-03-08 | RSJD Dr. Amino Gondohutomo | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0145 | Maintenance | Preventive Maintenance | Ready to use | Disarankan penggantian O2 sensor | Amna
6000628 | 2023-03-07 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0057 | Service | Pemasangan O2 sensor dan flow sensor | Ready to use |  | Afrizal
6000628 | 2023-03-07 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0050 | Service | Pemasangan O2 sensor dan flow sensor | Ready to use |  | Afrizal
6000627 | 2023-03-07 | RSUD Dr. Soeselo Slawi | Ventilator | Draeger/Babylog VN500 | ASJJ0103 | Service | Pemasangan flow sensor, neo flow dan hose heater | Ready to use |  | Afrizal
6000627 | 2023-03-07 | RSUD Dr. Soeselo Slawi | Ventilator | Draeger/Babylog VN500 | ASJJ0104 | Service | Pemasangan flow sensor, neo flow dan hose heater | Ready to use |  | Afrizal
8000529 | 2023-03-07 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Service | Layar blank. Calibration LCD | Ready to use |  | Amna
7000376 | 2023-03-07 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKL0094 | Survey | pembacaan tekanan tidak akurat, baterai gagal | Ready to use | Disarankan penggantian baterai internal | Wahyu
7000375 | 2023-03-06 | RSUD Prambanan | Ventilator | Draeger/Babylog VN800 | ASPF0049 | Maintenance | Preventinve Maintenance (Total Care) | Ready to use |  | Wahyu
7000375 | 2023-03-06 | RSUD Prambanan | Ventilator | Draeger/Babylog VN800 | ASPF0050 | Maintenance | Preventinve Maintenance (Total Care) | Ready to use |  | Wahyu
14000218 | 2023-03-06 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Service | Perbaikan kunci busy dan cover buka tutup busy | Ready to use |  | Purwono
2000158 | 2023-03-02 | RSK Dharmais | Digital Radiografi Fluoroskopi | Apelem/Platinum | J216F0664 | Survey | Remote Platinum dengan principle | Ready to use |  | Eddy
11000295 | 2023-03-02 | RSUD I Lagaligo | ventilator | Draeger/Babylog VN500 | ASKJ0005 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Nasrul
2000160 | 2023-03-02 | RSUD Sultan Suriansyah | X-ray Digital | Apelem/Camargue DR | CPD33483019 | Service | Applikasi slow respon, penghapusan data pasien 2021 | Ready to use |  | Heru
18000418 | 2023-02-28 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0418 | Service | Battery internal fault | Pending | Disarankan penggantian baterai internal | Irfan R
18000418 | 2023-02-28 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASHD0005 | Service | Battery internal fault | Pending | Disarankan penggantian baterai internal | Irfan R
18000420 | 2023-02-28 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASFL0313 | Survey | Power supply error | Pending | Disarankan penggantian power supply dan battery | Irfan R
8000528 | 2023-02-28 | RSU Gunung Sawo | X-ray Digital | Draeger/Polaris 200 | ASNF0033 | Maintenance | Reset lampu | Ready to use |  | Amna
15000217 | 2023-02-27 | RSUP Dr. Mohammad Hoesin | X-ray Digital | Apelem/Camargue DR | CPD36817K20 | Service | Tidak bisa bridging. Melakukan setting komputer | Ready to use |  | Sutan
14000217 | 2023-02-27 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Service | Unit tidak bisa fungsi, pengunci busi patah | Ready to use |  | Purwono
18000421 | 2023-02-27 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASHD0029 | Survey | Error pada sensor PATO | Pending | Disarankan penggantian O2 sensor PATO | Irfan R
18000422 | 2023-02-27 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASHD0074 | Survey | Error pada black PGDM | Pending | Disarankan penggantian ILCA2 | Irfan R
18000422 | 2023-02-27 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASJF0224 | Survey | Error pada black PGDM | Pending | Disarankan penggantian ILCA2 | Irfan R
7000433 | 2023-02-24 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
7000432 | 2023-02-24 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
7000373 | 2023-02-24 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0091 | Survey | O2 supply, air supply dan kalibrasi O2 gagal | Pending | Disarankan penggantian PATO SA | Wahyu
8000527 | 2023-02-24 | RSI Sultan Hadlirin | Mesin Anasthesia | Draeger/Fabius Plus | ASND0059 | Maintenance | General check | Ready to use |  | Amna
10000154 | 2023-02-23 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Penggantian kabel extension karna joystick console error | Ready to use |  | Heru
13000152 | 2023-02-23 | RSUD Dr. Fauziah | ESWL | Medispec/EM1000 | 0716249 | Service | Isi ulang air EMH karna EMH tidak bisa menembak | Ready to use |  | Sofian
19000563 | 2023-02-23 | RSUP dr. Kariadi | Dental Unit | Kavo/1058 | 1002722 | Service | Air tidak keluar, mengganti membran | Ready to use |  | Nanda
3000253 | 2023-02-23 | RSUD Gunung Jati Cirebon | X-ray Digital | Apelem/Camargue DR | IMA704014 | Survey | Masuk ke software program lama | Pending | Dibuatkan penawaran upgrade software dan storage | Imam
7000372 | 2023-02-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Maintenance | Preventinve Maintenance (Total Care) | Ready to use |  | Wahyu
13000147 | 2023-02-22 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Service | Penggantian HVPS | Ready to use |  | Sofian
6000626 | 2023-02-22 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASJF0034 | Survey | General check | Ready to use | Disarankan penggantian service set 6 tahun, spirolog flow sensor, waterlock, prmus breathing set 2tahun
6000626 | 2023-02-22 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASMH0115 | Survey | General check | Ready to use | Disarankan penggantian service set 3 tahun, spirolog flow sensor, waterlock, prmus breathing set 2tahun
6000626 | 2023-02-22 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASNJ0039 | Survey | General check | Ready to use
6000626 | 2023-02-22 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0039 | Survey | General check | Ready to use | Disarankan penggantian O2 sensor
6000626 | 2023-02-22 | RSUD Dr. R. Goeteng Taroenadibrata | Incubator Transport | Draeger/TI500 | CR12075 | Survey | General check. Tukar dengan baterai unit lain | Ready to use
19000562 | 2023-02-21 | RSUD dr. Gunawan Mangunkusumo | Incubator Transport | Draeger/Isolette TI500 | KY11399 | Service | Baterai internal rusak | Ready to use | Disarankan penggantian baterai internal | Nanda
6000625 | 2023-02-21 | RSUD Ajibarang | Patient Monitor | Draeger/Vista 120 | VGSNE5023 | Survey | Software bootloop. Install ulang fail | Pending | Disarankan penggantian display drive dan main pcbx | Afrizal
6000625 | 2023-02-21 | RSUD Ajibarang | Patient Monitor | Draeger/Vista 120 | V3SHK0114 | Survey | display hang | Pending | Disarankan penggantian front shellkit (inc. display) | Afrizal
3000250 | 2023-02-21 | RSUD Malingping | X-ray Digital | Apelem/Camargue DR | IMA703916 | Survey | Software lemot dan layar berkedip | Ready to use | Dibuatkan penawaran upgrade software dan storage | Imam
10000153 | 2023-02-20 | RSUD Ulin | ESWL | Medispec/E3000 | 712861 | Survey | Adjusment titik shock pada pointer | Ready to use |  | Heru
7000371 | 2023-02-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Wahyu
19000561 | 2023-02-17 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Service | Balon pecah | Pending | Disarankan penggantian EMH | Nanda
18000310 | 2023-02-17 | RSUD Toto Kabila | Ventilator | Draeger/Savina V300 | ASND0131 | Service | Penggantian board monitor (switch) | Ready to use |  | Faugi
13000151 | 2023-02-16 | RSUD Arifin Achmad | Dental Unit | Kavo/E30TM | 10111156 | Service | LCD dan Tombol dentist element mati | Pending |  | Sofian
19000560 | 2023-02-16 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
7000370 | 2023-02-16 | RSUD KOTA YOGYAKARTA | Ventilator | Draeger/Babylog VN800 | ASRK0005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
6000624 | 2023-02-16 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASJH0124 | Service | Penggantian battery pack, pba theraphy control unit, expiration valve, kit air filter, safety valve, particle filter, fan unit | Ready to use |  | Afrizal
6000695 | 2023-02-16 | RSUD Brebes | Ventilator | Dra
14000216 | 2023-02-15 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 0720015 | Service | Penggantian unit USG lama (SN: 6045408WXO - Logiq V5 expert) ke USG baru (SN:6063144WXO - Versana Essensial VA) | Ready to use |  | Purwono
19000559 | 2023-02-15 | RSUP dr. Kariadi | Baby Incubator | Draeger/Caleo | ARUC0047 | Service | bocor dan bau gosong, water heater kotor berkerak dan patah | Pending | Disarankan penggantian conv. Water heater kit | Nanda
16000214 | 2023-02-15 | Kesumah Dental Clinic | Dental Unit | Kavo/E30 | 10111157 | Maintenance | Preventive Maintenance | Ready to use |  | Ipan
14000215 | 2023-02-15 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 0720015 | Survey | USG mati total | Pending |  | Purwono
19000558 | 2023-02-13 | RSUD Ibu Fatmawati Surakarta | Trolley Emergency | SYASTECH | 230269501 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2024-02-13
8000526 | 2023-02-13 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0110 | Maintenance | General check | Ready to use |  | Amna
8000526 | 2023-02-13 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0108 | Maintenance | General check | Ready to use |  | Amna
8000526 | 2023-02-13 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASLL0147 | Maintenance | General check | Ready to use |  | Amna
8000526 | 2023-02-13 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASHD0034 | Maintenance | General check | Ready to use |  | Amna
8000526 | 2023-02-13 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASNB0020 | Maintenance | General check | Ready to use |  | Amna
8000526 | 2023-02-13 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASNB0021 | Maintenance | General check | Ready to use |  | Amna
7000369 | 2023-02-13 | RSA UGM | Ventilator | Draeger/Oxylog 2000 Plus | ASKH0020 | Service | Alarm leakage. Penggantian breathing valve dan flow sensor | Ready to use |  | Wahyu
4000401 | 2023-02-10 | RSUD Dr. H. Chasan Boesoirie | ESWL | Medispec/E3000 | 0922002 | Survey | Pelaksanaan uji paparan dan uji kesesuaian | Ready to use |  | Bayu
2000155 | 2023-02-09 | RSUD Kab. Temanggung | Mobile X-ray | Apelem/Rafale EV30 | 009/19/01306 | Maintenance | Preventive Maintenance | Ready to use |  | Eddy
8000525 | 2023-02-09 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Primus | ASMA0193 | Service | alat mati total, general check dan di charge | Ready to use |  | Amna
19000556 | 2023-02-09 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Service | Kontrak service | Ready to use |  | Nanda
19000557 | 2023-02-09 | RSUP dr. Kariadi | Dental Unit | Kavo/Primus 1058 | 1002722 | Service | air di bowl tidak keluar lama, setting dan adjustment | Ready to use |  | Nanda
19000555 | 2023-02-08 | RSUP Dr. Kariadi | ventilator | Draeger/Babylog VN800 | ASRF0014 | Service | Layar mati nyala | Pending |  | Nanda
2000153 | 2023-02-08 | RSUD Kab. Temanggung | X-Ray Analog | Apelem/Camargue Analog | CPD30673C18 | Service | Angle tube tidak sesuai pembacaannya; Perbaiki koneksi soket input sudut angle | Ready to use |  | Eddy
8000524 | 2023-02-08 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Primus | ASMA0193 | Service | alat mati total | Pending | Alat ditarik ke draeger | Amna
10000151 | 2023-02-08 | RSUD Sultan Suriansyah | X-ray Digital | Apelem/Camargue DR | CPD33483019 | Survey | Lampu kolimator mati | Pending | Disarankan penggantian lampu kolimator | Heru
3000251 | 2023-02-08 | RSUD Ulin | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/048/CB1 | Service | hasil foto tidak bagus | Ready to use | Disarankan menggunakan panel sendiri | Imam
2000154 | 2023-02-07 | RSUD Kab. Temanggung | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/011/C0 | Maintenance | Preventive Maintenance | Ready to use |  | Eddy
19000554 | 2023-02-07 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHA0114 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000554 | 2023-02-07 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHC0076 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
7000368 | 2023-02-07 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | FH36369 | Service | Nilai kelembapan tidak sesuai. Pembersihan heater | Ready to use |  | Wahyu
7000368 | 2023-02-07 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | FH36368 | Service | Nilai kelembapan tidak sesuai. Pembersihan heater | Ready to use |  | Wahyu
3000252 | 2023-02-07 | RSUD Ulin | Bone Mineral Densitometer | Medilink/Medix 90 | J15016M269 | Service | Penarikan spare part detector, monoblock, inverter | Ready to use |  | Imam
18000307 | 2023-02-06 | RSUD Samratulangi Tondano | ventilator | Draeger/Evita V300 | ASML0255 | Survey | Refresh training | Ready to use |  | Faugi
18000307 | 2023-02-06 | RSUD Samratulangi Tondano | ventilator | Draeger/Evita V300 | ASML0256 | Survey | Refresh training | Ready to use |  | Faugi
18000307 | 2023-02-06 | RSUD Samratulangi Tondano | ventilator | Draeger/Evita V300 | ASMJ0392 | Survey | Refresh training | Ready to use |  | Faugi
7000367 | 2023-02-06 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Caleo | ASKL0037 | Service | Alarm water empty, selang terlalu panjang. Penggantian selang | Ready to use |  | Wahyu
2000152 | 2023-02-06 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Service | Sering muncul error table | Ready to use |  | Eddy
11000296 | 2023-02-03 | RSUD Sawerigading | Mesin Anasthesia | Draeger/Perseus A500 | ASNJ0071 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Nasrul
8000523 | 2023-02-03 | RS ISLAM PATI | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0114 | Service | Penggantian housing asm | Ready to use |  | Amna
6000623 | 2023-02-03 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0047 | Service | Pemasangan fabius GS service set 3years | Ready to use |  | Afrizal
18000311 | 2023-02-02 | RSUD Ainun Habibie | Patient Monitor | Draeger/Infinity | ASPH0326 | Service | Penggantian kabel lcd monitor | Ready to use |  | Faugi
6000622 | 2023-02-02 | RSUD Brebes | Baby Incubator | Draeger/Caleo | ASJM0056 | Service | Pemasangan conv kit water heater | Ready to use |  | Afrizal
17000156 | 2023-02-02 | RSUD Gabriel Manek | Digital Radiografi Fluoroskopi | Apelem/Platinum | E226F0723 | Survey | Training alat | Ready to use |  | Bahar
18000305 | 2023-02-02 | RSUD Walanda Maramis | Mesin Anasthesia | Draeger/Primus | ASMH0120 | Service | Penggantian flow sensor | Ready to use |  | Faugi
7000366 | 2023-02-01 | RSUD Bagas Waras | Incubator Transport | Draeger/TI500 | DY11070 | Service | Penggantian baterai | Ready to use |  | Wahyu
13000149 | 2023-02-01 | RSUP Dr. H. Adam Malik Medan | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/083/C1 | Survey | Perbaikan printer dan cek koneksi | Ready to use |  | Sofian
19000552 | 2023-02-01 | RSUP Dr. Kariadi | Baby Incubator | Draeger/Isolette 8000 | MR0778 | Service | humidity tidak naik | Pending | Disarankan penggantian heater water tank | Nanda
2000151 | 2023-01-31 | RSUP dr. Ben Mboi | Digital Radiografi Fluoroskopi | Apelem/Platinum | F226F0736 | Service | Penggantian detector | Ready to use |  | Eddy
7000364 | 2023-01-31 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Maintenance | Maintenance total care | Ready to use |  | Wahyu
19000551 | 2023-01-31 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0304 | Service | Ukuran tegangan pressure tidak valid, kalibrasi tekanan | Ready to use |  | Nanda
18000304 | 2023-01-30 | RSUD Labuha | X-ray Digital | Apelem/Camargue DR | CPD36860K20 | Survey | Pemasangan digital wifi | Ready to use |  | Faugi
13000148 | 2023-01-30 | RSUD Dr. Djasamen Saragih | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/137/C2 | Survey | Pendampingan uji kesesuaian | Ready to use |  | Sofian
7000363 | 2023-01-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0007 | Service | Penggantian part ILCA2 | Ready to use |  | Wahyu
7000361 | 2023-01-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Service | Penggantian part conv. Kit display | Ready to use |  | Wahyu
7000365 | 2023-01-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Service | Gas delivery tidak terbaca, penggantian part ILCA2 | Ready to use |  | Wahyu
3000249 | 2023-01-30 | RSUD Berkah Pandeglang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0497 | Survey | mA low | Ready to use | Disarankan untuk perbaikan listrik 3 phase menggunakan UPS | Imam
6000621 | 2023-01-27 | RSUD Banyumas | ventilator | Draeger/Evita V300 | ASFL0120 | Survey | gas supply sensor calibration merah | Ready to use | Disarankan penggantian internal baterai, eksternal baterai, filter air o2 dan pcb | Afrizal
17000155 | 2023-01-27 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0917012 | Service | Penggantian EMH | Ready to use |  | Bahar
18000302 | 2023-01-26 | RSUD Dr. Hallussy | Mamografi | Metaltronica/Lilyum | 66K010 | Survey | Generator rusak | Pending |  | Faugi
15000215 | 2023-01-26 | RSUD Dr. M. Yunus | ESWL | Medispec/EM1000 | 0917011 | Survey | Dilakukan uji kesesuaian ulang | Ready to use |  | Sutan
19000548 | 2023-01-26 | RSUD Dr. Soehadi Prijonegoro | Baby Incubator | Draeger/Isolette C2000 | AC37611 | Survey | Controller failure II | Pending | Disarankan penggantian CPU board | Nanda
19000547 | 2023-01-26 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Fabius Plus | ASJM0264 | Service | Layar bleret, cleaning and repositioning cable | Ready to use |  | Nanda
19000547 | 2023-01-26 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Fabius Plus | ASJM0126 | Service | Selang udara bocor, ada yang robek. Penggantian selang | Ready to use |  | Nanda
6000620 | 2023-01-25 | RSI Harapan Anda | Baby Incubator | Draerger/Isolette C2000 | DSRJ0061 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000620 | 2023-01-25 | RSI Harapan Anda | Baby Incubator | Draerger/Isolette C2000 | DSRJ0067 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
19000546 | 2023-01-24 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0125 | Maintenance | Preventive Maintenance | Ready to use |  | Nanda
19000545 | 2023-01-24 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0124 | Service | FG outlet berwarna kuning, kemungkinan ada rusak pada pembacaaan gas outlet | Ready to use | Alat masih bisa digunakan | Nanda
8000522 | 2023-01-24 | RSI Siti Khadijah | Infant Warmer | Draeger/Isolette C2000 | EM31828 | Service | control failure 14 | Ready to use |  | Amna
19000544 | 2023-01-21 | RSUP Dr. Kariadi | ventilator | Draeger/Babylog VN800 | ASRF0014 | Service | Layar kedap kedip nyala hidup, penggantian kabel koneksi | Ready to use |  | Nanda
8000521 | 2023-01-20 | RSUP Dr. Kariadi | ventilator | Draeger/Babylog VN800 | ASRF0014 | Service | Penggantian kabel system karna layar kedip kedip | Ready to use |  | Amna
16000213 | 2023-01-20 | RS Dr. Suyoto | Dental Unit | Kavo/1058 TM | 10207194 | Service | Penggantian katup waterblock | Ready to use |  | Ipan
18000383 | 2023-01-19 | RSUD Ainun Habibie | Patient Monitor | Draeger/Infinity | ASPH0326 | Service | LCD Error. Dilakukan perbaikan konektor LCD | Ready to use |  | Irfan R
19000542 | 2023-01-19 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Service | balon kendor | Ready to use | Disarankan penggantian EMH dan kontrak service | Nanda
6000619 | 2023-01-19 | RSI Harapan Anda | ventilator | Draeger/Evita 4 edition | ASCC0210 | Survey | VT tidak teratur, PEEP minus | Pending | Disarankan penggantian pressure measurement unit, electrowave, PEEP valve, PCB pneumatic controller | Afrizal
2000150 | 2023-01-19 | RSUD Malingping | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0498 | Service | Setting dicom configuration | Ready to use |  | Eddy
8000520 | 2023-01-19 | RSUD RA Kartini | C-arm | Apelem/Evo R+ | 55120027813 | Service | Alat tidak bisa expose, general check | Ready to use | Overheating karna terlalu banyak expose | Amna
8000519 | 2023-01-18 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | Service | system check fail, general check | Ready to use |  | Amna
17000152 | 2023-01-18 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0486 | Service | Meja tidak bisa bergerak,cek koneksi di compression | Ready to use |  | Bahar
10000152 | 2023-01-18 | RSUD Ulin | ESWL | Medispec/E3000 | 712861 | Service | Shock loncat, pembersihan board console | Ready to use |  | Heru
7000360 | 2023-01-18 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | FH36367 | Survey | Engsel dan pengunci pintu chamber patah | Ready to use | Alat masih bisa digunakan | Wahyu
7000360 | 2023-01-18 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | FH36368 | Survey | Engsel dan pengunci pintu chamber patah | Ready to use | Alat masih bisa digunakan | Wahyu
7000360 | 2023-01-18 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | FH36369 | Survey | Engsel dan pengunci pintu chamber patah | Ready to use | Alat masih bisa digunakan | Wahyu
7000360 | 2023-01-18 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | FH36366 | Survey | Engsel dan pengunci pintu chamber patah | Ready to use | Alat masih bisa digunakan | Wahyu
19000541 | 2023-01-17 | RSUD Pandan Arang | Incubator Transport | Draeger/TI500 | FY11157 | Survey | supply baterai tidak masuk | Ready to use | alat masih dapat digunakan namun disarankan penggantian baterai | Nanda
19000540 | 2023-01-17 | RSUD Pandan Arang | Infant Warmer | Draeger/Resuscitaire | CF10509 | Service | Error 05, pemasangan heater | Ready to use |  | Nanda
17000153 | 2023-01-17 | RSUD Dr. Drajat Prawiranegara | Meja Pasien | SYAS/Standard Bed 1 crank | 2211622 | Survey | Roda tidak bisa berfungsi, sudah dilakukan penggantian roda | Ready to use |  | Bahar
16000212 | 2023-01-17 | RS Dr. Suyoto | Dental Unit | Kavo/1058TM | 10207194 | Survey | Muncul error pada waterblock | Ready to use |  | Ipan
3000248 | 2023-01-17 | RSUD Toto Kabila | X-ray Digital | Apelem/Camargue DR | CPD327474K18 | Service | Upgrade ssd dan update software | Ready to use |  | Imam
7000358 | 2023-01-16 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Survey | LCD rusak | Pending | Disarankan penggantian kabel fleksible dan connectornya | Wahyu
600549 | 2023-01-14 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHL0160 | Service | Pemasangan kit air filter dan baterai eksternal | Ready to use |  | Afrizal
6000548 | 2023-01-13 | RSI Harapan Anda | Lampu Operasi | Draeger/Polaris 600 | ASRK1014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
7000359 | 2023-01-13 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0007 | Survey | Gas delivery tidak terbaca, water trap berarir dan ganti watertrap | Ready to use |  | Wahyu
6000655 | 2023-01-13 | RSI Harapan Anda | ventilator | Draeger/Evita 4 edition | ASCC0210 | Survey | Alat error | Pending | Disarankan penggantian repl-kit LP CPU, valve air, valve O2, PCB pneumatic motherboard | Afrizal
19000536 | 2023-01-12 | RSU Banyumanik 2 | Ventilator | Draeger/Savina 300 | ASPJ0110 | Survey | General check | Ready to use |  | Nanda
17000151 | 2023-01-12 | RSUD Dr. Drajat Prawiranegara | X-ray Digital | Apelem/Camargue DR | 2510220562 | Survey | Training alat | Ready to use |  | Bahar
6000550 | 2023-01-12 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHM0104 | Service | Pemasangan baterai eksternal | Ready to use |  | Afrizal
8000518 | 2023-01-12 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0033 | Service | gas analyzer failed | Pending | Disarankan penggantian flow power 2, ILCA 3, dan motor pump | Amna
8000518 | 2023-01-12 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0037 | Service | gas analyzer failed | Pending | Disarankan penggantian flow power 2, ILCA 3, dan motor pump | Amna
5000179 | 2023-01-12 | RSUP dr. Ben Mboi | Digital Radiografi Fluoroskopi | Apelem/Platinum | F226F0736 | Service | Detector tidak menampilkan gambar saat expose. Sudah dilakukan remote oleh Apelem. Detector harus diganti | Ready to use | Proses claim warranty detector | Bilal
3000245 | 2023-01-12 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | FPD lost, reload software detector | Ready to use |  | Imam
7000355 | 2023-01-11 | RSJD Dr. RM. Soedjarwadi | Baby Incubator | Draeger/Caleo | ASJM0054 | Survey | Cover LCD pecah | Ready to use |  | Wahyu
2000149 | 2023-01-11 | RSK Dharmais | Digital Radiografi Fluoroskopi | Apelem/Platinum | L216F0664 | Service | Update DSA | Ready to use |  | Eddy
6000545 | 2023-01-11 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0013 | Survey | APL macet | Ready to use |  | Afrizal
6000547 | 2023-01-11 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Savina 300 | ARWB0347 | Survey | volume tidal tidak sesuai | Pending | Disarankan penggantian mainboard, service set 6years, blower internal, humidifier | Afrizal
6000546 | 2023-01-11 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Savina 300 | ARWC0062 | Service | volume tidal tidak sesuai | Pending | Disarankan penggantian mainboard, service set 6years, blower internal, humidifier | Afrizal
19000535 | 2023-01-11 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 2004002722 | Service | air kumur tidak keluar, penggantian switch S8 pada tanki air | Ready to use |  | Nanda
6000545 | 2023-01-11 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0013 | survey | APL macet. Melakukan pengurangan sodalime | Ready to use |  | Afrizal
19000534 | 2023-01-10 | RSUP Dr. Kariadi | ESWL | Medispec/EM1000 | 0922164 | Maintenance | Kuras air sebelum digunakan | Ready to use |  | Nanda
18000388 | 2023-01-10 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | BY33329 | Maintenance | Preventive Maintenance | Ready to use |  | Irfan R
3000247 | 2023-01-10 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | reload software dan kalibrasi detector | Ready to use |  | Imam
14000214 | 2023-01-09 | RSUD Dr. Saiful Anwar | Bone Mineral Densitometer | Medilink/Medix DR | J207H0735 | Service | Koneksikan alat ke PACS dan RIS | Ready to use |  | Purwono
7000354 | 2023-01-06 | RS Duta Mulya | Ventilator | Draeger/Savina 300 classic | ASRJ0121 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Wahyu
19000533 | 2023-01-05 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE5094 | Service | Layar mati, ada kerusakan pada LCD board dan screen driver LCD | Pending | Disarankan penggantian LCD dan screen driver LCD | Nanda
7000353 | 2023-01-05 | RSUD Dr. Soedirman - Kebumen | Infant Warmer | Draeger/Isolette C2000 | AH36595 | Service | Perbaikan pada pemanas humidifier | Ready to use |  | Wahyu
7000353 | 2023-01-05 | RSUD Dr. Soedirman - Kebumen | Infant Warmer | Draeger/Isolette C2000 | AH36598 | Service | Perbaikan pada pemanas humidifier | Ready to use |  | Wahyu
17000150 | 2023-01-05 | RSUD Ulin | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150176 | Service | Alat mati sendiri saat dinaikan power, Alat harus menggunakan mcb 16A | Ready to use |  | Bahar
19000532 | 2023-01-04 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0039 | Service | sensor check merah, ada kerusakan sensor pdiif low dan sensor cylinder pressure | Pending | Disarankan penggantian sensor pdiif low high & sensor cylinder pressure | Nanda
7000352 | 2023-01-04 | RSUD Bagas Waras | Infant Warmer | Draeger/Resuscitaire | JY03386 | Survey | gas bocor, kerusakan pada sambungan selang valve | Pending |  | Wahyu
16000211 | 2023-01-04 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058 | 1007769 | Service | Pembersihan separator selang, penggantian separator speed low | Ready to use |  | Ipan
8000517 | 2023-01-03 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Maintenance | gas analyzer failed | Ready to use |  | Amna
19000531 | 2023-01-03 | RSUP Dr. Kariadi | Infant Warmer | Draeger/Resuscitaire | AF10368 | Survey | Error 08, kerusakan pada heater | Pending | Disarankan penggantian heater
19000531 | 2023-01-03 | RSUP Dr. Kariadi | Infant Warmer | Draeger/Resuscitaire | AF10368 | Survey | Error 08 | Pending | Disarankan penggantian heater | Nanda
7000351 | 2023-01-03 | RSUP Dr. Sardjito | Vaporizer | Draeger/Vapor 2000 | ASLC1417 | Survey | hasil keluaran gas rendah | Pending |  | Wahyu
9000397 | 2022-12-31 | RSUD Ainun Habibie | ventilator | Draeger/Evita V300 | ASLD0172 | Maintenance | Kalibrasi dan maintenance | Ready to use |  | Faugi
3000244 | 2022-12-30 | RSUD Sumedang | UPS | Yoshiga/GPM10 | 501110038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2023-12-30
19000530 | 2022-12-29 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0039 | Service | gas delivery sensor check merah, ada kerusakan pada valve | Pending |  | Nanda
9000396 | 2022-12-28 | RSUD Ainun Habibie | ventilator | Draeger/Babylog VN800 | ASRH0015 | Maintenance | FIO2 tidak sesuai setting, penggantian input kompressor dari portable ke sentral | Ready to use |  | Faugi
3000243 | 2022-12-28 | RSUD Ciracas | UPS | Yoshiga/PENTA1000 | 31003078 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2023-12-28
19000529 | 2022-12-27 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Fabius Plus | ASHH0122 | Service | gas delivery sensor check merah, setting adjusment | Ready to use |  | Nanda
8000516 | 2022-12-26 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | ASML0039 | Maintenance | System check, penggantian O2 sensor | Ready to use |  | Amna
8000516 | 2022-12-26 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | r | Service | System check, penggantian O2 sensor | Ready to use |  | Amna
8000515 | 2022-12-23 | RS Aisyiyah Kudus | Medical Compressor | Durr/Sicolab Med | DPSE0048 | Service | water trap terlepas, telah dipasang kembali | Ready to use |  | Amna
13000146 | 2022-12-23 | RSUD Djasamen Saragih | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/137/C2 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sofian | 2024-12-23
6000543 | 2022-12-23 | RSUD Kardinah Tegal | Ventilator | Medispec/ED1000 | 03111118 | Service | Pemasangan applicator | Ready to use |  | Afrizal
6000544 | 2022-12-23 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASEF0108 | Service | Pemasangan spare part pressure regulator, filter pad, particle filter, baterai | Ready to use |  | Afrizal
6000544 | 2022-12-23 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASHL0166 | Service | Pemasangan spare part filter pad, particle filter, baterai pack | Ready to use |  | Afrizal
2000146 | 2022-12-23 | RSUP Fatmawati | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0493 | Service | update DSA | Ready to use |  | Eddy
14000212 | 2022-12-22 | RSUD RA Basoeni | Dental Unit | Kavo/1058TM | 20007477 | Service | lampu mati, kabel putus dan sudah diganti | Ready to use |  | Purwono
9000395 | 2022-12-21 | RSUD Datoe Binangkang | Lampu Operasi | Draeger/Polaris 600 | ASLL0245 | Survey | Penyambungan listrik ke unit dari panel utama | Ready to use |  | Faugi
9000395 | 2022-12-21 | RSUD Datoe Binangkang | Lampu Operasi | Draeger/Polaris 600 | ASMA0338 | Survey | Penyambungan listrik ke unit dari panel utama | Ready to use |  | Faugi
9000395 | 2022-12-21 | RSUD Datoe Binangkang | Lampu Operasi | Draeger/Polaris 600 | ASLL0239 | Service | Penyambungan listrik ke unit dari panel utama | Ready to use |  | Faugi
6000540 | 2022-12-21 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Survey | error driver motor table | Ready to use | Disarankan reload software | Afrizal
5000177 | 2022-12-20 | RSAU Dr. Esnawan Antariksa | Mobile X-ray | Apelem/Rafale EV30 | 0051801391 | Service | Penggantian kabel inverter | Ready to use |  | Bilal
20000226 | 2022-12-20 | RSKD Dadi | Examination Bed | SYAS | 22107394384 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana | 2023-12-20
14000211 | 2022-12-20 | RSUD Dr. R. Sosodoro Djatikoesoemo Bojonegoro | Dental Unit | Kavo/E30TM | 10111154 | Service | suction mati, diperbaiki | Ready to use |  | Purwono
3000242 | 2022-12-20 | RSUD Dr. Soediran Mangun Sumarso | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/134/C2 | Service | Penggantian processor interl core i7 | Ready to use |  | Imam
9000394 | 2022-12-19 | RSU Kinapit | ventilator | Draeger/Evita V300 | ASMN0109 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi
5000176 | 2022-12-19 | RSAU Dr. Esnawan Antariksa | Mobile X-ray | Apelem/Rafale EV30 | 0051801391 | Service | Error inverter fault, ada kabel yg error digigit tikus | Ready to use |  | Bilal
6000539 | 2022-12-19 | RS Nirmala | Ventilator | Draeger/Savina 300 | ASRE0307 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000538 | 2022-12-16 | RSUD Kajen | Ventilator | Draeger/Babylog VN800 | ASRD0040 | Service | Nilai FIO2 tidak sesuai setting, dialihkan menggunakan kompresor | Ready to use |  | Afrizal
15000211 | 2022-12-16 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Mic console tidak berfungsi, penggantian mic | Ready to use |  | Sutan
18000308 | 2022-12-15 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Babyleo TI500 | ASMR0048 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi
18000308 | 2022-12-15 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Babyleo TI500 | ASMR0043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi
11000290 | 2022-12-15 | RSKD Dadi | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0050 | Survey | O2 sensor fail | Pending | Disarankan penggantian O2 sensor dan service set 3 tahun | Nasrul
11000290 | 2022-12-15 | RSKD Dadi | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0031 | Service | O2 sensor fail | Pending | Disarankan penggantian O2 sensor dan service set 3 tahun | Nasrul
2000145 | 2022-12-15 | RSUD Malingping | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/131/C2 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2024-12-15
9000392 | 2022-12-15 | RSUD Ainun Habibie | UPS | Yoshiga | 1100258199 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi | 2023-12-15
9000391 | 2022-12-15 | RSUD Otohana | ventilator | Draeger/Evita V800 | ASPL0095 | Survey | General check | Ready to use | Disarankan menggunakan kompresor central | Faugi
9000391 | 2022-12-15 | RSUD Otohana | ventilator | Draeger/Evita V300 | ASPA0007 | Survey | General check | Ready to use | Disarankan menggunakan kompresor central | Faugi
9000391 | 2022-12-15 | RSUD Otohana | ventilator | Draeger/Evita V300 | ASPE0310 | Maintenance | General check | Ready to use | Disarankan menggunakan kompresor central | Faugi
3000241 | 2022-12-14 | RSUD Sumedang | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/132/C2 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2024-12-14
9000374 | 2022-12-14 | RSUD Dr. H. Chasan Boesoirie | ESWL | Medispec/E3000 | 0922002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2024-12-14
18000309 | 2022-12-13 | RSUD Walanda Maramis | Mesin Anasthesia | Draeger/Primus | ASMH0120 | Service | O2 sensor failure, kalibrasi unit/selft test | Ready to use
7000348 | 2022-12-13 | RSUD Muntilan | ventilator | Draeger/Babylog VN800 | ASRF0013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana
20000225 | 2022-12-13 | RSUD Morowali | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/126/C2 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2024-12-13
16000210 | 2022-12-13 | Poli Gigi Klinik DPRD | Dental Unit | Kavo/1058 | 10206373 | Maintenance | general check | Ready to use |  | Ipan
20000223 | 2022-12-12 | RSUD Kota Kendari | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/127/C2 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana | 2024-12-12
15000210 | 2022-12-12 | RSUD Hasanuddin Damrah | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/138/C2 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan | 2024-12-12
7000347 | 2022-12-12 | RSJD Dr. RM. Soedjarwadi | Mesin Anasthesia | Draeger/Atlan A350 | ASPJ0045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
9000390 | 2022-12-12 | Dinas Kesehatan Kab. Gorontalo | Lemari Obat | SYAS | 221132923004005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi | 2023-12-12
6000536 | 2022-12-12 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0043 | Survey | Pengecekan Vaporizer | Ready to use |  | Afrizal
9000390 | 2022-12-12 | Dinas Kesehatan Kab. Gorontalo | Cabinet | SYAS | 221132923004006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi | 2023-12-12
9000388 | 2022-12-10 | RSUD Ainun Habibie | ventilator | Draeger/Evita V300 | ASLE0403 | Maintenance | General check | Ready to use | Disarankan mengganti baterai dan filter O2 | Faugi
6000532 | 2022-12-09 | RSUD Banyumas | Ventilator | Draeger/Savina 300 | ASRM0113 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000532 | 2022-12-09 | RSUD Banyumas | Ventilator | Draeger/Savina 300 | ASRM0114 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000532 | 2022-12-09 | RSUD Banyumas | Ventilator | Draeger/Savina 300 | ASRM0115 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000532 | 2022-12-09 | RSUD Banyumas | Ventilator | Draeger/Savina 300 | ASRM0116 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000532 | 2022-12-09 | RSUD Banyumas | Ventilator | Draeger/Savina 300 | ASRM0087 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
8000514 | 2022-12-09 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0034 | Service | Penggantian flow power, ILCA 3 dan update firmware | Ready to use |  | Amna
17000149 | 2022-12-08 | RSUD Dr. Drajat Prawiranegara | X-ray Digital | Apelem/Camargue DR | 2510220562 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2024-12-08
6000533 | 2022-12-08 | RSUD Ajibarang | UPS | Yoshiga/GPH33-30 | 501212003460N | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2023-12-08
8000509 | 2022-12-07 | RSUD Andi Makassau | ESWL | Medispec/EM1000 | 0622173 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2024-12-07
6000531 | 2022-12-07 | RSI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASRJ0080 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
13000144 | 2022-12-07 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Service | Perbaikan EMH dan penggantian selang dan HDMI | Ready to use | EMH masih bocor, menunggu spare part Power supply HVPS | Sofian
6000541 | 2022-12-07 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0025 | Service | Pemasangan service set 3 thn, maintenance kit 2thn | Ready to use |  | Afrizal
15000209 | 2022-12-07 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Service | CPU bunyi beep, motherboard rusak | Pending | Sudah diajukan penawaran harga | Sutan
7000334 | 2022-12-05 | RS Aisyiyah Kudus | Bubble CPAP | Draeger/Oxymixer | QM20215174 | Survey | Alat berdengung | Ready to use | Disarankan penambahan pressure regulator | Jarwo
19000527 | 2022-12-05 | RSUD dr. Soediran Mangun Sumarso | Ventilator | Draeger/Savina 300 | ASJM0264 | Service | Baterai internal failure, pemasangan servoce set 6 tahun | Ready to use |  | Nanda
2000148 | 2022-12-05 | RSUD Kab. Temanggung | X-Ray Analog | Apelem/Camargue Analog | CPD30673C18 | Service | Adjust dan setting parameter potensio | Ready to use |  | Eddy
2000147 | 2022-12-05 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Survey | Table error dan stiching terpotong, reload parameter driver | Ready to use |  | Eddy
6000530 | 2022-12-02 | RSKBD Panti Nugroho | Pendant | Draeger/Movita | ASMA0061 | Survey | gerakan pendant terkunci | Ready to use | Dianjurkan stabilkan suhu pendant dan ruangan dan tidak lembab | Afrizal
6000530 | 2022-12-02 | RSKBD Panti Nugroho | Pendant | Draeger/Movita | ASMK0350 | Survey | gerakan pendant terkunci | Ready to use |  | Afrizal
6000529 | 2022-12-02 | RSU An Ni'mah | ESMR | Draeger/CPAP | QM20215459 | Service | cek humidifier | Ready to use |  | Afrizal
15000208 | 2022-12-02 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Service | Penggantian applicator | Ready to use |  | Sutan
18000366 | 2022-12-02 | RSUD Ainun Habibie | Infant Warmer | Draeger/Resuscitaire | DY03438 | Service | Penggantian heater | Ready to use |  | Irfan R
18000366 | 2022-12-02 | RSUD Ainun Habibie | Infant Warmer | Draeger/Resuscitaire | DY03435 | Service | Penggantian heater | Ready to use |  | Irfan R
18000366 | 2022-12-02 | RSUD Ainun Habibie | Mesin Anasthesia | Draeger/Atlan | ASMM0118 | Service | Penggantian O2 sensor | Ready to use |  | Irfan R
9000375 | 2022-12-01 | Dinkes Kab. Gorontalo Utara | Mobile X-ray | Apelem/Rafale EV30 | 1890192200001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2024-12-01
3000240 | 2022-12-01 | RSUD Gabriel Manek | Digital Radiografi Fluoroskopi | Apelem/Platinum | E226F0723 | Survey | Uji fungsi printer dicom | Ready to use |  | Imam
7000333 | 2022-11-30 | RSUD Dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Primus | ASKJ0050 | Service | Error power supply, self test | Ready to use |  | Jarwo
7000332 | 2022-11-30 | RSUD Dr. Soedirman - Kebumen | UPS | Yoshiga/HPT2 | 501160KA41 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-11-30
19000524 | 2022-11-30 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Fabius Plus | ASJM0264 | Service | Penggantian service set 2 tahun dan ribbon cable O2 | Ready to use |  | Nanda
7000331 | 2022-11-30 | RSUP dr. Soeradji Tirtonegoro | ventilator | Draeger/Babylog VN500 | ASKL0067 | Service | Penggantian filament flow sensor | Ready to use |  | Jarwo
7000331 | 2022-11-30 | RSUP dr. Soeradji Tirtonegoro | ventilator | Draeger/Babylog VN500 | ASKL0486 | Service | Penggantian filament flow sensor | Ready to use |  | Jarwo
7000331 | 2022-11-30 | RSUP dr. Soeradji Tirtonegoro | ventilator | Draeger/Babylog VN500 | ASKL0460 | Service | Penggantian filament flow sensor | Ready to use |  | Jarwo
7000331 | 2022-11-30 | RSUP dr. Soeradji Tirtonegoro | ventilator | Draeger/Babylog VN500 | ASKL0091 | Service | Penggantian filament flow sensor | Ready to use |  | Jarwo
18000364 | 2022-11-29 | RS Mata Sulawesi Utara | Mesin Anasthesia | Draeger/Perseus A500 | ASJM0178 | Survey | General check, namun baterai error | Ready to use | Disarankan penggantian baterai | Irfan R
14000213 | 2022-11-29 | RSUD Sidoarjo Barat | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0755 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono | 2024-11-29
6000652 | 2022-11-29 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASRJ0002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
13000142 | 2022-11-29 | RSUD Simeulue | Bone Mineral Densitometer | Medilink/Medix DR | F207H0702 | Survey | Pendampingan ukes dengan BPFK Jkt | Ready to use |  | Sofian
9000387 | 2022-11-29 | RSUD Labuha | X-ray Digital | Apelem/Camargue DR | IMA704030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi | 2024-11-29
6000653 | 2022-11-25 | RS Prima Medika | Ventilator | Draeger/Savina 300 | ASHN0071 | Service | O2 sensor failed, internal baterai tidak ada | Pending | Disarankan penggantian power supply, oxytrace dan service set savina 3years kit | Afrizal
7000315 | 2022-11-25 | RSUD Dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Primus | ASKJ0050 | Service | Penggantian power supply | Ready to use |  | Jarwo
18000386 | 2022-11-25 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0410 | Service | Perbaikan kabel harness | Ready to use |  | Irfan R
18000386 | 2022-11-25 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0172 | Service | Perbaikan kabel harness | Ready to use |  | Irfan R
14000210 | 2022-11-24 | RSUD Sidoarjo Barat | C-arm | Apelem/Evo R+ | 55121062913 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono | 2024-11-24
3000239 | 2022-11-24 | RSUD Pringsewu | Digital Radiografi Fluoroskopi | Apelem/Platinum | J226F0756 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2024-11-24
11000289 | 2022-11-24 | RSKD Dadi | Bed Patient | Syas/Syas Premium | 221034219254 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2023-11-24
19000522 | 2022-11-24 | RSUP Dr. Kariadi | Jaudience Meter | Draeger/JM 105 | B3502043 | Service | circuit error 05, ada kerusakan pada baterai internal | Pending | Butuh penggantian baterai internal dan dilakukan kalibrasi internal | Nanda
11000289 | 2022-11-24 | RSKD Dadi | Bed Patient | Syas/Syas Premium | 221034219252 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2023-11-24
6000657 | 2022-11-23 | RS Dadi Keluarga | ventilator | Draeger/Babylog VN800 | ASPF0030 | Service | Pemasangan kit proportional valve | Ready to use |  | Afrizal
6000656 | 2022-11-23 | RSKBD Panti Nugroho | Pendant | Draeger/Movita | ASLL0154 | Survey | sistem pengereman tidak berfungsi baik | Pending | Disarankan penggantian module control electronic mainboard, valve block brake arm, gas connector | Afrizal
6000656 | 2022-11-23 | RSKBD Panti Nugroho | Pendant | Draeger/Movita | ASMK0349 | Service | sistem pengereman tidak berfungsi baik | Pending | Disarankan penggantian module control electronic mainboard, valve block brake arm, gas connector, jose 12x3 | Afrizal
16000204 | 2022-11-23 | RSUD Dr. Drajat Prawiranegara | Operating table | SYAS | 221167221144 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan | 2023-11-23
16000206 | 2022-11-23 | RSUD Dr. Drajat Prawiranegara | Operating table | SYAS | 221167221147 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan | 2023-11-23
16000205 | 2022-11-23 | RSUD Dr. Drajat Prawiranegara | Operating table | SYAS | 221167221148 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan | 2023-11-23
19000518 | 2022-11-23 | RSUD Dr. Soeratno Gemolong | Jaudience Meter | Draeger/JM 105 | B3503803 | Service | Kalibrasi jaudience meter | Ready to use |  | Nanda
7000330 | 2022-11-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Service | Error gas delivery. Penggantian ILCA dan flow power | Ready to use |  | Jarwo
8000508 | 2022-11-22 | RSI Siti Khadijah | Baby Incubator | Draeger/Isolette C2000 | EM31828 | Survey | Penggantian PCB assy, interface | Ready to use |  | Amna
18000363 | 2022-11-22 | RSUD Ainun Habibie | Resuscitaire | Draeger/Resuscitaire | DY03434 | Service | Kerusakan pada tombol on/off | Pending | Disarankan perbaikan tombol on/off | Irfan R
7000326 | 2022-11-21 | RSUP dr. Soeradji Tirtonegoro | ventilator | Draeger/Babylog VN500 | ASLL0067 | Service | Penggantian kabel flow sensor | Ready to use |  | Jarwo
7000327 | 2022-11-21 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Isolette 8000 | AC09112 | Service | Penggantian engsel pintu | Ready to use |  | Jarwo
7000325 | 2022-11-21 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Isolette 8000 | AC09111 | Service | Penggantian engsel pintu | Ready to use |  | Jarwo
7000328 | 2022-11-19 | RSUD Wonosari | ESWL | Medispec/E3000 | 0722027 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2024-11-19
3000238 | 2022-11-18 | RSUD Merah Putih | X-Ray Analog | Apelem/Camargue Analog | CPD34328J19 | Service | Rotor fault | Ready to use |  | Imam
7000329 | 2022-11-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Survey | error gas delivery sensor | Ready to use | Alat masih bisa digunakan namun untuk gas tidak bisa terminator. Perlu penggantian flow power dan ILCA | Jarwo
6000654 | 2022-11-16 | RSI Harapan Anda | Ventilator | Draeger/Savina 300 | ASFJ0001 | Service | Pemasangan servoce set 2 tahun kit dan oxytrace | Ready to use |  | Afrizal
3000237 | 2022-11-16 | RSUD Dr. Adjidarmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Service | Penggantian PC Console | Ready to use |  | Imam
8000507 | 2022-11-15 | RSUD RA Kartini | Ventilator | Draeger/Savina 300 | ASEK0197 | Service | Penggantian baterai internal | Ready to use |  | Amna
8000506 | 2022-11-14 | RSUD K.R.M.T Wongsonegoro | ventilator | Draeger/Babylog VN500 | ASNL0109 | Service | Penggantian baterai | Ready to use |  | Amna
14000209 | 2022-11-14 | RSUD Dr. R. Sosodoro Djatikoesoemo Bojonegoro | Dental Unit | Kavo/E30TM | 10111154 | Service | suction pump tidak berfungsi, motor dibongkar dan diperbaiki | Ready to use |  | Purwono
14000207 | 2022-11-10 | RSUD Bangil | C-arm | Apelem/Evo R+ | 55121061912 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Purwono | 2024-11-10
20000213 | 2022-11-10 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Primus | ASNK0060 | Service | Penggantian vaporizer | Ready to use |  | Sujana
19000512 | 2022-11-09 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | Penggantian shock dan power supply | Ready to use |  | Bahar
7000324 | 2022-11-09 | RSUD Dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Atlan A300 | ASRK0304 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
3000236 | 2022-11-09 | RS Immanuel | Digital Radiografi Fluoroskopi | Apelem/Platinum | H226F0745 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2024-11-09
8000505 | 2022-11-08 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Primus | ASJE0010 | Service | error gas analyzer | Pending | Disarankan penggantian ILCA2 dan baterai | Amna
19000513 | 2022-11-08 | RSUD Merah Putih | X-Ray Analog | Apelem/Camargue Analog | CPD34328J19 | Service | Rotor fault | Ready to use |  | Nanda
7000323 | 2022-11-08 | RSUP dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Fabius MRI | ASRJ0026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
18000362 | 2022-11-06 | RSUP Ratatotok Buyat | X-Ray Analog | Apelem/Camargue Analog | CPD36816520 | Service | Penggantian lampu kolimator | Ready to use |  | Irfan R
8000504 | 2022-11-04 | RS ISLAM PATI | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0114 | Service | General check, system check, running test | Ready to use |  | Amna
7000322 | 2022-11-04 | RSUD KOTA YOGYAKARTA | Jaudience Meter | Draeger/JM 105 | B3503416 | Maintenance | Kalibrasi jaudience meter | Ready to use |  | Jarwo
19000511 | 2022-10-31 | RSUP Dr. Kariadi | ESMR | Draeger/Resuscitaire | MC13337 | Service | Error 05 | Ready to use | Penggantian heater
15000207 | 2022-10-31 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Maintenance | penggantian applicator | Ready to use |  | Sutan
6000696 | 2022-10-28 | RSUD Brebes | ventilator | Draeger/Evita V300 | ASJH0125 | Service | Device check incomplete | Ready to use | Disarankan penggantian battery pack, kit air filter, particle filter, filter pad | Afrizal
19000510 | 2022-10-27 | RSUP Dr. Kariadi | Infant Warmer | Draeger/Resuscitaire | MC13338 | Service | Error 05, heater putus | Ready to use | Disarankan penggantian water assy. Sudah dilakukan penggantian | Nanda
16000202 | 2022-10-27 | Poli Gigi Klinik DPRD | Dental Unit | Kavo/Primus 1058 | 10206373 | Service | general check | Ready to use |  | Ipan
3000235 | 2022-10-27 | RSUD Alimuddin Umar | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0402 | Service | Table side control tidak konek | Ready to use |  | Imam
19000509 | 2022-10-26 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0145 | Service | O2 sensor fail, penggantian O2 sensor dan kabelnya | Ready to use |  | Nanda
19000508 | 2022-10-26 | RSUD Dr. Soeratno Gemolong | Baby Incubator | Draeger/Isolette C2000 | AC37611 | Service | control failure II | Pending | Akan dicoba update software | Nanda
18000384 | 2022-10-26 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0173 | Service | Penggantian pressure regulator | Ready to use |  | Irfan R
18000384 | 2022-10-26 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0415 | Service | Penggantian pressure regulator | Ready to use |  | Irfan R
18000384 | 2022-10-26 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0166 | Service | Penggantian pressure regulator | Ready to use |  | Irfan R
19000506 | 2022-10-25 | RS Nasional Diponegoro | Infant Warmer | Draeger/Resuscitaire | AF10307 | Service | Gas supply bocor | Pending | Disarankan penggantian control mini 5 way | Nanda
8000503 | 2022-10-25 | RSU Fastabiq Sehat | Mesin Anasthesia | Draeger/Fabius Plus | ASRC0070 | Service | General check, system check, running test | Ready to use |  | Amna
7000314 | 2022-10-25 | RSUD Dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Primus | ASKJ0050 | Maintenance | Penggantian service set 3 thn dan MV2 | Ready to use |  | Jarwo
19000507 | 2022-10-25 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0180 | Service | gas delivery alarm | Pending | Ada air masuk mixer, perlu penggantian sensor pembacaan gas | Nanda
19000505 | 2022-10-24 | RSUD Dr. Moewardi | UPS | Yoshiga/PENTA1000 | 31002582 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2023-10-24
7000313 | 2022-10-24 | RSUD KOTA YOGYAKARTA | Jaudience Meter | Draeger/JM 105 | B3503376 | Maintenance | Kalibrasi jaudience meter | Ready to use |  | Jarwo
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001183 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001195 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001172 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001205 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001190 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001171 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001177 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001202 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001199 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001196 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001185 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001198 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001189 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001174 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001180 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001193 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001179 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001182 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001173 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001186 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001169 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001067 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001168 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001178 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001192 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001197 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001181 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001203 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001191 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001176 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001206 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001200 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001187 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001184 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001170 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001188 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001194 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
7000317 | 2022-10-23 | FKG Universitas Gajah Mada | Dental Simulator | Kavo/DSE Expert | 10001204 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-23
13000141 | 2022-10-22 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Service | Balon EMH bocor, konektor EMH rusak | Ready to use | Menunggu spare part | Sofian
5000175 | 2022-10-22 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058 TM | 1007769 | Service | perbaikan 3 ways syringe | Ready to use |  | Bilal
19000504 | 2022-10-21 | RS JIH | Mesin Anasthesia | Draeger/Atlan | ASMM0125 | Maintenance | Maintenance total care | Ready to use |  | Nanda
19000504 | 2022-10-21 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Maintenance | Maintenance total care | Ready to use |  | Nanda
20000212 | 2022-10-21 | RSUD Bahteramas | Bed Patient | Syas/Noble Bed | 2209791012640080 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sujana | 2023-10-21
15000205 | 2022-10-21 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Service | perbaikan komputer detektor | Ready to use |  | Sutan
18000367 | 2022-10-21 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0166 | Service | Penggantian power supply dan service set 1thn | Ready to use |  | Irfan R
8000502 | 2022-10-20 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Service | flow tube O2 tidak sesuai, setting ulang dan system check | Ready to use |  | Amna
9000385 | 2022-10-20 | RS Siloam Paal Dua | Bubble CPAP | Draeger/CPAP | QM20190223 | Service | Pemasangan water trap yang lepas | Ready to use |  | Faugi
6000661 | 2022-10-20 | RSUD Dr. R. Goeteng Taroenadibrata | Incubator Transport | Draeger/TI500 | CR12075 | Survey | Pemasangan probe 1 skin temp, Iris port | Ready to use |  | Afrizal
6000660 | 2022-10-20 | RSUD Dr. R. Goeteng Taroenadibrata | UPS | Yoshiga | 3100258 | Service | general check | Ready to use | Disarankan standby nyala terus saja | Afrizal
18000359 | 2022-10-20 | RSUD Toto Kabila | X-ray Digital | Apelem/Camargue DR | CPD327474K18 | Survey | rel meja karat, meja macet | Ready to use |  | Irfan R
18000368 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0414 | Service | Penggantian service set 1thn | Ready to use |  | Irfan R
18000368 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0411 | Service | Penggantian service set 1thn | Ready to use |  | Irfan R
18000368 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0419 | Service | Penggantian service set 1thn | Ready to use |  | Irfan R
18000368 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0415 | Service | Penggantian service set 1thn | Ready to use |  | Irfan R
18000368 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0412 | Service | Penggantian service set 1thn | Ready to use |  | Irfan R
18000368 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0183 | Service | Penggantian service set 1thn | Ready to use |  | Irfan R
18000368 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0410 | Service | Penggantian service set 1thn | Ready to use |  | Irfan R
18000368 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0168 | Service | Penggantian service set 1thn | Ready to use |  | Irfan R
18000373 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASHD0006 | Service | Penggantian service set 1 thn | Ready to use |  | Irfan R
18000372 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0165 | Service | Penggantian service set 1 thn | Ready to use |  | Irfan R
18000372 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0167 | Service | Penggantian service set 1 thn | Ready to use |  | Irfan R
18000372 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0413 | Service | Penggantian service set 1 thn | Ready to use |  | Irfan R
18000372 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASEL0164 | Service | Penggantian service set 1 thn | Ready to use |  | Irfan R
18000372 | 2022-10-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASEL0166 | Service | Penggantian service set 1 thn | Ready to use |  | Irfan R
3000234 | 2022-10-20 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | upgrade software, ganti ssd | Ready to use |  | Imam
9000384 | 2022-10-19 | RS Gunung Maria Tomohon | Baby Incubator | Draeger/Isolette C2000 | HC37698 | Maintenance | sensor skin tidak berbaca, modul rusak | Pending |  | Faugi
19000502 | 2022-10-19 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0124 | Maintenance | Maintenance total care | Ready to use |  | Nanda
19000502 | 2022-10-19 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0125 | Maintenance | Maintenance total care | Ready to use |  | Nanda
18000377 | 2022-10-19 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASHD0032 | Service | Pemasangan power supply dan baterai internal | Ready to use |  | Irfan R
18000377 | 2022-10-19 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASHD0029 | Service | Pemasangan power supply dan baterai internal | Ready to use |  | Irfan R
18000377 | 2022-10-19 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASHD0033 | Service | Pemasangan power supply dan baterai internal | Ready to use |  | Irfan R
18000377 | 2022-10-19 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASFA0224 | Service | Pemasangan power supply dan baterai internal | Ready to use |  | Irfan R
18000377 | 2022-10-19 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASHA0100 | Survey | Pemasangan power supply dan baterai internal | Ready to use |  | Irfan R
3000233 | 2022-10-19 | RSUD Tidar Magelang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0496 | Service | Upgrade software | Ready to use |  | Imam
18000379 | 2022-10-18 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | BY33326 | Service | Penggantian part hinge, access pnl right left | Ready to use |  | Irfan R
18000379 | 2022-10-18 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | BY33328 | Service | Hinge, hood/access pnl left right dan service set 3thn | Ready to use |  | Irfan R
5000174 | 2022-10-18 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058 TM | 1007769 | Service | perbaikan selang dan saluran air | Ready to use |  | Bilal
8000501 | 2022-10-17 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina 300 | ASPJ0382 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000501 | 2022-10-17 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina 300 | ASPJ0383 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
5000173 | 2022-10-17 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Survey | pengcekan applicator | Ready to use |  | Bilal
19000501 | 2022-10-17 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDL0003 | Service | Penggantian knop pengganti vaporizer | Ready to use |  | Nanda
19000501 | 2022-10-17 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDK0300 | Service | Penggantian knop pengganti vaporizer | Ready to use |  | Nanda
19000501 | 2022-10-17 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDL0002 | Service | Penggantian knop pengganti vaporizer | Ready to use |  | Nanda
9000381 | 2022-10-17 | RSUD ODSK Provinsi Sulawesi Utara | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0066 | Maintenance | Kalibrasi | Ready to use |  | Faugi
18000380 | 2022-10-17 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | AR35152 | Service | Penggantian hing, acces pnl right left dan service set 3thn | Ready to use |  | Irfan R
18000381 | 2022-10-17 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | BY33331 | Service | Penggantian part hinge, access pnl right left | Ready to use |  | Irfan R
18000381 | 2022-10-17 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | BY33329 | Service | Penggantian part hinge, access pnl right left | Ready to use |  | Irfan R
18000381 | 2022-10-17 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | BY33323 | Service | Penggantian part hinge, access pnl right left | Ready to use |  | Irfan R
18000381 | 2022-10-17 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | KM04081 | Service | Penggantian part hinge, access pnl right left | Ready to use |  | Irfan R
18000381 | 2022-10-17 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | KM04800 | Service | Penggantian part hinge, access pnl right left | Ready to use |  | Irfan R
18000381 | 2022-10-17 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | CU03934 | Service | Penggantian part hinge, access pnl right left | Ready to use |  | Irfan R
18000381 | 2022-10-17 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | BY33330 | Service | Penggantian part hinge, access pnl right left | Ready to use |  | Irfan R
18000381 | 2022-10-17 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | CR35567 | Maintenance | Penggantian part hinge, access pnl right left | Ready to use |  | Irfan R
6000662 | 2022-10-15 | RS Dadi Keluarga | ventilator | Draeger/Babylog VN500 | ASPF0030 | Service | air supply, o2 supply dan internal gas sensor tidak berfungsi | Pending | Software bermasalah. Akan didiskusikan dengan kantor pusat | Afrizal
20000211 | 2022-10-15 | RSIA Defina | Medical Compressor | Draeger/Medical Compressor | 2107103087 | Survey | Penggantian kompresor cpap | Ready to use |  | Sujana
18000382 | 2022-10-15 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | BY333332 | Service | Penggantian spare part Hing, access pal right left | Ready to use |  | Irfan R
18000382 | 2022-10-15 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | BY33327 | Service | Penggantian spare part Hing, access pal right left | Ready to use |  | Irfan R
18000382 | 2022-10-15 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C8000 | AY34025 | Service | Penggantian spare part Hing, access pal right left | Ready to use |  | Irfan R
6000663 | 2022-10-14 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASEF0108 | Survey | Bootlooping | Ready to use | Disarankan penggantian baterai internal, filter pad, filter air, pressure regulator | Afrizal
6000663 | 2022-10-14 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASEF0030 | Survey | general check | Pending | Disarankan penggantian PATO O2 sensor, pressure regulator, alarm system, baterai internal, filter O2 dan filter air | Afrizal
3000232 | 2022-10-14 | RSUD Berkah Pandeglang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0497 | Survey | Update software adam version | Ready to use |  | Imam
6000664 | 2022-10-13 | RSUD Brebes | Jaudience Meter | Draeger/JM 105 | B3503457 | Survey | time periodic calibration tidak sesuai | Pending | Disarankan dikalibrasi dengan draeger dan penggatian baterai | Afrizal
6000664 | 2022-10-13 | RSUD Brebes | Baby Incubator | Draeger/Isolette C2000 | DSNM0005 | Service | Humidty berair luber | Ready to use |  | Afrizal
19000500 | 2022-10-13 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Service | Setting water and cleaning | Ready to use |  | Nanda
8000500 | 2022-10-12 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina 300 | ASPJ0382 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000500 | 2022-10-12 | RS PKU Muhammadiyah Mayong | Ventilator | Draeger/Savina 300 | ASPJ0383 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
9000380 | 2022-10-11 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120 | ASPH0326 | Service | Pemasangan sensor saturasi O2 diperbaiki | Ready to use |  | Faugi
6000673 | 2022-10-11 | RSUD Brebes | Baby Incubator | Draeger/Caleo | ASJM0056 | Survey | air exceeds beneath aggregrate housing | Ready to use | Disarankan penggantian MX46563 | Afrizal
6000665 | 2022-10-11 | RSUD Dr. Soeselo Slawi | Ventilator | Draeger/Babylog VN500 | ASJJ0104 | Survey | General check | Ready to use | Disarankan penggantian MK31424, touchscreen | Afrizal
6000665 | 2022-10-11 | RSUD Dr. Soeselo Slawi | Ventilator | Draeger/Babylog VN500 | ASJJ0103 | Service | General check | Ready to use | Disarankan penggantian baterai dan filter | Afrizal
11000286 | 2022-10-11 | RSUD H. Andi Sulthan Daeng Radja | ventilator | Draeger/Babylog VN500 | ASMK0030 | Survey | Penggantian kit ptoportional valve air dan O2, dan penggantian flow sensor | Ready to use |  | Nasrul
18000355 | 2022-10-11 | RSUD Tobelo | Delivery Bed | Paramount Bed | 9007050 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
16000201 | 2022-10-10 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Maintenance | General maintenance | Ready to use |  | Ipan
6000692 | 2022-10-10 | RS Dadi Keluarga | Baby Incubator | Draeger/Babylog VN800 | ASPF0030 | Maintenance | air supply dan nebulizer alarm merah | Ready to use |  | Afrizal
7000312 | 2022-10-10 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Service | General maintenance total care | Ready to use |  | Jarwo
6000672 | 2022-10-10 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius CE | ARWC0079 | Survey | power supply error, penggantian fuse | Ready to use |  | Afrizal
18000356 | 2022-10-10 | RSUD Tambulilato | X-ray Digital | Apelem/Camargue DR | CPD36814320 | Service |  | Ready to use |  | Irfan R | 2024-10-10
18000360 | 2022-10-10 | RSUD Toto Kabila | X-ray Digital | Apelem/Camargue DR | CPD327474K18 | Maintenance | Penggantian tombol on/off pada panel | Ready to use |  | Irfan R
19000499 | 2022-10-10 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Service | O2 fail dan gas measurement fail | Pending | Disarankan penggantian flow power | Nanda
9000379 | 2022-10-07 | RSUD Dr. M. M. Dunda | Mesin Anasthesia | Draeger/Primus | ASLD0029 | Maintenance | Penggantian flow sensor | Ready to use |  | Faugi
8000499 | 2022-10-07 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0049 | Maintenance | General maintenance | Ready to use |  | Amna
8000499 | 2022-10-07 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0148 | Maintenance | General maintenance | Ready to use |  | Amna
8000499 | 2022-10-07 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0115 | Service | General maintenance | Ready to use |  | Amna
7000345 | 2022-10-06 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHL0059 | Service | Penggantians service set 2 tahun | Ready to use |  | Jarwo
7000336 | 2022-10-06 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Service | Penggantians service set 2 tahun | Ready to use |  | Jarwo
7000339 | 2022-10-06 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0002 | Service | Kebocoran saat pasien bayi. Penggantian circuit dan breathing bag khusus bayi | Ready to use |  | Jarwo
7000337 | 2022-10-06 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0026 | Service | Penggantian service set 2 tahun | Ready to use |  | Jarwo
9000378 | 2022-10-06 | RSUD Dr. M. M. Dunda | ventilator | Draeger/Evita V300 | ASLD0169 | Survey | update software | Ready to use |  | Faugi
18000303 | 2022-10-05 | RSUD Tani dan Nelayan | Baby Incubator | Draeger/Isolette TI500 | KY11396 | Survey | General check | Ready to use | Disarankan penggantian baterai | Faugi
18000301 | 2022-10-05 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASJL0135 | Service | Pengecekan baterai | Ready to use | Disarankan penggantian baterai | Faugi
2000351 | 2025-10-08 | RSUP Jayapura | Digital Radiografi Fluoroskopi | Apelem/Platinum | F226F0736
19000498 | 2022-10-04 | RSUD Dr. Soeratno Gemolong | Jaudience Meter | Draeger/JM 105 | B3503803 | Service | Error 05 | Pending | Disarankan penggantian charger dock, adaptor | Nanda
2000144 | 2022-10-02 | RSUP Fatmawati | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0493 | Survey | Setting parameter drv dan parameter table | Ready to use |  | Eddy
18000496 | 2022-10-01 | RSJ Ratumbuysang Kalasey | Bed Patient | Paramount Bed | PRO89IL | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi
6000668 | 2022-10-01 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASJF0034 | Survey | General check | Ready to use | Dianjurkan penggantian service set 6 tahun | Afrizal
6000668 | 2022-10-01 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASMH0115 | Survey | General check | Ready to use | Dianjurkan ganti service set 3 tahun, waterlock | Afrizal
6000668 | 2022-10-01 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0039 | Maintenance | General check | Ready to use |  | Afrizal
6000668 | 2022-10-01 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASNB0071 | Service | General check | Ready to use |  | Afrizal
16000200 | 2022-09-30 | RS Abdi Waluyo | ESMR | Medispec/ED1000 | 0909876 | Survey | Selang getas, roda rusak, pengecekan secara umum | Pending | Disarankan penggantian spare part | Ipan
6000667 | 2022-09-30 | RSU PKU Muhammadiyah Purbalingga | Mesin Anasthesia | Draeger/Fabius Tiro | ARZM0025 | Survey | Pemasangan maintenance kit, Fabius service set 1 tahun | Ready to use |  | Afrizal
20000209 | 2022-09-30 | RSUD Syekh Yusuf | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0024 | Service | O2 sensor fail, flow sensor fail | Pending | Disarankan penggantian service set 3thn | Sujana
20000209 | 2022-09-30 | RSUD Syekh Yusuf | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0028 | Service | O2 sensor fail, flow sensor fail | Pending | Disarankan penggantian service set 3thn | Sujana
7000310 | 2022-09-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Service | Penggantian flow meter | Ready to use |  | Jarwo
7000306 | 2022-09-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Survey | Ada kebocoran jika digunakan untuk pasien bayi. Pasang dengan circuit bayi | Ready to use |  | Jarwo
7000307 | 2022-09-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Survey | Ada kebocoran jika digunakan untuk pasien bayi. Pasang dengan circuit bayi | Ready to use |  | Jarwo
4000345 | 2022-09-29 | RS Bhayangkara Jambi | Mobile X-ray | Apelem/Rafale EV30 | 005/19/01531 | Survey | Pelaksanaan ukes | Ready to use |  | Bayu
6000666 | 2022-09-29 | RSUD Banyumas | C-arm | Apelem/Evo R+ | 55116055585 | Service | Pendampingan ukes LPFK Solo | Ready to use |  | Afrizal
18000358 | 2022-09-29 | RSUD Toto Kabila | X-ray Digital | Apelem/Camargue DR | CPD327474K18 | Service | Hasil print hitam | Ready to use |  | Irfan R
7000309 | 2022-09-28 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Survey | EtCO2 tinggi, setting ulang dan penggantian waterlock | Ready to use |  | Jarwo
8000498 | 2022-09-28 | RSI Harapan Anda | Ventilator | Draeger/Savina 300 | ARZF0049 | Service | TV tidak stabil | Ready to use |  | Amna
8000498 | 2022-09-28 | RSI Harapan Anda | Ventilator | Draeger/Savina 300 | ASFJ0001 | Service | General check | Ready to use |  | Amna
19000497 | 2022-09-28 | RSUD Pandan Arang | Incubator Transport | Draeger/Isolette TI500 | BH12345 | Service | Penggantian power board dan baterai | Ready to use |  | Nanda
7000308 | 2022-09-28 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMN0118 | Survey | EtCO2 tinggi, setting ulang dan penggantian waterlock | Ready to use |  | Jarwo
6000700 | 2022-09-27 | RSUD Banyumas | C-arm | Apelem/Evo R+ | 55116055585 | Service | Gambar X-ray tidak keluar | Ready to use |  | Afrizal
19000496 | 2022-09-27 | RSUD Merah Putih | X-Ray Analog | Apelem/Camargue Analog | W0E003312 | Service | knop kolimator seret, dilakukan pelumasan | Ready to use |  | Nanda
9000373 | 2022-09-27 | RSUD Tambulilato | X-ray Digital | Apelem/Camargue DR | CPD36814320 | Service |  | Ready to use |  | Irfan R
4000344 | 2022-09-26 | Dinas Kesehatan Kab. Lombok Timur | Mobile X-ray | Apelem/Rafale EV30 | 1876552200001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2024-09-26
20000207 | 2022-09-26 | RSIA Pratiwi | Infant Warmer | Draeger/Resuscitaire | ASMM0036 | Service | Error 01, WT power rusak | Pending | Disarankan penggantian WT power | Sujana
1500206 | 2022-09-26 | RSUD Dr. M. Yunus | ESWL | Medispec/EM1000 | 0917011 | Service | console mcu dan emh sudah diperbaiki | Ready to use |  | Sutan
7000259 | 2022-09-26 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Caleo | ASKL0038 | Service | Error 51, penggantian keypad yg baru | Ready to use |  | Jarwo
20000206 | 2022-09-23 | RSUD Kota Makassar | Mesin Anasthesia | Draeger/Fabius Plus | ASJD0033 | Service | Pemasangan O2 turn switch | Ready to use |  | Sujana
17000148 | 2022-09-22 | RSUD Ulin | ESWL | Medispec/E3000 | 09131173 | Survey | Tidak bisa menembak. Pengecekan kabel smabungan antar board | Ready to use |  | Bahar
5000172 | 2022-09-21 | Klinik Ultimo | ESMR | Medispec/ED1000 | 09111368 | Survey | Pengecekan unit | Pending | Disarankan penggantian roda, applicator, dan kabel sensor. Sudah diajukan penawaran harga | Bilal
11000288 | 2022-09-21 | RS Muhammadiyah Makassar | Bubble CPAP | Draeger/CPap | QM20215173 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
19000494 | 2022-09-21 | RSUD dr. H. Soewondo Kendal | Infant Warmer | Draeger/Resuscitaire | CF10549 | Service | Error 05 | Pending | Heater dan board SSR harus diganti | Nanda
6000699 | 2022-09-21 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Babylog VN500 | ASEF0030 | Survey | Rusak total | Pending | Disarankan penggantian PATO, pressure regulator, baterai internal,basic board | Afrizal
5000171 | 2022-09-20 | RS Abdi Waluyo | EECP | Vasomedical/Lumenair | 5TC1034 | Service | Pengecekan unit | Ready to use | Disarankan penggantian cpu board dan kabel power. Sudah dilakukan penggantian | Bilal
17000147 | 2022-09-19 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0486 | Service | Penggantian kontaktor | Ready to use |  | Bahar
6000697 | 2022-09-19 | RSUD Dr. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CU03923 | Service | Pemasangan hinge access panel | Ready to use |  | Afrizal
6000697 | 2022-09-19 | RSUD Dr. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CU03921 | Survey | Pemasangan hinge access panel | Ready to use |  | Afrizal
14000205 | 2022-09-19 | RSUD Gambiran Kediri | C-arm | Apelem/Evo R+ | 55121001850 | Service | Alat tidak bisa nyala karna board rusak di gigit tikus (banyak kotoran hewan) | Pending | Sudah diajukan penawaran harga | Purwono
6000671 | 2022-09-19 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0073 | Service | Pemasangan pressure regulator, pressure sensor | Ready to use |  | Afrizal
18000488 | 2022-09-19 | RSUD Sanana | Dental Unit | Kavo/1058TM | 20007381 | Service | Perbaikan valve udara di port switch | Ready to use |  | Faugi
7000304 | 2022-09-17 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Service | Penggantian HVPS | Ready to use |  | Jarwo
14000203 | 2022-09-16 | RSUD Padangan | UPS | Yoshiga/GPI33 | 50120110113 | Maintenance | alat mati total, board seluruh unit mati, terdapat banyak kotoran hewan | Pending |  | Purwono
14000202 | 2022-09-16 | RSUD Padangan | UPS | Yoshiga/GPH33100 | 500190109095 | Service | General check | Ready to use |  | Purwono
15000203 | 2022-09-16 | RSUD Raden Mattaher | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/026/C0 | Service | Grid rusak | Pending | Disarankan penggantian grid | Sutan
14000204 | 2022-09-16 | RSUD Dr. R. Sosodoro Djatikoesoemo Bojonegoro | Dental Unit | Kavo/E30 | 10111154 | Survey | General check | Ready to use |  | Purwono
15000202 | 2022-09-15 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Service | Grid nyangkut, komputer detektor rusak | Pending |  | Sutan
9000370 | 2022-09-14 | RS Manembo-nembo Bitung | Bed Patient | Paramount Bed | 9003135 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
8000496 | 2022-09-14 | RSUD RA Kartini | Ventilator | Draeger/Savina 300 | ASEK0197 | Survey | Baterai tidak berfungsi | Pending | Disarankan penggantian baterai | Amna
15000201 | 2022-09-14 | RSUD Raden Mattaher | X-Ray Analog | Apelem/Camague Analog | IMA703863 | Service | Spring roll putus | Ready to use | Disarankan penggantian spring roll | Sutan
19000491 | 2022-09-14 | RSUD Tidar Magelang | Resuscitaire | Draeger/Resuscitaire | DSPF0064 | Service | Penggantian heater | Ready to use |  | Nanda
9000383 | 2022-09-14 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASFL0313 | Survey | pengecekan power supply | Pending |  | Faugi
19000490 | 2022-09-13 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Draeger/Fabius Plus | ARZM00051 | Maintenance | Pemasangan maintenance kit | Ready to use |  | Nanda
19000490 | 2022-09-13 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Draeger/Fabius Plus | ASHD0007 | Maintenance | Pemasangan maintenance kit | Ready to use |  | Nanda
7000303 | 2022-09-13 | RS Pratama Kelingi | Ventilator | Draeger/Babylog VN800 | ASRD0041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
11000287 | 2022-09-13 | RSUD Andi Makassau | Resuscitaire | Draeger/Resuscitaire | HR12028 | Maintenance | Penggantian heater | Ready to use |  | Nasrul
7000344 | 2022-09-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMM0007 | Service | Penggantian service set 2 tahun | Ready to use |  | Jarwo
7000341 | 2022-09-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASME0234 | Service | Penggantian service set 2 tahun | Ready to use |  | Jarwo
8000495 | 2022-09-12 | RSUD Dr. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | ASLJ0092 | Maintenance | General maintenance | Ready to use | Disarankan penggantian service set 3 tahun | Amna
8000495 | 2022-09-12 | RSUD Dr. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0041 | Maintenance | General maintenance | Ready to use | Disarankan penggantian service set 3 tahun | Amna
8000495 | 2022-09-12 | RSUD Dr. Soeprapto Cepu | Ventilator | Draeger/Savina 300 | ASJK0211 | Maintenance | General maintenance | Ready to use | Disarankan penggantian service set 3 tahun | Amna
7000273 | 2022-09-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Service | Penggantian service set 2 tahun | Ready to use |  | Jarwo
7000335 | 2022-09-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Maintenance | Preventive Maintenance (Total care) | Ready to use |  | Jarwo
7000302 | 2022-09-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASPL0087 | Service | Penggantian monitor | Ready to use |  | Jarwo
3000231 | 2022-09-09 | RSUD Tambulilato | X-ray Digital | Apelem/Camargue DR | CPD36814320 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2024-09-09
8000494 | 2022-09-09 | RSI Sunan Kudus | Ventilator | Draeger/Savina 300 | ARUE0022 | Service | FIO2 dan filter tidak berfungsi, VT Low, battery fail | Pending | Disarankan dilakukan penggantian oxytrace, service set 6 tahun, flow sensor | Amna
8000494 | 2022-09-09 | RSI Sunan Kudus | Ventilator | Draeger/Savina 300 | ARUM0308 | Survey | FIO2 dan filter tidak berfungsi, VT Low, battery fail | Pending | Disarankan dilakukan penggantian oxytrace, service set 6 tahun, flow sensor | Amna
6000693 | 2022-09-09 | RSUD dr Azhari | Mesin Anasthesia | Draeger/Primus | ASMA0131 | Survey | General check | Ready to use | Disarankan penggantian engsel breathing bag, maintenance primus service set 3thn, maintenance kit 2 tahun | Afrizal
6000693 | 2022-09-09 | RSUD dr Azhari | Mesin Anasthesia | Draeger/Primus | ASMF0114 | Survey | General check | Ready to use | Dianjurkan ganti maintenance kit 2thn | Afrizal
6000693 | 2022-09-09 | RSUD dr Azhari | Mesin Anasthesia | Draeger/Primus | ASLD0025 | Maintenance | General check | Ready to use | Disarankan penggantian primus 3tahun, maintenance kit, breathing circuit | Afrizal
6000694 | 2022-09-09 | RSI Harapan Anda | Phototherapy | Draeger/Bililux | 111352020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
17000146 | 2022-09-08 | RSUD Gabriel Manek | Digital Radiografi Fluoroskopi | Apelem/Platinum | E226F0723 | Survey | Pelaksanaan Ukes dan Upar | Ready to use |  | Bahar
7000301 | 2022-09-07 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMN0118 | Service | Penggantian service set 2 tahun | Ready to use |  | Jarwo
7000272 | 2022-09-07 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Service | Penggantian service set 2 tahun | Ready to use |  | Jarwo
8000493 | 2022-09-07 | RSUD Muntilan | ESMR | Draeger/Perseus A500 | ASRH0184 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
15000200 | 2022-09-06 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Survey | penggantian applicator | Ready to use | Disarankan penggran module control electronic mainboard, valve block brake arm, g06 connector | Sutan
8000492 | 2022-09-05 | RS Mitra Bangsa Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0041 | Maintenance | penggantian O2 sensor | Ready to use |  | Amna
18000497 | 2022-09-05 | RSUD Ainun Habibie | Infant Warmer | Draeger/Resuscitaire | IB22653 | Service | Pengecekan masukan volt tegangan ke kompressor | Pending | Harus dilakukan perbaikan heater | Faugi
4000346 | 2022-09-04 | RSUD Banyumas | C-arm | Apelem/Evo R+ | 55116055585 | Service | Uji paparan ruangan | Ready to use |  | Bayu
17000145 | 2022-09-02 | RSUD Dr. M. Yunus | ESWL | Medispec/EM1000 | 0917011 | Service | Penggantian console dan HVPS | Ready to use | EMH dibawa untuk diperbaiki | Bahar
4000343 | 2022-09-02 | RSUP Dr. Hasan Sadikin | Jaudience Meter | Medispec/ESMR | 06111184 | Service | Penggantian applicator | Ready to use |  | Bayu
7000271 | 2022-09-02 | RSUP dr. Soeradji Tirtonegoro | Jaudience Meter | Draeger/JM 105 | B3502776 | Service | Pengembalian alat setelah di kalibrasi | Ready to use |  | Afrizal
9000367 | 2022-09-01 | RSUD Maria Walanda Maramis | Ventilator | Draeger/Babylog VN500 | ASPF0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
3000230 | 2022-09-01 | RSUD Banyumas | C-arm | Apelem/Evo R+ | 55116055585 | Service | Penggantian sigma board | Ready to use |  | Imam
18000499 | 2022-08-31 | RSUD Ainun Habibie | UPS | Yoshiga | 1100315 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi | 2023-08-31
9000366 | 2022-08-30 | RS Siloam Paal Dua | Bubble CPAP | Draeger/CPap | QM20190223 | Service | Penggantian flow meter | Ready to use |  | Irfan R
6000691 | 2022-08-29 | RS PKU Purbalingga | Mesin Anasthesia | Draeger/Fabius tiro | ARZM0025 | Service | Pemasangan O2 sensor capsule, blue set anasthesia P/A, spirolog flow sensor | Ready to use |  | Afrizal
8000491 | 2022-08-29 | RSUD Dr. R. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | ASLJ0092 | Survey | Penggantian O2 sensor | Ready to use |  | Amna
18000498 | 2022-08-29 | RSUD Samratulangi Tondano | Dental Unit | Kavo/1058TM | 10205322 | Service | penggantian suction | Ready to use |  | Faugi
3000229 | 2022-08-26 | RSUD Gabriel Manek | Digital Radiografi Fluoroskopi | Apelem/Platinum | E226F0723 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2024-08-26
11000283 | 2022-08-26 | RS Muhammadiyah Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF1513 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000283 | 2022-08-26 | RS Muhammadiyah Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF1517 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000281 | 2022-08-26 | RS Muhammadiyah Makassar | Mesin Anasthesia | Draeger/Atlan A350 | ASPE0199 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000281 | 2022-08-26 | RS Muhammadiyah Makassar | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0127 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000282 | 2022-08-26 | RS Muhammadiyah Makassar | Ventilator | Draeger/Savina 300 | ASPH0298 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
6000689 | 2022-08-26 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASHM0015 | Survey | mixer ventilator tidak berfungsi | Pending | Disarankan penggantian Pba CPU Primus, waterlock dan sample line | Afrizal
6000690 | 2022-08-26 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0015 | Survey | O2 sensor fail | Ready to use | Disarankan penggantian O2 sensor, fabius service set, cozy service set dan flow sensor | Afrizal
6000690 | 2022-08-26 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0057 | Survey | O2 sensor fail | Ready to use | Disarankan penggantian O2 sensor capsule, fabius servoce set 3 tahun, cozy service set | Afrizal
6000690 | 2022-08-26 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0058 | Survey | O2 sensor fail | Ready to use | Disarankan penggantian O2 sensor capsule, fabius servoce set 3 tahun, cozy service set | Afrizal
6000690 | 2022-08-26 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASHN0015 | Survey | selft test, running test | Ready to use | Dianjurkan penggantian perawatan berkala | Afrizal
6000687 | 2022-08-26 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHL0156 | Survey | selft test, running test | Ready to use | Disarankan penggantian pressure regulator, baterai PS500, baterai internal, kit air filter, O2 filter particle, membran complete | Afrizal
6000687 | 2022-08-26 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08777 | Survey | selft test, running test | Ready to use | Disarankan penggantian filter, engsel depan kanan dan engsel pintu 4pcs | Afrizal
6000687 | 2022-08-26 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASHA0100 | Survey | selft test, running test | Ready to use | Dianjurkan penggantian pressure regulator, baterai internal, karet expiratory valve dan kit air filter | Afrizal
6000687 | 2022-08-26 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08779 | Survey | selft test, running test | Ready to use | Disarankan penggantian filter, engsel depan kanan dan engsel pintu 4pcs | Afrizal
6000687 | 2022-08-26 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08776 | Survey | selft test, running test | Ready to use | Disarankan penggantian filter, engsel belakang kanan, engsel depan kanan dan engsel pintu 4pcs | Afrizal
11000282 | 2022-08-26 | RS Muhammadiyah Makassar | Ventilator | Draeger/Savina 300 | ASPH0299 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
9000386 | 2022-08-25 | RS Siloam Paal Dua | Bubble CPAP | Draeger/CPAP | QM20190223 | Maintenance | Kerusakan pada mixer bubble cpap | Ready to use |  | Faugi
6000688 | 2022-08-25 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASEF0108 | Survey | self test, running test | Ready to use | Dianjurkan penggantian pressure regulator, kit air filter, O2 particle filter, membran complete, filter pad | Afrizal
6000688 | 2022-08-25 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | DSNF0062 | Survey | self test, running test | Ready to use | Dianjurkan perawatan berkala (filter dan engsel pintu) | Afrizal
6000688 | 2022-08-25 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | AF32541 | Survey | self test, running test | Ready to use | Dianjurkan perawatan filter & penggantian engsel pintu belakang kanan, depan kiri dan pintu 4pcs | Afrizal
7000270 | 2022-08-25 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Survey | General maintenance | Ready to use |  | Jarwo
6000686 | 2022-08-24 | RSUI Harapan Anda | Ventilator | Draeger/Evita 4 edition | ASCC0210 | Service | Pemasangan O2 snesor | Ready to use |  | Afrizal
6000685 | 2022-08-24 | RSUI Harapan Anda | Ventilator | Draeger/Babylog VN500 | ASMK0055 | Service | Pemasangan baterai internal dan neonatal flow sensor cable | Ready to use |  | Afrizal
16000199 | 2022-08-24 | Poli Gigi Klinik DPRD | Dental Unit | Kavo/1058 Lite | 10206373 | Service | general check | Ready to use |  | Ipan
19000307 | 2022-08-19 | RSUD dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0049 | Survey | mv rendah | Ready to use |  | Nanda
7000266 | 2022-08-18 | RSUD Sleman | Jaudience Meter | Draeger/JM 105 | B3503649 | Survey | Penggantian baterai | Ready to use |  | Jarwo
9000365 | 2022-08-18 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0481 | Survey | Unit sering restart sendiri | Pending |  | Irfan R
16000198 | 2022-08-18 | RS Dr. Suyoto | Dental Unit | Kavo/Primus 1055 | 10207193 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan | 2023-08-18
16000197 | 2022-08-18 | RS Dr. Suyoto | Dental Unit | Kavo/Primus 1058 | 10207194 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan | 2023-08-18
18000495 | 2022-08-16 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120 | VGSNF1696 | Service | EtCO2 tinggi, setting ulang | Ready to use |  | Faugi
7000265 | 2022-08-16 | RSUP dr. Soeradji Tirtonegoro | Jaudience Meter | Draeger/JM 105 | B3502776 | Service | error calibration | Pending | Unit dikirim ke Draeger untuk di kalibrasi | Jarwo
7000269 | 2022-08-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Service | sensor gas tidak sesuai, penggantian ILCA | Ready to use |  | Jarwo
18000494 | 2022-08-15 | RSUD Ainun Habibie | Infant Warmer | Draeger/Resuscitaire | DY0348 | Service | Setting ulang regulator | Ready to use |  | Faugi
6000683 | 2022-08-15 | RSUD Hj. Anna Lasmanah | Ventilator | Draeger/Evita V300 | ASHL0169 | Maintenance | Pemasangan baterai | Ready to use |  | Afrizal
3000227 | 2022-08-15 | RSUD Kab. Bekasi | C-arm | Apelem/Evo R+ | 55118026644 | Service | Penggantian SSD & Kalibrasi detector | Ready to use |  | Imam
4000341 | 2022-08-12 | RSK Dharmais | Digital Radiografi Fluoroskopi | Apelem/Platinum | L216F0664 | Maintenance | general check | Ready to use |  | Bayu
11000279 | 2022-08-12 | RSKD IBU DAN ANAK PERTIWI | Resuscitaire | Draeger | HC13575 | Service | error 05, heater tidak panas | Pending | Disarankan penggantian heater | Nasrul
15000199 | 2022-08-12 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Survey | Penggantian kabel extension karna joystick console error | Ready to use |  | Sutan
18000493 | 2022-08-12 | RSUD Toto Kabila | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150173 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi | 2024-08-12
11000278 | 2022-08-11 | RSI Faisal Makassar | ESWL | Medispec/EM1000 | 1114075 | Survey | EMH tidak bisa menembak | Pending | Disarankan penggantian EMH dan HVG | Nasrul
7000264 | 2022-08-10 | RSPAU dr. S. Hardjolukito | Ventilator | Draeger/Savina 300 | ASDL0131 | Survey | PC console mati | Pending | Disarankan penggantian service set 2 tahun | Jarwo
7000261 | 2022-08-10 | RSPAU dr. S. Hardjolukito | Ventilator | Draeger/Savina 300 | ASAC0044 | Survey | General check | Pending | Disarankan penggantian service set 2 tahun & O2 sensor | Jarwo
7000262 | 2022-08-10 | RSPAU dr. S. Hardjolukito | Ventilator | Draeger/Savina 300 | ASDL0132 | Survey | General check | Pending | Disarankan penggantian service set 2 tahun & O2 sensor | Jarwo
7000263 | 2022-08-10 | RSPAU dr. S. Hardjolukito | Ventilator | Draeger/Savina 300 | ASDL0128 | Maintenance | General check | Pending | Disarankan penggantian service set 2 tahun & O2 sensor | Jarwo
13000132 | 2022-08-10 | RSU Haji Medan | Bone Mineral Densitometer | Medilink/Medix DR | D197H0560 | Survey | Keyboard error, loading lama | Ready to use |  | Sofian
8000497 | 2022-08-10 | RSUD Kayen Pati | IACS Monitoring | Draeger/IACS Standard | ASRJ0473 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000497 | 2022-08-10 | RSUD Kayen Pati | IACS Monitoring | Draeger/IACS Standard | ASRJ0450 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000497 | 2022-08-10 | RSUD Kayen Pati | IACS Monitoring | Draeger/IACS Standard | ASRJ0451 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000497 | 2022-08-10 | RSUD Kayen Pati | IACS Monitoring | Draeger/IACS Standard | ASRD0241 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000497 | 2022-08-10 | RSUD Kayen Pati | IACS Monitoring | Draeger/IACS Standard | ASRJ0452 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
18000492 | 2022-08-10 | RSUD Dr. M. M. Dunda | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/048/C0 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi | 2024-08-10
3000225 | 2022-08-10 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helianthus BYM | HELI/0023/CB0 | Survey | Pemindahan alat ke ruangan baru | Ready to use |  | Imam
3000226 | 2022-08-10 | RSUP Dr. Hasan Sadikin | USG | Esaote/MyLab30Gold | 04323 | Service | General check | Pending |  | Imam
8000489 | 2022-08-08 | RSUD Kayen Pati | Patient Monitor | Draeger/Delta XL | 6010357777 | Maintenance | penggantian baterai | Ready to use |  | Amna
9000362 | 2022-08-08 | RSUD Otohana | Ventilator | Draeger/Evita V300 | ASPA0007 | Maintenance | General check, function check | Ready to use |  | Irfan R
9000362 | 2022-08-08 | RSUD Otohana | Ventilator | Draeger/Evita V800 | ASPL0095 | Service | General check, function check | Ready to use |  | Irfan R
9000362 | 2022-08-08 | RSUD Otohana | Ventilator | Draeger/Evita V800 | ASPF0165 | Survey | General check, function check | Ready to use |  | Irfan R
3000224 | 2022-08-08 | RSUD Dr. Adjidarmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Survey | PC Console mati | Ready to use | Dibuatkan penawaran | Imam
13000131 | 2022-08-05 | RSUD Simeulue | Bone Mineral Densitometer | Medilink/Medix DR | F207H0702 | Service | Penggantian applicator | Ready to use |  | Sofian
19000487 | 2022-08-05 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Survey | O2 fail dan gas measurement fail | Pending | Tarik error | Nanda
7000258 | 2022-08-05 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMM0118 | Service | Pemindahan BMD ke ruangan baru | Ready to use |  | Jarwo
19000486 | 2022-08-04 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0145 | Service | Penggantian maintenance kit 2years dan O2 sensor housing | Ready to use |  | Nanda
9000361 | 2022-08-03 | RSUD Toto Kabila | ESMR | Apelem/Camargue DR | CPD327474K18 | Survey | Detektor tidak tersambung. Restart unit | Ready to use |  | Irfan R
15000197 | 2022-08-03 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Service | interface board sigma tidak konek | Ready to use |  | Sutan
3000223 | 2022-08-03 | RSUD Banyumas | C-Arm | Apelem/Evo R+ | 55116055585 | Service | Engsel patah | Ready to use | Spare part diretur karna belum di program | Imam
19000485 | 2022-08-02 | RSUD Dr. Soeratno Gemolong | Baby Incubator | Draeger/Isolette C2000 | AC37611 | Service | Error controller 11 failure | Pending | Disarankan penggantian board interface PCB Assembly | Nanda
8000488 | 2022-08-02 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKK0117 | Service | Penggantian baterai internal | Ready to use |  | Amna
7000255 | 2022-08-02 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Isolette 8000 | AC09111 | Survey | Hasil expose tidak bagus bersihkan grid dan kalibrasi detector | Ready to use | Masih bisa digunakan namun perlu penggantian | Jarwo
7000254 | 2022-08-02 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Isolette 8000 | AC09112 | Service | Engsel patah | Ready to use | Masih bisa digunakan namun perlu penggantian | Jarwo
11000275 | 2022-08-02 | RSUD Haji Makassar | Dental Unit | Kavo/1058TM | 10205316 | Service | perbaikan suction | Ready to use |  | Nasrul
18000491 | 2022-08-01 | RS Bunda Pengharapan | Bed Patient | Paramount Bed | 900505800002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi
9000360 | 2022-08-01 | RSUD KOTAMOBAGU | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/043/C0 | Survey | tampilan layar booting terus | Pending | Disarankan penggantian hardist dan reload software | Irfan R
19000484 | 2022-08-01 | RSUD Tidar Magelang | Baby Incubator | Draeger/Caleo | ASKB0078 | Maintenance | Maintenance total care | Ready to use |  | Nanda
19000484 | 2022-08-01 | RSUD Tidar Magelang | Baby Incubator | Draeger/Caleo | ASKB0076 | Maintenance | Maintenance total care | Ready to use |  | Nanda
19000484 | 2022-08-01 | RSUD Tidar Magelang | Baby Incubator | Draeger/Caleo | ASKB0077 | Maintenance | Maintenance total care | Ready to use |  | Nanda
19000483 | 2022-08-01 | RSUD Tidar Magelang | Infant Warmer | Draeger/Resuscitaire | DSPF0064 | Survey | Error 05 | Pending | Disarankan penggantian heater | Nanda
8000487 | 2022-07-31 | RSI Sultan Agung | Lampu Operasi | Draeger/Polaris 600 | ASMA0369 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000487 | 2022-07-31 | RSI Sultan Agung | Lampu Operasi | Draeger/Polaris 600 | ASMA0367 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
9000359 | 2022-07-29 | RSUD Ainun Habibie | Resuscitaire | Draeger/Resuscitaire | DY03434 | Maintenance | General Check | Ready to use |  | Irfan R
9000358 | 2022-07-29 | RSUD Ainun Habibie | Ventilator | Draeger/Evita V300 | ASDL0172 | Maintenance | General Check | Ready to use |  | Irfan R
8000486 | 2022-07-28 | RS Panti Wilasa | Ventilator | Draeger/Savina 300 | ASMM0026 | Service | General check, system check, running test | Ready to use |  | Amna
4000339 | 2022-07-28 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Service | General check | Ready to use |  | Bayu
13000128 | 2022-07-28 | RSUD Kota Langsa | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/057/C1 | Survey | Software lemot, Setting auto delete dan hapus data base pasien | Ready to use |  | Sofian
7000252 | 2022-07-28 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASKJ0043 | Service | Penggantian resistor break | Pending | Disarankan penggantian part MV1 dan MV2 | Jarwo
3000222 | 2022-07-28 | RSUD Tidar Magelang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0496 | Maintenance | Sensor check failed | Ready to use |  | Imam
7000253 | 2022-07-27 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
90000357 | 2022-07-27 | RSUD KOTAMOBAGU | C-arm | Apelem/Evo R+ | 55118031654 | Service | Setting menu windows, monitor ke-3 tidak tampil gambar | Ready to use |  | Irfan R
8000485 | 2022-07-27 | RSUD Kajen | Ventilator | Draeger/Babylog VN800 | ASRD0040 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
4000338 | 2022-07-27 | RSUP Fatmawati | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0493 | Service | Monitor ke-3 (samsung) tidak tampil gambar; Setting D1 menu windows | Ready to use |  | Bayu
8000484 | 2022-07-26 | RSU Fastabiq Sehat | Mesin Anasthesia | Draeger/Fabius Plus | ASRC0070 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
18000490 | 2022-07-26 | RS Mitra Masyarakat Papua | Stretcher | Paramount Bed | 187793000007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi
7000300 | 2022-07-26 | RSUD Sleman | Ventilator | Draeger/BabylogVN800 | ASRD0053 | Survey | Re-install software dan penggantian sigma board | Ready to use |  | Jarwo | 2024-07-26
6000682 | 2022-07-26 | RSUI Harapan Anda | Ventilator | Draeger/Evita 4 edition | AFHA0095 | Service | Pemasangan aksesoris expiratory valve | Ready to use |  | Afrizal
6000680 | 2022-07-25 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0057 | Survey | self test, running test | Ready to use | Disarankan penggantian O2 sensor dan service set 3 tahun | Afrizal
6000678 | 2022-07-25 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0050 | Survey | leak test gagal, O2 sensor expired | Pending | Disarankan penggantian O2 sensor dan selang dari central base mesin ke flow nasal | Afrizal
6000676 | 2022-07-25 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0056 | Survey | O2 sensor fail, selft test running test oke, baterai drop | Ready to use | Disarankan penggantian O2 sensor dan service set 3 tahun | Afrizal
6000677 | 2022-07-25 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASJE0005 | Service | self test alat oke | Ready to use | Disarankan penggantian waterlock dan sample line | Afrizal
6000679 | 2022-07-25 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASHM0015 | Service | mixer tidak berfungsi, self test mixer hijau | Ready to use | Disarankan penggantian waterlock dan sample line | Afrizal
8000483 | 2022-07-25 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
11000277 | 2022-07-25 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 0720017 | Service | Penggantian motor actuator | Ready to use |  | Nasrul
4000337 | 2022-07-25 | RSUP Fatmawati | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0493 | Survey | Re-install software dan penggantian sigma board | Ready to use |  | Bayu
6000675 | 2022-07-23 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0072 | Service | Pemasangan PATO sensor, board PCB AMO-MFM, service set 3 tahun | Ready to use | Penambahan perlu penggantian Kit LP AMO Flow ILCA | Afrizal
7000340 | 2022-07-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0002 | Maintenance | Preventive Maintenance (Total Care) | Ready to use |  | Jarwo
7000338 | 2022-07-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0026 | Maintenance | Preventive Maintenance (Total Care) | Ready to use |  | Jarwo
7000343 | 2022-07-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMH0007 | Maintenance | Preventive Maintenance (Total Care) | Ready to use |  | Jarwo
7000346 | 2022-07-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHL0059 | Maintenance | Preventive Maintenance (Total Care) | Ready to use |  | Jarwo
7000342 | 2022-07-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Maintenance | Preventive Maintenance (Total Care) | Ready to use |  | Jarwo
6000674 | 2022-07-22 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0073 | Service | Pressure regulator error | Ready to use | Back up pressure regulator dengan punya kantor | Afrizal
6000670 | 2022-07-22 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0072 | Service | Pemasangan service set 6 tahun, PATO, PCB | Ready to use |  | Afrizal
18000487 | 2022-07-22 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0167 | Survey | FIO2 tidak sesuai setting | Pending |  | Faugi
7000299 | 2022-07-22 | RSI Klaten | Dental Unit | Kavo/E30 | 10111152 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-07-22
9000354 | 2022-07-21 | RSUD Morotal | Bed Patient | Paramount | PA6918CEC01 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000354 | 2022-07-21 | RSUD Morotal | Bed Patient | Paramount | PA6415CBBABA | Service | Penggantian selang suction | Ready to use |  | Irfan R
9000354 | 2022-07-21 | RSUD Morotal | Bed Patient | Paramount | PB2102 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000354 | 2022-07-21 | RSUD Morotal | Bed Patient | Paramount | KA8261BF | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000354 | 2022-07-21 | RSUD Morotal | Bed Patient | Paramount | PB1110D | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000355 | 2022-07-21 | RSUD Morotai | Emergency Stretcher | Paramount | KA8261 BF | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
16000196 | 2022-07-21 | Klinik Drg. Hedy Soleman | Dental Unit | Kavo/Primus 1058 | 1002324 | Service | Penggantian PCB Dentist Element dan reload software | Ready to use |  | Ipan
11000273 | 2022-07-21 | RSKD Gigi dan Mulut | Dental Unit | Kavo/Primus 1058 | 10207187 | Service | Realod software, perbaikan motor grid dan penggantian SSD | Ready to use |  | Nasrul
4000336 | 2022-07-20 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2024-07-20
15000198 | 2022-07-20 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Service | Penggantian SSD dan upgrade software | Ready to use |  | Sutan
9000353 | 2022-07-19 | RSUD Weda | Bubble CPAP | Draeger | QM20215462 | Service | Pelumasan handpiece | Ready to use |  | Irfan R
9000353 | 2022-07-19 | RSUD Weda | Bubble CPAP | Draeger | QM20215463 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
16000195 | 2022-07-18 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan | 2023-07-18
19000479 | 2022-07-18 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/EM1000 | 0319201 | Service | Balon tidak mengembang maksimal, dilakukan pengurasan air dan penggantian air yang baru | Ready to use |  | Nanda
7000297 | 2022-07-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASPL0087 | Service | PU BOX mati, tombol emergency tertekan | Ready to use |  | Jarwo | 2024-07-15
7000297 | 2022-07-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASPL0088 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2024-07-15
19000478 | 2022-07-14 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0145 | Service | LPM tidak bisa naik - Setting knop dan adjusment | Ready to use |  | Nanda
15000196 | 2022-07-14 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Pengecekan PU Box, test pakai donggel ok, tombol emergency tertekan | Ready to use |  | Sutan
8000480 | 2022-07-13 | RS Aisyiah Jepara | Mesin Anasthesia | Draeger/Fabius Plus | ASND0058 | Service | General check, O2 sensor check | Ready to use |  | Amna
19000477 | 2022-07-13 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Isolette 8000 | MC09026 | Service | Penggantian engsel | Ready to use |  | Nanda
7000295 | 2022-07-13 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Isolette 8000 | ASKL0038 | Service | Penggantian board display colimator | Pending |  | Jarwo
7000294 | 2022-07-13 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Primus | AC09112 | Service | Penggantian heater dan wi2 board, namun keypad masih error | Ready to use |  | Jarwo
11000272 | 2022-07-12 | RSUD Labuang Baji | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0494 | Service | Display collimator error; Pemasangan board display collimator | Ready to use |  | Nasrul
19000476 | 2022-07-12 | RSUP Dr. Kariadi | Ventilator | Draeger/Isolette 8000 | ASML0094 | Survey |  | Pending | Disarankan penggantian PATO sensor | Nanda
19000476 | 2022-07-12 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0094 | Service | O2 calibration failed | Pending | Disarankan penggantian PATO sensor | Nanda
7000292 | 2022-07-12 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Babytherm 8010 | AC09111 | Service | O2 calibration failed | Ready to use |  | Jarwo
7000291 | 2022-07-12 | RSUP dr. Soeradji Tirtonegoro | Infant Warmer | Draeger/Caleo | ASMM0045 | Survey | Pemasangan engsel yang baru | Ready to use |  | Jarwo
8000479 | 2022-07-11 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Service | Penggantian power pack | Ready to use |  | Amna
6000578 | 2022-07-11 | RSUD Dr. Soedirman - Kebumen | ESMR | Draeger/Primus | ASKJ0050 | Service | Gas sensor dan baterai tidak berfungsi | Pending | Disarankan penggantian service set 3 tahun, 2/2 distributing valve. Sensor ILCA 3 dan power pack | Afrizal
15000195 | 2022-07-11 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Service | Penggantian applicator | Ready to use |  | Sutan
13000127 | 2022-07-07 | RSUD Kota Langsa | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/057/C1 | Service | Detektor berair | Ready to use | Sudah dikeringkan | Sofian
18000474 | 2022-07-07 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Perseus | ASKJ0081 | Service | Gas failure | Pending | Disarankan diganti Flow power dan ILCA | Irfan R
13000126 | 2022-07-06 | RSUD Dr. Fauziah | ESWL | Medispec/E3000 | 0716249 | Service | ESWL tidak bisa menembak | Ready to use | Isi air full EMH | Sofian
8000478 | 2022-07-05 | RSI Sultan Agung | Baby Incubator | Draeger/Isolette C2000 | DSNK0640 | Service | Penggantian filter bakteri | Ready to use | Disarankan penggantian O2 sensor | Amna
8000478 | 2022-07-05 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0126 | Service | Error RVI fresh gas valve | Pending | Disarankan penggantian piston valve | Amna
8000478 | 2022-07-05 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0409 | Service | Error RVI fresh gas valve | Pending | Disarankan penggantian piston valve | Amna
4000335 | 2022-07-05 | RSUD Alimuddin Umar | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0402 | Service | Tidak bisa print, sensor coilision tidak berfungsi | Ready to use |  | Bayu
3000220 | 2022-07-05 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | Back up data, re-install windows 10 dan reload software | Ready to use |  | Afrizal
18000458 | 2022-07-05 | RSUD dr. Zainal Umar Sidiki | Bed Patient | Paramount Bed | PD150 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
11000271 | 2022-07-05 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASJJ0408 | Service | Baterai internal tidak ada, VT tidak stabil | Pending | Disarankan penggantian service set 3 tahun | Nasrul
14000176 | 2022-07-04 | RSUD Dr. R. Sosodoro Djatikoesoemo Bojonegoro | Dental Unit | Kavo/E30 | 10111154 | Survey | footswitch bocor, ganti karet heksagonal | Ready to use |  | Purwono
6000577 | 2022-07-01 | RSUD Dr. R. Goeteng Taroenadibrata | Ventilator | Draeger/Evita V300 | ASMJ0011 | Maintenance | Pengecekan general | Ready to use |  | Afrizal
8000477 | 2022-07-01 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000476 | 2022-06-30 | RS Aisyiyah Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0048 | Service | General check | Ready to use |  | Amna
19000468 | 2022-06-30 | RSUD Dr. Soediran Mangun Sumarso | Ventilator | Draeger/Evita V300 | ARWL0203 | Service | Baterai tidak ada, O2 inop mixer bocor | Pending | Disarankan penggantian baterai internal, O2 sensor, bacterial filter, pressure regulator dan kabel flow sensor | Nanda
19000469 | 2022-06-30 | RSUD Dr. Soediran Mangun Sumarso | Resuscitaire | Draeger/Resuscitaire | BC13132 | Survey | Error 05 (Heater rusak) | Pending | Disarankan penggantian heater | Nanda
6000576 | 2022-06-30 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0073 | Service | General check | Pending | Disarankan penggantian pressure sensor 2pcs untuk O2 dan udara, pressure regulator 3pcs untuk O2, udara dan N2O. Sudah diberikan penawaran, proses PO | Afrizal
7000288 | 2022-06-30 | RSUP Dr. Sardjito | ESWL | Medispec/E3000 | 0719139 | Survey | C-arm tidak bisa expose | Ready to use | Power supply dan main board diambil untuk di cek | Jarwo
8000475 | 2022-06-29 | RS Bhina Bhakti Husada | Mesin Anasthesia | Draeger/Primus | ASPL0086 | Survey | General check. Penggantian water lock | Ready to use |  | Amna
7000283 | 2022-06-29 | RSUP Dr. Sardjito | Jaudience Meter | Draeger/JM 105 | H3505880 | Maintenance | Error time period calibration | Pending | Alat perlu dikalibrasi | Jarwo
7000284 | 2022-06-29 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMH0118 | Maintenance | General check (Service kontrak) | Ready to use |  | Jarwo |  | √
7000287 | 2022-06-29 | RSUP Dr. Sardjito | Ventilator | Draeger/Savina 300 | ASLD0041 | Service | General check (Service kontrak) | Ready to use |  | Jarwo |  | √
8000473 | 2022-06-28 | RST dr. Soedjono Magelang | Baby Incubator | Draeger/Isolette C2000 |  | Service
7000286 | 2022-06-28 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 |  | PEP pressure drop. Pindah ke outline lain | Ready to use |  | Jarwo
7000282 | 2022-06-24 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0128 | Maintenance | Penggantian service set 2 tahun | Ready to use |  | Jarwo
8000472 | 2022-06-24 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0039 | Maintenance | General check, system check, running test | Ready to use | Disarankan stock water trap dan sample line | Amna
8000472 | 2022-06-24 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0044 | Maintenance | General check, system check, running test | Ready to use | Disarankan stock water trap dan sample line | Amna
8000474 | 2022-06-24 | RST dr. Soedjono Magelang | Baby Incubator | Draeger/Isolette C2000 | DSRB0017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000474 | 2022-06-24 | RST dr. Soedjono Magelang | Baby Incubator | Draeger/Isolette C2000 | DSPM0035 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
7000342 | 2022-06-24 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Maintenance | General check (Service kontrak) | Ready to use |  | Jarwo |  | √
7000280 | 2022-06-24 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0113 | Service | General check (Service kontrak) | Ready to use |  | Jarwo |  | √
7000281 | 2022-06-23 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0144 | Survey | Pemasangan service set 2 tahun | Ready to use |  | Jarwo
6000572 | 2022-06-23 | RSUD Brebes | Ventilator | Draeger/Primus | ASHA0095 | Service | Pengecekan general | Pending | Disarankan penggantian baM13, batterypack, Kit air filter, particle folter, microfilter, battery P5000, dan filter pad. Sudah diberikan penawaran | Afrizal
19000467 | 2022-06-23 | RSUD Pandan Arang | Incubator Transport | Draeger/TI500 | BH12345 | Maintenance | PSU short, coba ganti komponen yang rusak = masih belum bisa | Pending | Penggantian power supply | Nanda
8000471 | 2022-06-22 | RSUD Kajen | Mesin Anasthesia | Draeger/Primus | ASMA1091 | Maintenance | General check, system check, running test | Ready to use |  | Amna
8000471 | 2022-06-22 | RSUD Kajen | Mesin Anasthesia | Draeger/Primus | ASLD0023 | Maintenance | General check, system check, running test | Ready to use |  | Amna
8000471 | 2022-06-22 | RSUD Kajen | Mesin Anasthesia | Draeger/Primus | ASHA0233 | Maintenance | General check, system check, running test | Ready to use |  | Amna
8000470 | 2022-06-21 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKK0117 | Service | General check, system check, running test | Ready to use |  | Amna
80000481 | 2022-06-20 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0120 | Survey | Penggantian PSU | Ready to use |  | Amna
8000469 | 2022-06-20 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Vapor 3000 | ASRB0246 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
3000219 | 2022-06-20 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Survey | Motor lifter blocking | Ready to use |  | Imam
19000463 | 2022-06-18 | RSUD Pandan Arang | Incubator Transport | Draeger/TI500 | BH12345 | Service | fuse mati, komponen di power supply rusak | Pending | Penggantian power supply | Nanda
3000218 | 2022-06-17 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | reload parameter DRV | Ready to use |  | Sutan
7000285 | 2022-06-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Evita V300 | ASKF0069 | Service | Vaporizer tidak lolos kalibrasi. Setting dan adjust vaporizer | Ready to use |  | Jarwo
16000193 | 2022-06-16 | Kesumah Dental Clinic | Dental Unit | Durr/Tornado I | C415137069 | Maintenance | Pemindahan kompressor | Ready to use |  | Ipan
8000468 | 2022-06-16 | RSI Sunan Kudus | ESMR | Draeger/Atlan A350 | ASNL0042 | Service | General check | Ready to use |  | Amna
15000193 | 2022-06-16 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Survey | Penggantian applicator | Ready to use |  | Jarwo
1100269 | 2022-06-16 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASNL0247 | Service | Alat mati total, cek pcb main, pba, power supply, pressure regulator | Pending | Disarankan penggantian PBA Power board dan pressure regulator | Nasrul
20000156 | 2022-06-15 | RSUD Sayang Rakyat | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0403 | Service | Penggantian resistor break | Ready to use |  | Nasrul
7000278 | 2022-06-14 | RSKBD Panti Nugroho | UPS | ICA/SE3100 | IHIKT1600545 | Service | Penggantian baterai | Ready to use |  | Afrizal
7000279 | 2022-06-14 | RSKBD Panti Nugroho | UPS | Uni Power/3RUAS | 900031711251103 | Survey | Penggantian baterai | Ready to use |  | Jarwo
6000573 | 2022-06-11 | RSUD Dr. M. Ashari | Mesin Anasthesia | Draeger/Primus | ASMA0131 | Survey | general check | Ready to use | Disarankan penggantian waterlock, maintenance service set 2thn, knob breathing bag | Ipan
6000574 | 2022-06-11 | RSUD Dr. M. Ashari | Mesin Anasthesia | Draeger/Primus | ASLD0050 | Survey | general check | Ready to use | Disarankan penggantian waterlock, maintenance service set 2thn, knob breathing bag | Afrizal
6000575 | 2022-06-11 | RSUD Dr. M. Ashari | Mesin Anasthesia | Draeger/Primus | ASMF0114 | Survey | General check | Ready to use | Disarankan penggantian waterlock 1 box, maintenance service set 2thn | Afrizal
18000451 | 2022-06-10 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina 300 | ASPJ0086 | Service | Refresh training ventilator | Ready to use |  | Irfan R
19000465 | 2022-06-10 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | Auto repair pada PC user | Ready to use | Disarankan ganti komputer dan OS baru | Nanda
19000461 | 2022-06-10 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Fabius Plus | ASJM0264 | Survey | layar blinking, O2 tidak bisa. Housing O2 dan connector rusak | Pending | Disarankan penggantian konektor layar LCD dan housing O2 sensor | Nanda
18000454 | 2022-06-10 | RSUD Maria Walanda Maramis | Ventilator | Draeger/Savina 300 | ASNL0878 | Service | Refresh training ventilator | Ready to use |  | Irfan R
17000144 | 2022-06-09 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Maintenance | Perbaikan koneksi selang dengan handpiece | Ready to use |  | Jarwo
8000467 | 2022-06-09 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASLD0028 | Maintenance | General check | Ready to use | Disarankan penggantian service set 3 tahun | Amna
16000192 | 2022-06-09 | Poli Gigi Klinik DPRD | Dental Unit | Kavo/1058 Lite | 10206373 | Maintenance | General check | Ready to use |  | Bahar
7000290 | 2022-06-08 | RSIA Harapan Ibu | Bubble CPAP | Draeger/Oximixer | QM20215509 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-06-08
8000466 | 2022-06-08 | RSUD dr. R. Soetijono Blora | Mesin Anasthesia | Draeger/Fabius Plus | ASJD0085 | Maintenance | general check | Ready to use | Disarankan penggantian O2 sensor dan flow sensor | Amna
8000466 | 2022-06-08 | RSUD dr. R. Soetijono Blora | Mesin Anasthesia | Draeger/Fabius Plus | ASJD0084 | Survey | general check | Ready to use | Disarankan penggantian O2 sensor dan flow sensor | Amna
7000276 | 2022-06-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMF0234 | Survey | Cek BHP, Cek flow sensor, cek waterlock | Ready to use |  | Jarwo
7000277 | 2022-06-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Survey | Cek BHP, Cek flow sensor, Cek water lock | Ready to use |  | Jarwo
7000275 | 2022-06-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Maintenance | Cek BHP, cek flow sensor, cek waterlock | Ready to use |  | Jarwo
8000463 | 2022-06-07 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 | ASMN0162 | Maintenance | General check | Ready to use |  | Amna
8000464 | 2022-06-07 | RSI Sultan Hadlirin | Mesin Anasthesia | Draeger/Fabius Plus | ASND0059 | Service | General check | Ready to use |  | Amna
18000461 | 2022-06-07 | RSUD Datoe Binangkang | X-ray Digital | Apelem/Camargue DR | CPD33569E19 | Service | Kabel sambungan tube terbakar karna listrik tidak stabil, x-ray tidak bisa expose ; Penggantian kabel sambungan tube | Ready to use |  | Irfan R
19000458 | 2022-06-07 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Caleo | ASJE0010 | Service | Penggantian water heater | Ready to use |  | Nanda
8000465 | 2022-06-07 | RSUD RA Kartini | UPS | Yoshiga/HPT2 | 501111003780N3300002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2023-06-07
8000465 | 2022-06-07 | RSUD RA Kartini | UPS | Yoshiga/HPR 1 | 501101206060KA100042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2023-06-07
8000462 | 2022-06-06 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Perseus | ASPJ0088 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2023-06-06
19000457 | 2022-06-06 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0129 | Maintenance | General check | Ready to use |  | Nanda
8000462 | 2022-06-06 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Perseus | ASRA0027 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2023-06-06
7000251 | 2022-06-06 | RS PKU Muhammadiyah Gamping (KSO) | CT Scan | Phillips/Ever exced | 320121 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
11000270 | 2022-06-06 | RSUD Kota Makassar | Mesin Anasthesia | Draeger/Perseus A500 | ASJD0033 | Survey | O2 turn switch rusak | Pending | Disarankan penggantian O2 turn switch dan service set 3tahun | Nasrul
6000670 | 2022-06-06 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius | ARWC0079 | Service | Adjust regulator O2 | Ready to use |  | Afrizal
7000274 | 2022-06-06 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Service | Pengecekan self test, volume tidal OK | Ready to use |  | Sutan
7000289 | 2022-06-05 | RSUD Dr. Soedirman - Kebumen | ESWL | Draeger/Isolette C2000 | AH36596 | Service | Penggantian engsel | Ready to use |  | Jarwo
15000189 | 2022-06-05 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/EM1000 | 1116095 | Service | Penggantian EMH dan kalibrasi unit | Ready to use |  | Sutan
15000192 | 2022-06-03 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Ganti fiber optic karna kabel fiber optic ke PU Box putus | Ready to use |  | Ipan
19000455 | 2022-06-02 | RSUD dr. Gunawan Mangunkusumo | Central Monitor | Draeger/Vista 120 | VLSMF0152 | Survey | General check | Ready to use |  | Nanda
16000191 | 2022-06-02 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Maintenance | handpiece sering nyala mati sendiri | Ready to use |  | Imam
3000217 | 2022-05-31 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | pemasangan unit power supply dan HVG | Ready to use |  | Jarwo
18000479 | 2022-05-31 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 0920036 | Service | Kalibrasi | Ready to use |  | Faugi
7000349 | 2022-05-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMH0007 | Survey | EtCO2 tinggi, setting ulang | Ready to use |  | Sutan
19000454 | 2022-05-27 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0124 | Survey | Training | Ready to use |  | Nanda
19000454 | 2022-05-27 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Primus | ASPH0125 | Service | Training | Ready to use |  | Nanda
15000191 | 2022-05-27 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Maintenance | Pembersihan board allegro karna aplikasi duet tidak terbuka | Ready to use |  | Afrizal
6000571 | 2022-05-24 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Perseus A500 | ASJE0005 | Service | Pasang spare part MV2 dan perbaiki koneksi sensor | Ready to use |  | Afrizal
8000461 | 2022-05-24 | RSUD Sunan Kalijaga | UPS | Yoshiga/HPT10 | 501111003790N3300001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2023-05-24
8000461 | 2022-05-24 | RSUD Sunan Kalijaga | UPS | Yoshiga/HPT10 | 50111003790N3300002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2023-05-24
18000478 | 2022-05-24 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0178 | Service | Kalibrasi ulang | Ready to use |  | Faugi
7000348 | 2022-05-21 | RSUD Bagas Waras | Resuscitaire | Draeger/Resuscitaire | JY03385 | Service | Penggantian heater unit | Ready to use |  | Jarwo
8000460 | 2022-05-20 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASKF0024 | Service | Pemasangan service set 3 tahun | Ready to use |  | Amna
6000568 | 2022-05-20 | RSU An Ni'mah | Bubble CPAP | Draeger/CPAP | QM20215459 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
19000452 | 2022-05-20 | RSUP Dr. Kariadi | Incubator Transport | Draeger/Isolette C2000 | CR35569 | Survey | sensor disconnected | Ready to use | Disarankan penggantuan kabul sensor modul | Nanda
11000268 | 2022-05-20 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Perseus A500 | ASNK0060 | Service | Isofluerence bocor | Pending | Disarankan penggantian cap bayonet lock dan filling system | Nasrul
8000458 | 2022-05-19 | RS Mitra Bangsa Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0041 | Maintenance | General check, flow sensor - O2 sensor leak | Ready to use |  | Amna
13000124 | 2022-05-19 | RS Mitra Medika Premiere | Laser Surgical | Jena Surgical/MultiPulse HoPlus | J191150174 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sofian | 2024-05-19
16000190 | 2022-05-19 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Maintenance | supply press bocor | Ready to use |  | Imam
6000567 | 2022-05-18 | RSUD Dr. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Primus | ASKJ0050 | Service | self test gas analyzer berwarna kuning | Ready to use | Disarankan penggantian sensor ILCA3 - alat bisa digunakan tanpa pembacaan EtCo2 dan N2O dari pasien | Afrizal
3000213 | 2022-05-18 | RSUD Provinsi NTB | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/032/C0 | Service | Upgrade software dan kalibrasi AEC | Ready to use |  | Eddy
2000143 | 2022-05-17 | RSUD Dr. Adjidarmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Service | Reprogram parameter DRV2 collom & pemasangan motor servo | Ready to use |  | Sutan
18000460 | 2022-05-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | ASJC0048 | Maintenance | Pendampingan pengecekan kontrak service | Ready to use |  | Irfan R
18000460 | 2022-05-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | ASJC0105 | Maintenance | Pendampingan pengecekan kontrak service | Ready to use |  | Irfan R
18000460 | 2022-05-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | ASJD0020 | Maintenance | Pendampingan pengecekan kontrak service | Ready to use |  | Irfan R
18000460 | 2022-05-17 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | ASJD0019 | Maintenance | Pendampingan pengecekan kontrak service | Ready to use |  | Irfan R
18000460 | 2022-05-17 | RSUP Prof. Dr. R. D. Kandou Manado | ESMR | Apelem/Platinum | ASEK0109 | Service | Pendampingan pengecekan kontrak service | Ready to use |  | Irfan R
15000190 | 2022-05-13 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Survey | Penggantian applicator | Ready to use |  | Bilal
8000459 | 2022-05-12 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Survey | Power supply system fail | Pending | Disarankan penggantian power peak (kontrak service) | Amna
19000451 | 2022-05-12 | RSUD Pandan Arang | Ventilator | Draeger/Isolette C2000 | DSPL0024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000451 | 2022-05-12 | RSUD Pandan Arang | Ventilator | Draeger/Isolette C2000 | DSPL0027 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000451 | 2022-05-12 | RSUD Pandan Arang | Ventilator | Draeger/Isolette C2000 | DSPL0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000451 | 2022-05-12 | RSUD Pandan Arang | Ventilator | Draeger/Isolette C2000 | DSPL0029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000451 | 2022-05-12 | RSUD Pandan Arang | Ventilator | Draeger/Isolette C2000 | DSPM0033 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
19000451 | 2022-05-12 | RSUD Pandan Arang | Ventilator | Draeger/Babylog VN500 | DSPM0034 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda
5000170 | 2022-05-12 | RSUD Malingping | Dental Unit | Kavo/1058TM | 10205319 | Service | Pemasangan monitor | Ready to use |  | Afrizal
6000566 | 2022-05-12 | RSUD Kardinah Tegal | Ventilator | Draeger/Isolette C2000 | ASHL0153 | Service | alat mati total | Ready to use | Service board error dan diganti milik unit lain | Imam
4000334 | 2022-05-11 | RSUD Dr. M. Yunus | ESWL | Medispec/EM1000 | 0917011 | Survey | Error 11 | Ready to use | HVPS dibawa ke jakarta untuk di cek | Bayu
3000212 | 2022-05-11 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | HVG pecah dan power supply mengeluarkan api | Ready to use | Sementara HVG dan Power supply dibawa untuk diperbaiki. Sudah diperbaiki | Eddy
2000142 | 2022-05-11 | RSUP Fatmawati | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0493 | Survey | Data loading lemot | Ready to use |  | Bayu
8000450 | 2022-05-10 | RSI Siti Khadijah | Baby Incubator | Draeger/Isolette C2000 | EM31828 | Survey | Alarm control fail 14 | Pending | Disarankan penggantian heater assy | Amna
7000343 | 2022-05-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0026 | Survey | Pengecekan flow sensor dan sampe line | Ready to use | Disarankan penambahan stok flow sensor, sample line, waterlock dan service set | Jarwo
7000344 | 2022-05-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0002 | Survey | Pengecekan flow sensor dan sampe line | Ready to use | Disarankan penambahan stok flow sensor, sample line, waterlock dan service set | Jarwo
7000345 | 2022-05-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHL0059 | Survey | Pengecekan flow sensor dan sample line | Ready to use |  | Jarwo
7000346 | 2022-05-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Service | Pengecekan flow sensor dan sampe liine | Ready to use | Disarankan penambahan stok flow sensor, sample line, waterlock dan service set | Jarwo
6000563 | 2022-04-28 | RSUD Dr. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEL0081 | Service | Pemasangan touch screen, pressure regulator dan baterai internal | Ready to use |  | Afrizal
6000565 | 2022-04-28 | RSUD Dr. R. Goeteng Taroenadibrata | Incubator Transport | Draeger/TI500 | CR12875 | Service | Pemasangan baterai internal | Ready to use |  | Afrizal
19000449 | 2022-04-28 | RSUP Dr. Kariadi | Ventilator | Draeger/TI500 | ASML0087 | Service | Setting adjusment membran | Ready to use | Disarankan oenggantian filter air, filter O2 dan baterai | Nanda
16000189 | 2022-04-27 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Maintenance | Pengecekan filter kompressor | Ready to use |  | Ipan
19000448 | 2022-04-27 | RS Nasional Diponegoro | Incubator Transport | Draeger/Isolette C2000 | AF32536 | Maintenance | General check | Ready to use |  | Nanda
19000448 | 2022-04-27 | RS Nasional Diponegoro | Incubator Transport | Draeger/Isolette C2000 | AF32539 | Maintenance | General check | Ready to use |  | Nanda
19000448 | 2022-04-27 | RS Nasional Diponegoro | Incubator Transport | Draeger/Isolette C2000 | AF32538 | Maintenance | General check | Ready to use |  | Nanda
19000448 | 2022-04-27 | RS Nasional Diponegoro | Incubator Transport | Metaltronica/ Helianthus DBT | AF32540 | Service | General check | Ready to use |  | Nanda
6000564 | 2022-04-27 | RSKBD Panti Nugroho | Incubator Transport | Draeger/Babylog VN500 | NC12870 | Service | Pemasangan baterai internal | Ready to use |  | Afrizal
8000449 | 2022-04-27 | RSUD Kajen | Lampu Operasi | Draeger/Polaris | 005272 | Service | Lampu operasi mati | Pending | Disarankan penggantian power supply | Amna
13000122 | 2022-04-27 | RSUD Kota Langsa | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/057/C1 | Service | Penggantian RAW display | Ready to use |  | Sofian
6000562 | 2022-04-26 | RSUD Dr. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CU03924 | Service | Pemasangan spare part Hinge access | Ready to use |  | Afrizal
6000562 | 2022-04-26 | RSUD Dr. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CU03925 | Service | Pemasangan spare part Hinge access | Ready to use |  | Afrizal
6000562 | 2022-04-26 | RSUD Dr. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette C2000 | CU03921 | Service | Pemasangan spare part Hinge access | Ready to use |  | Afrizal
8000448 | 2022-04-26 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASNB0020 | Service | Penggantian O2 sensor | Ready to use |  | Amna
11000266 | 2022-04-26 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/VN500 | ASLA0080 | Service | Penggantian kit gas inlet cover | Ready to use |  | Nasrul
7000340 | 2022-04-25 | RSUD Bagas Waras | Ventilator | Draeger/Babylog VN500 | ASJD0018 | Service | Penggantian power supply | Ready to use |  | Jarwo
19000445 | 2022-04-25 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/E3000 | 1016250 | Service | Balon tidak mengembang maksimal, dilakukan pengurasan air dan penggantian air yang baru | Ready to use |  | Nanda
15000188 | 2022-04-25 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Ganti lampu kolimator | Ready to use |  | Sutan
19000453 | 2022-04-23 | RSUD Dr. Soeratno Gemolong | Patient Monitor | Draeger/Vista 120 | V2SHN0075 | Service | LCD bergaris blok | Pending | Disarankan penggantian LCD | Nanda
7000339 | 2022-04-23 | RSUD Sleman | Mesin Anasthesia | Draeger/Perseus A500 | ASMC0128 | Service | Penggantian water lock, pembersihan flow sensor | Ready to use |  | Jarwo
19000447 | 2022-04-23 | RSUP Dr. Kariadi | Ventilator | Draeger/Fabius Plus | ASML0079 | Service | Adjusting board and grid | Ready to use |  | Nanda
6000251 | 2022-04-22 | RSUD Suradadi Kab. Tegal | Mesin Anasthesia | Draeger/Babylog VN500 | ASKL0182 | Survey | O2 sensor fail | Ready to use | Dianjurkan penggantian O2 sensor | Afrizal
19000441 | 2022-04-22 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASML0079 | Service | kalibrasi dengan LPFK Solo | Ready to use |  | Nanda
19000444 | 2022-04-22 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASMK0322 | Service | Bug pada sistem | Ready to use |  | Nanda
8000447 | 2022-04-21 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHL0168 | Survey | General check | Ready to use |  | Amna
8000447 | 2022-04-21 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHL5958 | Service | General check | Ready to use |  | Amna
14000175 | 2022-04-21 | RSUD RA Basoeni | Dental Unit | Kavo/1058TM | 10205460 | Service | Pemasangan power supply back up karna power supply rusak kena Tikus | Ready to use |  | Satria
15000187 | 2022-04-20 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Survey | Software tidak bisa terbuka | Ready to use | Sudah dibuatkan penawaran reload software. ( Sudah diberikan penawaran. RS sudah PO. Sedang dalam perbaikan) | Sutan
7000337 | 2022-04-20 | RSUD Sleman | Jaudience Meter | Draeger/JM 105 | B3503649 | Survey | Error 5 | Pending | Disarankan penggantian baterai. Sudah penawaran, RS sudah PO, menunggu spare part datang | Jarwo
11000253 | 2022-04-20 | RSUD Tenriawaru Bone | Mesin Anasthesia | Draeger/Primus | ASJM0022 | Survey | O2 sensor fail dan Gas sensor fail | Pending | Disarankan penggantian sensor ILCA 3 dan O2 sensor | Nasrul
11000251 | 2022-04-20 | RSUD Tenriawaru Bone | Mesin Anasthesia | Draeger/Primus | ASHK0159 | Service | Ventilator error | Pending | Disarankan penggantian sensor ILCA 2 dan service set 2 tahun | Nasrul
19000441 | 2022-04-19 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0079 | Service | Penggantian PATO filter air, partical filter, batterai | Ready to use |  | Nanda
15000186 | 2022-04-19 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Komputer workstation tidak menyala. Tombol emergency table platinum tertekan | Ready to use |  | Sutan
14000174 | 2022-04-18 | RSUD Dr. R. Sosodoro Djatikoesoemo Bojonegoro | Dental Unit | Kavo/E30TM | 101111654 | Service | Penggantian board pcb | Ready to use |  | Satria
8000445 | 2022-04-14 | RS ISLAM PATI | UPS | Yoshiga/HPT2 | 501101 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2023-04-14
19000440 | 2022-04-14 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/EM1000 | 0319201 | Service | Penggantian air karna air keruh | Ready to use |  | Nanda
3000210 | 2022-04-14 | RSUD Kab. Temanggung | ESWL | Medispec/E3000 | 1115056 | Service | penggantian ssd | Ready to use |  | Imam
19000439 | 2022-04-13 | RSUD dr. Gunawan Mangunkusumo | Incubator Transport | Draeger/Fabius Plus | KY11399 | Survey | air flow low | Ready to use | Alat masih bisa digunakan, namun disarankan penggantian karet chamber hood | Nanda
6000555 | 2022-04-13 | RSUI Harapan Anda | Ventilator | Draeger/Babylog VN500 | ASJC0050 | Survey | pengecekan dasar, perlu penggantian internal battery, filter O2 | Ready to use |  | Afrizal
6000556 | 2022-04-13 | RSUI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0009 | Survey | volume tidal tidak sesuai | Pending | Dianjurkan penggantian flow sensor dan O2 sensor. Sudah diberikan penawaran | Afrizal
6000557 | 2022-04-13 | RSUI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0011 | Survey | General check | Ready to use |  | Afrizal
6000558 | 2022-04-13 | RSUI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0164 | Survey | Kalibrasi O2 error | Pending | Dianjurkan penggantian flow sensor dan O2 sensor. Sudah diberikan penawaran | Afrizal
6000559 | 2022-04-13 | RSUI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ARZF0073 | Survey | battery fail, O2 sensor fail | Pending | Disarankan ganti O2 sensor, flow sensor, maintenance kit 2 years. Sudah diberikan penawaran | Afrizal
6000560 | 2022-04-13 | RSUI Harapan Anda | Mesin Anasthesia | Draeger/TI500 | ASEL0010 | Service | VT tidak sesuai | Ready to use |  | Afrizal
7000330 | 2022-04-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0059 | Survey | Penggantian flow sensor | Ready to use |  | Jarwo
7000331 | 2022-04-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMH0007 | Survey | Cek service set, flow sensor, sample line | Ready to use |  | Jarwo
7000332 | 2022-04-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMH0118 | Survey | Cek service set, flow sensor, sample line | Ready to use |  | Jarwo
7000333 | 2022-04-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Survey | Cek service set, flow sensor, sample line | Ready to use | Disarankan penggantian service set, water lock, flow sensoro dan sample line | Jarwo
7000334 | 2022-04-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHL0059 | Survey | Pengecekan service set, flow sensor, sample line, dan water lock | Ready to use | Disarankan penggantian service set, water lock, flow sensoro dan sample line | Jarwo
7000336 | 2022-04-12 | RSUP Dr. Sardjito | Ventilator | Draeger/Perseus | ASDE0045 | Survey | Alat mati total | Ready to use |  | Jarwo
7000335 | 2022-04-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Babylog VN500 | ASKF0077 | Service | Pengecekan aksesoris | Ready to use | Disarankan penambahan stok flow sensor, sample line, waterlock dan service set | Jarwo
5000169 | 2022-04-11 | RS Adhyaksa | Sybron Endo | System B/1005 | 15-63757 | Service | Alat tidak bisa charging | Pending | Syrlu penggdhyaksaan module control electronic mainboard, valve block brake arm, g15 connector | Bilal
8000444 | 2022-04-11 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLN0011 | Survey | Penggantian baterai | Ready to use |  | Amna
19000438 | 2022-04-08 | RS Telogorejo | Scio | Draeger/Sciofour | ASDK0068 | Survey | Gas measurement error | Pending | Disarankan penggantian ILCA | Nanda
18000476 | 2022-04-08 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0173 | Service | Tarik error, alat rusak pada pressure regulator | Pending |  | Irfan R
3000209 | 2022-04-07 | RSUD Kab. Temanggung | ESWL | Medispec/EM1000 | 1115056 | Service | OS C-arm error, reload software dan OS | Ready to use |  | Imam
8000443 | 2022-04-07 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 edition | ASEL0142 | Survey | peep high. General check | Ready to use |  | Amna
6000551 | 2022-04-07 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0042 | Survey | General check | Ready to use |  | Afrizal
6000551 | 2022-04-07 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASMH0337 | Survey | error tombol body pneumatic | Ready to use | Disarankan penggjan module control electronic mainboard, valve block brake arm, gas connector | Afrizal
6000552 | 2022-04-07 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASLD0042 | Survey | Pengecekan Vaporizer | Ready to use |  | Afrizal
6000552 | 2022-04-07 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASMH0336 | Survey | Switch body tidak bisa tapi pneumatic aman | Ready to use | Disarankan penggjan module control electronic mainboard, valve block brake arm, gas connector | Afrizal
6000553 | 2022-04-07 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0043 | Survey | General check | Ready to use |  | Afrizal
6000554 | 2022-04-07 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASMH0335 | Service | Pendant anasthesi aman, pendant surgery Disarankan ganti pneumatic | Ready to use | Disarankan penggjan module control electronic mainboard, valve block brake arm, gas connector | Afrizal
6000600 | 2022-04-07 | RSUD Hj. Anna Lasmanah | Baby Incubator | Draeger/Isolette C2000 | CF32805 | Service | Pemasangan chamber humidify baru | Ready to use |  | Afrizal
19000436 | 2022-04-06 | RS Telogorejo | Mesin Anasthesia | Draeger/Primus | ASDK0223 | Service | Gas Agent tidak keluar | Ready to use |  | Nanda
5000168 | 2022-04-05 | Kesumah Dental Clinic | Dental Unit | Kavo/1058 | 10111157 | Service | Penggantian clip dan clamp, penggantian copling handpiece | Ready to use |  | Bilal
7000328 | 2022-04-05 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0144 | Survey | Heater error | Ready to use | Disarankan penggantian VGC board. Sudah diberikan penawaran. Sudah diperbaiki | Jarwo
11000267 | 2022-04-04 | RSUD H. Andi Sulthan Daeng Radja | Ventilator | Draeger/Babylog VN500 | ASMK0030 | Survey | Alat tidak bisa dikalibrasi | Pending | Disarankan penggantian kit proportional valve dan valve drive. Sudah diajukan penawaran dan sudah PO | Nasrul
6000599 | 2022-04-01 | RSUD Dr. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | AH36596 | Survey | Engsel incubator patah | Ready to use | Disarankan penggantian hinge access left & right rear. Sudah diperbaiki | Afrizal
6000599 | 2022-04-01 | RSUD Dr. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | AH36598 | Service | Engsel incubator patah | Pending | Disarankan penggantian hinge access left & right rear | Afrizal
3000208 | 2022-03-31 | RSUD DR. AGOESDJAM | Digital Radiografi Fluoroskopi | Apelem/Platinum | E196F0342 | Survey | Detector lost conncetion,ws di console tidak ada | Ready to use | Indikasi board allegro rusak. Sudah dibuatkan penawaran | Imam
6000598 | 2022-03-31 | RS Purwogondo | Ventilator | Draeger/Savina 300 | ASPJ0376 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
3000165 | 2022-03-31 | RS Purwogondo | Ventilator | Draeger/Savina 300 | ASPJ0376 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
19000434 | 2022-03-31 | RSUD Dr. Moewardi | Ventilator | Draeger/Babylog VN500 | ASEJ0053 | Service | Running test karna alat lama tidak dipakai | Ready to use | Disarankan penggantian flow sensor neonate, dan pembelian breathing circuit yang lengkap | Nanda
19000435 | 2022-03-31 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Isolette C2000 | ASKL0022 | Service | Alat humidifier lepas, sudah di pasang kembali | Ready to use |  | Nanda
5000167 | 2022-03-30 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Service | Penggantian applicator | Ready to use |  | Bilal
15000184 | 2022-03-29 | RS Bhayangkara Jambi | Mobile X-ray | Apelem/Rafale EV30 | 005/19/01531 | Service | Training User | Ready to use |  | Sutan
7000329 | 2022-03-29 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Survey | Error gas delivery | Ready to use | Penggantian ILCA | Jarwo
6000597 | 2022-03-29 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0071 | Service | Lubang udara kemasukan air | Ready to use |  | Afrizal
18000292 | 2022-03-29 | Siloam Hospital Pal II | Bubble CPAP | Draeger/CPAP | QM20190223 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Faugi
7000327 | 2022-03-29 | Rumah Sakit St. Elisabeth | Ventilator | Draeger/Evita V300 | ASRA0001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
8000441 | 2022-03-28 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Survey | Layar blank, bersihkan connector cable backlight LCD | Ready to use |  | Amna
14000173 | 2022-03-25 | Klinik Pramitha Jemursari | Bone Mineral Densitometer | DMS/Stratos | J08015D134 | Maintenance | Pengecekan unit | Pending | Banyak komponen yang sudah lemah | Satria
19000431 | 2022-03-25 | RSJ Prof Dr Soerojo Magelang | ESMR | Draeger/Primus | ASLD0039 | Service | General check | Ready to use |  | Nanda
6000595 | 2022-03-25 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Survey | Penggantian applicator esmr | Ready to use |  | Sutan
18000477 | 2022-03-25 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASFF0224 | Survey | Battery dan power supply rusak | Pending | Sudah diberikan penawaran | Irfan R
18000477 | 2022-03-25 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Fabius tiro | ASHD0029 | Service | Battery dan power supply rusak | Pending | Sudah diberikan penawaran | Irfan R
15000183 | 2022-03-24 | RSUD Bumiayu | Infant Warmer | Draeger/Resuscitaire | DSPF0044 | Service | Pemasangan powersuppy board | Ready to use |  | Afrizal
8000440 | 2022-03-24 | RSUD Dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0109 | Survey | Penggantian O2 sensor capsule | Ready to use |  | Amna
7000326 | 2022-03-24 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Survey | error gas analyzer | Ready to use |  | Jarwo
7000347 | 2022-03-24 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Service | Pengecekan flow sensor, sampe line dan service set | Ready to use | Disarankan penggantian water lock, flow sensor, sample line dan service set | Jarwo
19000430 | 2022-03-24 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHC0076 | Service | main control board error | Ready to use | Disarankan penggantian control main board. Sementara di tarik error dahulu | Nanda
19000433 | 2022-03-24 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASHC0076 | Service | Pemasangan service set | Ready to use |  | Nanda
5000166 | 2022-03-23 | Kesumah Dental Clinic | Dental Unit | Kavo/E30 | 10111157 | Service | Pengecekan koneksi selang dan pengeleman pipa | Ready to use |  | Bilal
19000429 | 2022-03-23 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0484 | Service | Penggantian humidifier | Ready to use |  | Nanda
6000594 | 2022-03-23 | RS Pratama Kelingi | Dental Unit | Kavo/E30 | 10111151 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan | 2023-03-23
19000428 | 2022-03-23 | RSUD Dr. Soediran Mangun Sumarso | UPS | Yoshiga/HPT2 | 501101236060KA10003 | Service | General check | Ready to use |  | Nanda
8000439 | 2022-03-23 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0033 | Service | Penggantian flow power dan ILCA 3 | Ready to use |  | Amna
8000439 | 2022-03-23 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0037 | Service | Penggantian flow power, ILCA 3 dan 2/2 distributing valve | Ready to use |  | Amna
8000439 | 2022-03-23 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0034 | Maintenance | Penggantian flow power dan ILCA 3 | Ready to use |  | Amna
6000594 | 2022-03-23 | RS Pratama Kelingi | Dental Unit | Kavo/E30 | 10111151 | Service | motor suction terbakar | Ready to use |  | Sutan
8000438 | 2022-03-23 | RS Mitra Bangsa Pati | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
19000425 | 2022-03-22 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C2000 | DNSL0047 | Maintenance | General check | Ready to use |  | Nanda
19000425 | 2022-03-22 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C8000 | CY34224 | Maintenance | General check | Ready to use |  | Nanda
19000425 | 2022-03-22 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C8000 | CF06005 | Service | General check | Ready to use |  | Nanda
19000426 | 2022-03-22 | RSUD Kota Salatiga | Infant Warmer | Draeger/Resuscitaire | FM09495 | Survey | Heater error | Pending | Disarankan penggantian heater. Sudah diberikan penawaran namun RS belum ada anggaran | Nanda
6000593 | 2022-03-22 | RSUI Harapan Anda | Baby Incubator | Draeger/Isolette C2000 | DSPL0022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2023-03-22
6000592 | 2022-03-21 | RSUD Ajibarang | Resuscitaire | Draeger/Resuscitaire | DSNJ0049 | Service | Pemasangan part heater resuscitaire | Ready to use |  | Afrizal
19000424 | 2022-03-21 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Maintenance | Balon tidak mengembang maksimal, dilakukan pengurasan air dan penggantian air yang baru | Ready to use |  | Nanda
7000324 | 2022-03-21 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Service | General check total care 2 tahun | Ready to use |  | Jarwo
7000325 | 2022-03-21 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMH0007 | Service | Penggantian 1 set valve | Ready to use |  | Jarwo
7000322 | 2022-03-18 | RS Panti Rapih | Ventilator Transport | Draeger/Oxylog 2000 Plus | 2M8673709 | Service | Error baterai | Ready to use |  | Jarwo
6000591 | 2022-03-18 | RSU An Ni'mah | Ventilator | Draeger/Savina 300 | ASPJ0377 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2023-03-18
7000319 | 2022-03-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Service | General check total care 2 tahun | Ready to use |  | Jarwo
7000323 | 2022-03-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0069 | Service | Gas O2, HO2O vaporizer tidak terbaca | Ready to use |  | Jarwo
6000589 | 2022-03-17 | RSIA Ummu Hani | Ventilator | Draeger/Evita4 | DRRN0039 | Service | Pemasangan O2 sensor, expiration valve, angular | Ready to use |  | Afrizal
6000583 | 2022-03-17 | RSKBD Panti Nugroho | Infant Warmer | Draeger/Resuscitaire | BV14104 | Maintenance | Pemasangan heater assy | Ready to use |  | Afrizal
19000417 | 2022-03-17 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0085 | Service | kalibrasi unit dengan bpfk | Ready to use |  | Nanda
19000423 | 2022-03-16 | RS Nasional Diponegoro | Incubator Transport | Draeger/Isolette TI500 | MR11799 | Service | Penggantian baterai | Ready to use |  | Nanda
19000423 | 2022-03-16 | RS Nasional Diponegoro | Incubator Transport | Draeger/Isolette TI500 | AF10588 | Service | Penggantian baterai | Ready to use |  | Nanda
7000321 | 2022-03-16 | RSUP Dr. Sardjito | Ventilator | Draeger/Savina 300 | ASPF0045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
7000321 | 2022-03-16 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN800 | ASPF0046 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
19000422 | 2022-03-16 | RSUP Dr. Kariadi | Compressor | Durr | L417422034 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2023-03-16
17000143 | 2022-03-16 | RSUI Harapan Anda | Ventilator | Draeger/Babylog VN800 | ASPJ0001 | Service | Error baterai internal | Ready to use |  | Afrizal
6000590 | 2022-03-16 | RSUI Harapan Anda | Ventilator | Draeger/Savina 300 | ARZF0049 | Service | Pemasangan service set 6thn | Ready to use |  | Afrizal
7000320 | 2022-03-16 | RSUD Prambanan | Ventilator | Draeger/Babylog VN500 | ASPF0049 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
7000320 | 2022-03-16 | RSUD Prambanan | Ventilator | Draeger/Babylog VN800 | ASPF0029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
7000320 | 2022-03-16 | RSUD Prambanan | Ventilator | Draeger/Babylog VN800 | ASPF0050 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo
11000267 | 2022-03-16 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 09111371 | Service | Penggantian EMH | Ready to use |  | Nasrul
11000222 | 2022-03-15 | RS Pelamonia Makassar | Bed Patient | Paramount Bed/PA66250D | 900710400062 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000222 | 2022-03-15 | RS Pelamonia Makassar | Bed Patient | Paramount Bed/PA66250D | 900710400063 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
6000584 | 2022-03-15 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASPL0085 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2024-03-15
6000588 | 2022-03-15 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0488 | Service | Kamera kolimator mati | Ready to use |  | Bahar
8000436 | 2022-03-15 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0120 | Service | Penggantian O2 sensor | Ready to use |  | Amna
8000436 | 2022-03-15 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ARYJ0111 | Service | Penggantian O2 sensor | Ready to use |  | Amna
8000436 | 2022-03-15 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0138 | Service | Penggantian O2 sensor | Ready to use |  | Amna
8000436 | 2022-03-15 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0011 | Survey | Penggantian O2 sensor dan service set 3 tahun | Ready to use |  | Amna
14000172 | 2022-03-15 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia |  | ASJE0005 | Service | Valve ventilator system fail, error mv2 | Ready to use | Disarankan penggantian 2/2 distributing valve. Sudah dilakukan perbaikan | Afrizal
19000421 | 2022-03-15 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Service | Penggantian Ilca 3 dual dan primus set 2 tahun | Ready to use |  | Nanda
8000435 | 2022-03-14 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASKF0024 | Service | Penggantian peep valve | Ready to use |  | Amna
6000586 | 2022-03-11 | RS Dadi Keluarga | Ventilator | Draeger/Babylog VN800 | ASPF0030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2024-03-11
8000434 | 2022-03-11 | RS ISLAM PATI | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0144 | Service | Penggantian O2 sensor | Ready to use |  | Amna
19000420 | 2022-03-11 | RSUP Dr. Kariadi | Baby Incubator | Draeger/Caleo | ASJM0048 | Service | fan motor tidak berfungsi | Ready to use | Disarankan penggantian fan motor. Alat sudah diperbaiki | Nanda
7000318 | 2022-03-11 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMH0007 | Service | Error valve karena pecah | Ready to use | Valve harus diganti 1 set. Sudah dilakukan perbaikan | Jarwo
6000587 | 2022-03-11 | RSUD RA Basoeni | Dental Unit | Kavo/1058TM | 10205460 | Service | Bangku pasient mengeluarkan asap | Ready to use |  | Satria
14000171 | 2022-03-10 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASNH0254 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2024-03-10
6000585 | 2022-03-10 | RS JIH | Mesin Anasthesia | Draeger/Fabius Plus | ASMJ0040 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2024-03-10
19000419 | 2022-03-09 | RS Triharsi | Mesin Anasthesia | Draeger/Fabius Plus | ASML0038 | Service | Ganti O2 sensor | Ready to use |  | Nanda
17000141 | 2022-03-09 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0488 | Service | Kamera kolimator terbalik dengan layar console | Ready to use |  | Bahar
7000316 | 2022-03-09 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Babylog VN500 | ASHJ0032 | Survey | error alarm dan nebulizer | Ready to use | Alat masih bisa digunakan tanpa nebulizer namun perlu penggantian speaker monitor dan pressure regulator | Jarwo
7000317 | 2022-03-09 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKG0070 | Service | error gas supply unit dan nebulizer | Ready to use | Disarankan penggantian pressure regulator | Jarwo
17000142 | 2022-03-09 | RSUD Dr. R. Sosodoro Djatikoesoemo Bojonegoro | Dental Unit | Kavo/E30TM | 101111654 | Survey | Unit display blank | Ready to use |  | Satria
3000207 | 2022-03-08 | RSUD Dr. Adjidarmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Service | motor colom macet | Ready to use |  | Imam
15000183 | 2022-03-08 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Service | Adapter komunikasi ke monitor operator rusak | Ready to use | Alat masih dapat berfungsi | Bahar
19000418 | 2022-03-08 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0125 | Survey | O2 failure, gas failure | Ready to use | Alat masih garansi. Penggantian pia sensor & flow power | Nanda
7000315 | 2022-03-07 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASKM0402 | Service | error gas delivery | Pending | Disarankan penggantian ILCA, flow power, dan MOBI. Sudah diberikan penawaran, RS belum ada budget | Jarwo
3000206 | 2022-03-02 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Service | keyboard mouse komputer ws not repsonding | Ready to use |  | Imam
7000314 | 2022-03-02 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASKJ0043 | Service | Error VGC power | Ready to use |  | Jarwo
15000180 | 2022-03-02 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Service | Penggantian applicator esmr | Ready to use |  | Sutan
19000416 | 2022-03-01 | RS Nasional Diponegoro | Central Monitor | Draeger/Vista 120 CMS | VLSNF0259 | Service | Setting bahasa default | Ready to use |  | Nanda
19000415 | 2022-03-01 | RSUD dr. Gunawan Mangunkusumo | Ventilator | Draeger/Savina 300 | ASLA0071 | Survey | baterai internal error | Pending | Disarankan penggantian baterai internal, filter & hepa filter. Sudah diberikan penawaran. RS tidak ada budget | Nanda
7000313 | 2022-03-01 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMH0007 | Service | Error valve | Ready to use | Valve pecah. Sudah dilakukan perbaikan | Jarwo
5000163 | 2022-02-26 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Survey | Unit bunyi, plumbing ada kebocoran | Ready to use |  | Bilal
7000312 | 2022-02-25 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0144 | Service | Heater error, VGC board rusak | Ready to use | Disarankan penggantian service set. Sudah diberikan penawaran. Sudah diperbaiki | Jarwo
20000110 | 2022-02-25 | RSUD Syekh Yusuf | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0028 | Service | Pasang service set, flow sensor dan O2 sensor | Ready to use |  | Nasrul
17000140 | 2022-02-25 | RSUP Dr. Hasan Sadikin | ESMR | Metaltronica/Helinthus BYM | HELI/0023/CB0 | Service | Tidak bisa isi data pasien, gambar bergaris, muncul error | Ready to use |  | Bahar
15000181 | 2022-02-25 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Service | Penggantian shock wave | Ready to use |  | Sutan
7000311 | 2022-02-25 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASHL0032 | Service | Error basis board dan inverter | Ready to use |  | Jarwo
3000203 | 2022-02-24 | RSUD Malingping | Digital Radiografi Fluoroskopi | Apelem/Platinum | J256F0489 | Maintenance | Rotor grid macet; Pengecekan sensor motor grid | Ready to use |  | Imam
3000205 | 2022-02-24 | RSUD Malingping | X-ray Digital | Apelem/ Camargue DR | IMA703916 | Maintenance | Back up data base system | Ready to use |  | Imam
7000310 | 2022-02-24 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Maintenance | General check | Ready to use |  | Jarwo
7000309 | 2022-02-24 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASLD0026 | Maintenance | General check | Ready to use |  | Jarwo
16000188 | 2022-02-23 | Kesumah Dental Clinic | Dental Unit | Kavo/E30TM | 10111157 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan | 2023-02-23
19000414 | 2022-02-23 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Primus | ASPH0125 | Survey | O2 sensor failure | Ready to use | Pengajuan claim warranty untuk sensor PIA & Flow power | Nanda
17000139 | 2022-02-22 | RSUD RA Kartini | C-Arm | Apelem/Evo R+ | 55120027813 | Survey | Training User | Ready to use |  | Bahar
20000108 | 2022-02-22 | RSUD Sawerigading | Mesin Anasthesia | Draeger/Perseus A500 | ASNJ0071 | Service | General check | Ready to use |  | Nasrul |  | √
19000413 | 2022-02-22 | RSUD Tugurejo | Baby Incubator | Draeger/Caleo | ASHK0035 | Service | layar garis garis dan bergelombang | Pending | Sudah diberikan penawaran. RS tidak ada budget karna spare part nya kemahalan | Nanda
19000412 | 2022-02-22 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0085 | Service | Pemasangan spare part PATO, internal baterai, filter air | Ready to use |  | Nanda
19000411 | 2022-02-21 | RSUD Tidar Magelang | Mesin Anasthesia | Draeger/Perseus A500 | ASPH0125 | Maintenance | selftest fail, gas measurement fail | Ready to use | Alat ditarik ke principle | Nanda
7000308 | 2022-02-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASNE0234 | Service | General check | Ready to use |  | Jarwo
17000138 | 2022-02-17 | RSUD Kab. Bekasi | C-Arm | Apelem/Evo R+ | 55118026649 | Service | Perbaikan handle pegangan | Ready to use |  | Bahar
7000306 | 2022-02-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHL0059 | Maintenance | General check | Ready to use |  | Jarwo
7000307 | 2022-02-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Maintenance | General check | Ready to use |  | Jarwo
19000410 | 2022-02-16 | Siloam Hospital Semarang | Baby Incubator | Draeger/Isolette C2000 | CR35565 | Maintenance | General check | Ready to use |  | Nanda
19000410 | 2022-02-16 | Siloam Hospital Semarang | Incubator Transport | Draeger/Isolette TI500 | CR12070 | Service | General check | Ready to use |  | Nanda
19000409 | 2022-02-16 | Siloam Hospital Semarang | Ventilator | Draeger/Savina 300 | ASKN0288 | Maintenance | General check | Ready to use | Disarankan penggantian service set 3 tahun | Nanda
19000405 | 2022-02-15 | RS Indriati Solo | Mesin Anasthesia | Draeger/Primus | ASJF0036 | Maintenance | General check | Ready to use |  | Nanda
19000406 | 2022-02-15 | RS Indriati Solo | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0063 | Service | General check | Ready to use | Alat masih bisa digunakan namun disarankan harus ganti O2 sensor | Nanda
19000408 | 2022-02-15 | RS Triharsi | Mesin Anasthesia | Draeger/Fabius Plus | ASML0038 | Maintenance | O2 sensor fail | Ready to use | Disarankan ganti O2 sensor, housing O2. Sudah dilakukan perbaikan | Nanda
19000407 | 2022-02-15 | RS Triharsi | Mesin Anasthesia | Draeger/Primus | ASMH0121 | Maintenance | General check | Ready to use | Alat sementara bisa digunakan. Disarankan penggantian service set 3 tahun | Nanda
8000433 | 2022-02-15 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | ASML0039 | Maintenance | General check | Ready to use | Disarankan penggantian O2 sensor | Amna
8000433 | 2022-02-15 | RSUD Ki Ageng Getas Pondowo Grobogan | Ventilator | Draeger/Savina 300 | ASMK0137 | Service | General check | Ready to use | Disarankan penggantian O2 sensor | Amna
19000404 | 2022-02-14 | RS Telogorejo | PGM Module | Draeger/Sciofour | ASDK0068 | Maintenance | Gas measurement error | Pending | Sudah diberikan penawaran. RS tidak ada budget | Nanda
8000432 | 2022-02-14 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0115 | Maintenance | General check | Ready to use | Disarankan penggantian O2 sensor dan service set 3 tahun | Amna
8000432 | 2022-02-14 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0148 | Maintenance | General check | Ready to use | Disarankan penggantian O2 sensor dan service set 3 tahun | Amna
8000432 | 2022-02-14 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Primus | ASNJ0049 | Service | General check | Ready to use | Disarankan penggantian O2 sensor dan service set 3 tahun | Amna
19000403 | 2022-02-14 | RSUP Dr. Kariadi | Ventilator | Draeger/Fabius Plus | ASMK0322 | Service | Error 42.0002 | Ready to use |  | Nanda
6000582 | 2022-02-10 | RSUD Dr. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2024-02-10
19000402 | 2022-02-10 | RSUD Pandan Arang | Resuscitaire | Draeger/Resuscitaire | CF10509 | Service | saklar on/off gas keluar air | Ready to use | Alat masih bisa dipakai namun disarankan penggantian saklar on/off | Nanda
7000305 | 2022-02-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Fabius Plus | ASKF0002 | Service | penggantian diafragma valve | Ready to use |  | Jarwo
7000303 | 2022-02-09 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Evita V300 | ASPC0206 | Maintenance | Penggantian touch screen dan basis board | Ready to use |  | Jarwo
8000431 | 2022-02-09 | RSUD BENDAN | Baby Incubator | Draeger/Isolette C2000 | CU29514 | Maintenance | General check | Ready to use | Disarankan penggantian filter | Amna
8000431 | 2022-02-09 | RSUD BENDAN | Baby Incubator | Draeger/Isolette C2000 | CB27888 | Maintenance | General check | Ready to use | Disarankan penggantian filter | Amna
8000431 | 2022-02-09 | RSUD BENDAN | Baby Incubator | Draeger/Isolette C2000 | CU29513 | Maintenance | General check | Ready to use | Disarankan penggantian filter | Amna
8000431 | 2022-02-09 | RSUD BENDAN | Baby Incubator | Medilink/Medix DR | CU29525 | Survey | General check | Ready to use | Disarankan penggantian filter | Amna
6000581 | 2022-02-09 | RSUD Kardinah Tegal | X-Ray Analog | Apelem/Camargue Analog | IMA703962 | Service | Rotor fault | Ready to use | Sudah diperbaiki | Afrizal
19000400 | 2022-02-09 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0079 | Service | PATO rusak | Ready to use | Saran penggantian PATO. Sudah diperbaiki | Nanda
7000304 | 2022-02-09 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Maintenance | Diafragma membran pecah | Ready to use | Disarankan penggantian diafragma membran. Sudah dilakukan perbaikan | Jarwo
5000162 | 2022-02-08 | RS Sentra Medika Cibinong | EECP | Lumenair/Vasomedical | 5TC1064 | Survey | General check | Ready to use |  | Bilal
14000170 | 2022-02-08 | RSUD Dr. Saiful Anwar | Bone Mineral Densitometer | Draeger/Isolette C2000 | J207H0735 | Service | Pemindahan unit | Ready to use |  | Satria
7000302 | 2022-02-07 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASMH0118 | Service | Penggantian baging | Ready to use |  | Jarwo
7000400 | 2022-02-04 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Evita V300 | ASPC0206 | Survey | Penggantian touch screen | Ready to use |  | Jarwo
6000580 | 2022-02-04 | RSUD Wonosobo | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/044/C0 | Service | Mamo tidak bisa RLMO | Ready to use | HDD dibawa untuk di cek | Afrizal
19000399 | 2022-02-04 | RSUP Dr. Kariadi | Infant Warmer | Draeger/Babytherm 8010 | ARUC0014 | Service | heater kurang panas | Ready to use | Sudah diganti dengan heater baru | Nanda
19000398 | 2022-02-04 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0079 | Survey | error O2 cell fail | Ready to use | Indikasi kerusakaan pada sensor PATO. Sudah diperbaiki | Nanda
6000579 | 2022-02-03 | RSUD Ajibarang | Resuscitaire | Draeger/Resuscitaire | DSNJ0049 | Service | Cek error | Ready to use | Akan diproses claim warranty heater. Sudah diperbaiki | Afrizal
19000397 | 2022-02-03 | RSUD dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0149 | Service | o2 no supply | Ready to use | Saran penggantian filter O2& air serta baterai internal | Nanda
19000397 | 2022-02-03 | RSUD dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0149 | Service | o2 no supply | Ready to use | Saran penggantian filter O2& air serta baterai internal | Nanda
14000169 | 2022-02-03 | RSUD Provinsi NTB | Mamografi | Metaltronica/Helianthus DBT | HELIDBT/032/C0 | Service | Windows tidak bekerja | Ready to use |  | Satria
7000301 | 2022-02-03 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Service | Penggantian ILCA dan flow power | Ready to use |  | Jarwo
3000204 | 2022-02-02 | RSUD Dr. Adjidarmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Service | Error ET120 | Ready to use |  | Imam
8000428 | 2022-02-02 | RS Graha Husada Jepara | Ventilator | Draeger/Savina 300 | ASPK0828 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2024-02-02
7000399 | 2022-02-02 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKG0069 | Service | Pengaturan ulang V tank karna ada kebocoran | Ready to use |  | Jarwo
14000168 | 2022-01-29 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 0720015 | Survey | Kalibrasi probe USG, test shockwave | Ready to use |  | Satria
6000482 | 2022-01-29 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0071 | Survey | O2 sensor fail | Ready to use | Disarankan penggantian O2 sensor | Afrizal
6000482 | 2022-01-29 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0073 | Service | O2 sensor fail | Ready to use | Disarankan penggantian O2 sensor | Afrizal
5000161 | 2022-01-28 | Klinik Drg. Cisca Alam | Dental Unit | Kavo/1058TM | 1002323 | Maintenance | waterblock bocor, handpiece bocor | Ready to use |  | Bilal
6000480 | 2022-01-28 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0001 | Survey | Pendant tidak mau mengerem | Pending | Disarankan penggantian module valve. Sudah diberikan penawaran, namun RS belum PO karna tidak urgent | Afrizal
6000480 | 2022-01-28 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0002 | Survey | Pendant tidak mau mengerem | Pending | Disarankan penggantian module valve. Sudah diberikan penawaran, namun RS belum PO karna tidak urgent | Afrizal
6000480 | 2022-01-28 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0018 | Survey | Pendant tidak mau mengerem | Pending | Disarankan penggantian module valve. Sudah diberikan penawaran, namun RS belum PO karna tidak urgent | Afrizal
6000480 | 2022-01-28 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0020 | Survey | Pendant tidak mau mengerem | Pending | Disarankan penggantian module valve. Sudah diberikan penawaran, namun RS belum PO karna tidak urgent | Afrizal
6000480 | 2022-01-28 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0049 | Survey | Pendant tidak mau mengerem | Pending | Disarankan penggantian module valve. Sudah diberikan penawaran, namun RS belum PO karna tidak urgent | Afrizal
6000480 | 2022-01-28 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0050 | Survey | Pendant tidak mau mengerem | Pending | Disarankan penggantian module valve. Sudah diberikan penawaran, namun RS belum PO karna tidak urgent | Afrizal
6000480 | 2022-01-28 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0122 | Survey | Pendant tidak mau mengerem | Pending | Disarankan penggantian module valve. Sudah diberikan penawaran, namun RS belum PO karna tidak urgent | Afrizal
4000329 | 2022-01-28 | RSUD Dr. M. Yunus | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/056/C1 | Service | Training User | Ready to use |  | Bayu
16000186 | 2022-01-28 | RSCM Kencana | Dental Unit | Kavo/1058TM | 1007780 | Survey | Pengecekan secara umum | Ready to use |  | Ipan
18000456 | 2022-01-27 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0171 | Service | Battery dan power supply rusak | Pending | Sudah diberikan penawaran | Irfan R
19000395 | 2022-01-27 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHA0114 | Service | Penggantian selang flow meter | Ready to use |  | Nanda
8000426 | 2022-01-26 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ARYJ0111 | Maintenance | General check | Ready to use |  | Amna
8000425 | 2022-01-26 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0138 | Maintenance | General check | Ready to use |  | Amna
8000425 | 2022-01-26 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0011 | Maintenance | General check | Ready to use |  | Amna
8000426 | 2022-01-26 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0120 | Service | General check | Ready to use |  | Amna
19000394 | 2022-01-26 | RSUD Tidar Magelang | Baby Incubator | Draeger/Caleo | ASJM0061 | Maintenance | Error 31 | Ready to use |  | Nanda
5000160 | 2022-01-26 | RSCM Kencana | Dental Unit | Kavo/E30TM | 10111155 | Maintenance | Perbaikan kompressor dental | Ready to use |  | Bilal
8000424 | 2022-01-24 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0118 | Maintenance | General check | Ready to use |  | Amna
8000424 | 2022-01-24 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0136 | Survey | General check | Ready to use |  | Amna
6000479 | 2022-01-24 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0071 | Service | leak notification | Ready to use |  | Afrizal
19000393 | 2022-01-24 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHA0114 | Service | Selang flow meter bocor | Ready to use | Sementara dipasang selang unit yang lama | Nanda
16000185 | 2022-01-22 | Klinik Drg. Hedy Soleman | Dental Unit | Kavo/1058TM | 1002324 | Service | waterblock bocor, handpiece bocor | Ready to use |  | Ipan
3000201 | 2022-01-21 | RSUD Ulin | Bone Mineral Densitometer | Medilink/Medix 90 | J13015M269 | Service | Feedback generator error | Ready to use |  | Imam
3000202 | 2022-01-21 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | Portable detector tidak menampilkan hasil gambar | Ready to use |  | Imam
5000159 | 2022-01-21 | FKG Universitas Trisakti | Dental Unit | Kavo/E30TM | 10111155 | Service | Footswitch bocor angin | Ready to use |  | Bilal
19000392 | 2022-01-21 | RSUD Tidar Magelang | Baby Incubator | Draeger/Caleo | ASJM0061 | Service | Error 31 | Ready to use |  | Nanda
19000391 | 2022-01-20 | RS Nasional Diponegoro | Incubator Transport | Draeger/Isolette TI500 | AF10588 | Service | Power DC low | Ready to use | Disarankan ganti baterai. Sudah diberikan penawaran. RS sudah PO. Sudah diperbaiki | Nanda
19000390 | 2022-01-20 | RS Nasional Diponegoro | Incubator Transport | Draeger/Isolette TI500 | MR11799 | Service | Indikator baterai low | Ready to use | Disarankan ganti baterai. Sudah diberikan penawaran. RS sudah PO. Sudah diperbaiki | Nanda
18000284 | 2022-01-20 | RSU Dharma Ibu | Bed Patient | Paramount Bed/2 crank | 901203500044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000284 | 2022-01-20 | RSU Dharma Ibu | Bed Patient | Paramount Bed/2 crank | 901203500047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
14000167 | 2022-01-20 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Service | Training User | Ready to use |  | Satria
8000423 | 2022-01-20 | RSUD Dr. R. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0041 | Service | O2 sensor fail | Ready to use |  | Amna
19000388 | 2022-01-19 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0085 | Service | O2 kalibrasi gagal, indikasi rusak pada PATO | Ready to use | Disarankan mengganti PATO. Sudah dilakukan perbaikan | Nanda
19000389 | 2022-01-19 | RSUP Dr. Kariadi | Ventilator Transport | Draeger/Oxylog 2000 Plus | SRUC0012 | Survey | Leakage & Paw high | Pending | Alat tidak bisa dicek karna tidak ada circuit pembanding | Nanda
7000398 | 2022-01-19 | RSUP dr. Soeradji Tirtonegoro | Infant Warmer | Draeger/Babytherm 8010 | ASMM0045 | Service | Error 35 | Ready to use | Disarankan penggantian heater. Sudah diberikan penawaran. Sudah diperbaiki | Jarwo
7000397 | 2022-01-18 | RS PKU Muhammadiyah Yogyakarta | Ventilator | Draeger/Babylog VN500 | ASLF0055 | Service | Penggantian O2 cell PATO dan air filter | Ready to use |  | Jarwo
8000422 | 2022-01-18 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASKF0024 | Service | Valve ventilator system fail | Pending | Disarankan penggantian deep valve & Service set 3 thn. Sudah diberikan penawaran | Amna
19000387 | 2022-01-18 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0322 | Service | VT low | Ready to use |  | Nanda
19000386 | 2022-01-15 | RS JIH | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0125 | Maintenance | Penggantian propotional valve | Ready to use |  | Nanda
11000265 | 2022-01-15 | RSUD I Lagaligo | Ventilator | Draeger/VN500 | ASKJ0005 | Maintenance | General check | Ready to use |  | Nasrul |  | √
8000421 | 2022-01-14 | RS Aisyiyah Kudus | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0048 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2024-01-14
8000421 | 2022-01-14 | RS Aisyiyah Kudus | Ventilator | Draeger/Babylog VN800 | ASPF0017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2024-01-14
8000421 | 2022-01-14 | RS Aisyiyah Kudus | Seatle PAP | Draeger/Oxymixer | QM20215174 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2023-01-14
16000184 | 2022-01-13 | Klinik Budihardja | Dental Unit | Kavo/1058 TM | 10010476 | Service | Air handpiece tidak keluar dengan lancar | Ready to use |  | Ipan
18000361 | 2022-01-13 | RS Siloam Paal Dua | Baby Incubator | Draeger/Isolette | DSNK0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000361 | 2022-01-13 | RS Siloam Paal Dua | Bubble CPAP | Draeger/CPap | QM20190223 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
7000395 | 2022-01-13 | RS JIH | Baby Incubator | Draeger/Isolette 8000 | DSPM0025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-01-13
4000328 | 2022-01-13 | RSUD Dr. Saiful Anwar | Bone Mineral Densitometer | Medilink/Medix DR | J207N0735 | Service | Melakukan uji paparan | Ready to use |  | Bayu
6000478 | 2022-01-12 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0043 | Maintenance | Pemasangan part 2/2 valve | Ready to use |  | Afrizal
4000327 | 2022-01-11 | RSUD Gambiran Kediri | C-Arm | Apelem/Evo R+ | 55121001850 | Survey | Melakukan uji paparan | Ready to use |  | Bayu
7000394 | 2022-01-10 | RS Panti Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASMB0006 | Service | O2 cell tidak berfungsi | Ready to use |  | Jarwo
19000385 | 2022-01-08 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Service | Kipas tidak bisa menyala, ada benda asing nyangkut | Ready to use |  | Nanda
7000393 | 2022-01-07 | Klinik Drg. Lenny | Dental Unit | Kavo/E30TM | 10111158 | Service | Penggantian piezo dan scaler | Ready to use |  | Jarwo
8000420 | 2022-01-07 | RSI Sultan Agung | Baby Incubator | Draeger/Isolette C2000 | CR35553 | Service | Penggantian hinge, acs pnl, dan right rear | Ready to use |  | Amna
8000419 | 2022-01-06 | RS PKU Aisyah Jepar | Mesin Anasthesia | Draeger/Fabius Plus | ASND0058 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2024-01-06
19000384 | 2022-01-06 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Service | General check | Ready to use |  | Nanda
6000476 | 2022-01-06 | RSUI Harapan Anda | Ventilator | Draeger/Evita 4 edition | ASCC0210 | Service | Pemasangan Service set 6thn | Ready to use |  | Afrizal
19000383 | 2022-01-05 | RS JIH | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0125 | Service | Internal FG flow gagal, ada valve yang rusak | Ready to use | Unit sementara bisa digunakan | Nanda
7000392 | 2022-01-05 | RSUD Wonosari | C-Arm | Apelem/Evo R+ | 55121060911 | Service | Penggantian monitor (claim warranty) | Ready to use |  | Jarwo
8000418 | 2022-01-04 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Primus | ASJE0011 | Service | Ganti water trap, sample line, flow sensor (maintenance kontrak Service) | Ready to use |  | Amna
8000417 | 2022-01-04 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Perseus A500 | ASNK0043 | Service | penggantian vaporizer (ASPM0752) | Ready to use |  | Amna
19000382 | 2022-01-03 | Klinik Drg. Lenny | Dental Unit | Kavo/E30TM | 10111158 | Survey | Handpiece tidak bisa, rusak pada EMS | Ready to use |  | Nanda
7000389 | 2022-01-03 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Caleo | ASKL0038 | Survey | Error 31 | Ready to use | Disarankan penggantian board VIT 2 controller. Sudah diberikan penawaran. Sudah diperbaiki | Jarwo
7000389 | 2022-01-03 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Isolette 8000 | AC09111 | Survey | Engsel patah | Ready to use | Alat masih bisa digunakan namun harus hati-hati | Jarwo
7000390 | 2022-01-03 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Isolette 8000 | AC09112 | Service | Engsel patah | Ready to use | Alat sudah diperbaiki | Jarwo
4000326 | 2021-12-30 | RSK Dharmais | Digital Radiografi Fluoroskopi | Apelem/Platinum | L216F0664 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2023-12-30
19000381 | 2021-12-30 | RSUD Dr. Soehadi Prijonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0047 | Service | pengecekan secara general | Ready to use |  | Nanda
9000192 | 2021-12-29 | RSUD Dr. M. M. Dunda | C-arm | Apelem/Evo R+ | 55118013636 | Service | Main board error | Ready to use |  | Irfan R
19000380 | 2021-12-29 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0079 | Service | alarm air way | Ready to use |  | Nanda
17000137 | 2021-12-28 | RSUD Kardinah Tegal | ESWL | Medispec/EM1000 | 0109117 | Service | HVG dan EMH rusak | Ready to use | Sudah dilanjutkan perbaikan | Bahar
6000474 | 2021-12-27 | RSUD KRT. Setjonegoro Wonosobo | Ventilator | Draeger/Savina 300 | ASLC0400 | Service | Pemasangan O2 oxyfrace VE | Ready to use |  | Afrizal
14000166 | 2021-12-27 | RSUD Dr. R. Sosodoro Djatikoesoemo Bojonegoro | Dental Unit | Kavo/E30TM | 1011154 | Service | handpiece tidak bisa dibuka pembuka mata bor nya | Ready to use |  | Satria
19000379 | 2021-12-24 | RS JIH | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0125 | Service | internal FG failure | Ready to use | Alat akan dicek lebih lanjut | Nanda
7000384 | 2021-12-24 | RS Panti Rapih | Ventilator | Draeger/Savina 300 | ARSC0003 | Survey | Error baterai, tidal volume rendah, PEEP error | Pending | Disarankan penggantian baterai, flow sensor, valve expirasi. Sudah diberikan penawaran | Jarwo
8000416 | 2021-12-24 | RSI Sunan Kudus | Ventilator | Draeger/Babylog VN800 | ASPF0016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-12-24
8000416 | 2021-12-24 | RSI Sunan Kudus | Ventilator | Draeger/Babylog VN800 | ASPF0018 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-12-24
8000416 | 2021-12-24 | RSI Sunan Kudus | Ventilator | Draeger/Babylog VN800 | ASPF0019 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-12-24
7000385 | 2021-12-23 | RSGM Prof. Soedomo | Lampu Operasi | Draeger/Polaris 600 | ASML1056 | Service | lampu mati semua | Ready to use | Disarankan penggantian power supply. Info dari RS pihak RS sudah ganti spare part sendiri | Jarwo
6000472 | 2021-12-23 | RSUD Bumiayu | Incubator Transport | Draeger/Isolette TI500 | AY11232 | Survey | Pemasangan baterai baru | Ready to use |  | Afrizal
6000473 | 2021-12-23 | RSUD Bumiayu | Resuscitaire | Draeger/Resuscitaire | DSPF0044 | Service | alat mati total | Ready to use | Akan diproses claim warranty power supply. Sudah dilakukan perbaikan | Afrizal
16000182 | 2021-12-23 | RSCM Kencana | Dental Unit | Kavo/1058TM | 1007780 | Maintenance | pengecekan secara berkala | Ready to use |  | Ipan
16000183 | 2021-12-23 | RSCM Kencana | Dental Unit | Kavo/E30TM | 10000659 | Survey | pengecekan secara berkala | Ready to use |  | Ipan
6000471 | 2021-12-22 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASHM0015 | Survey | Pemasangan power supply pack | Ready to use |  | Afrizal
14000165 | 2021-12-22 | RSUD Gambiran Kediri | C-Arm | Apelem/Evo R+ | 55121001850 | Survey | Pendampingan ukes LPFK Solo | Ready to use |  | Satria
7000384 | 2021-12-22 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Survey | HVPS keluar percikan api | Pending | Sudah diberikan penawaran namun RS belum ada anggaran untuk PO | Jarwo
7000386 | 2021-12-22 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | FY06854 | Service | Error sensor disconnect | Ready to use | Disarankan penggantian kabel sensor modul. Sudah diberikan penawaran namun pihak RS perbaiki sendiri dan alat sudah bisa kembali | Jarwo
9000187 | 2021-12-21 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina 300 | ASPJ0085 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000187 | 2021-12-21 | RS Gunung Maria Tomohon | Ventilator | Draeger/Savina 300 | ASPJ0086 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
19000378 | 2021-12-21 | RS JIH | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0125 | Survey | internal FG flow failure | Ready to use | Alat akan dicek lebih lanjut | Nanda
7000383 | 2021-12-21 | RS PKU Muhammadiyah Yogyakarta | Ventilator | Draeger/Babylog VN500 | ASLF0055 | Service | perlu penggantian filter O2 cell dan sensor O2 cell | Ready to use | Sudah dilakukan perbaikan 7000397 | Jarwo
8000415 | 2021-12-21 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASKF0024 | Service | system valve fail | Ready to use |  | Amna
8000414 | 2021-12-20 | RS Mitra Bangsa Pati | Ventilator | Draeger/Savina 300 | ASPJ0184 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-12-20
9000188 | 2021-12-20 | RS Gunung Maria Tomohon | Wheel Chair | Paramount bed | 90100850051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
9000188 | 2021-12-20 | RS Gunung Maria Tomohon | Bed Patient | Paramount bed | 901100800001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
9000188 | 2021-12-20 | RS Gunung Maria Tomohon | Examination Table | Paramount Bed/PK-1115 | PK112A02211 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
13000117 | 2021-12-18 | RSU Haji Medan | ESWL | Medispec/E3000 | 0719138 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sofian | 2023-12-18
19000377 | 2021-12-18 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0079 | Service | Alat tidak ada masalah | Ready to use |  | Nanda
19000376 | 2021-12-18 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0152 | Service | O2 gas supply tidak akurat | Ready to use |  | Nanda
7000382 | 2021-12-17 | RSUD Bagas Waras | Ventilator | Draeger/Babylog VN500 | ASJD0017 | Service | Penggantian baterai | Ready to use |  | Jarwo
7000382 | 2021-12-17 | RSUD Bagas Waras | Ventilator | Draeger/Babylog VN500 | ASJD0018 | Service | Penggantian baterai | Ready to use |  | Jarwo
7000382 | 2021-12-17 | RSUD Bagas Waras | Ventilator | Draeger/Babylog VN500 | ASJD0021 | Service | Penggantian baterai | Ready to use |  | Jarwo
15000177 | 2021-12-17 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Service | Kerusakan #103 #6 | Ready to use | Disarankan mengganti terminal board, konsul MCU sudah dipulangkan | Sutan
4000325 | 2021-12-17 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Survey | alat tidak bisa expose | Ready to use |  | Bayu
20000113 | 2021-12-16 | RSUD Haji Makassar | Dental Unit | Kavo/1058TM | 10205316 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2022-12-16
14000164 | 2021-12-16 | RSUD Dr. Saiful Anwar | Bone Mineral Densitometer | Medilink/Medix DR | J207H0735 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2023-12-16
8000413 | 2021-12-16 | RSU Fastabiq Sehat | UPS | Yoshiga/HPT10 | 501101310210M9300002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-12-16
8000413 | 2021-12-16 | RSU Fastabiq Sehat | UPS | Yoshiga/HPT10 | 501101310210M9300003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-12-16
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | S2SNM0818 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-12-16
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | S2SNM0819 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-12-16
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2161 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-12-16
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2165 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-12-16
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2166 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-12-16
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2167 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-12-16
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2168 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-12-16
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2171 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-12-16
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120s | S2SNM0818 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120s | S2SNM0819 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2167 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2168 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2171 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
7000380 | 2021-12-16 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Evita V300 | ASPC0206 | Survey | touch screen tidak berfungsi | Pending | Disarankan penggantian touch screen dan basis board. Sudah claim warranty, menunggu spare part | Jarwo
7000381 | 2021-12-16 | RS PKU Muhammadiyah Yogyakarta | Ventilator | Draeger/Babylog VN500 | ASLF0055 | Service | Error O2 calibration | Ready to use | Kerusakan pada O2 sensor (PATO). Sudah dilakukan perbaikan 7000397 | Jarwo
19000375 | 2021-12-16 | RS Triharsi | Ventilator | Draeger/Evita V300 | ASJH0064 | Service | flow sensor neo failed | Ready to use | Disarankan penggantian kabel flow sensor neonatal dan flow sensor neonatal. Sudah diperbaiki | Nanda
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2161 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2165 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000265 | 2021-12-16 | RS PIP Makassar | Patient Monitor | Draeger/Vista 120 | VGSNF2166 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
6000470 | 2021-12-16 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0043 | Service | MV2 selenoid valve rusak | Ready to use | Disarankan penggantian part 2/2 port distributing valve. Sudah dilakukan perbaikan | Afrizal
9000349 | 2021-12-16 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASNH0244 | Survey | Claim warranty power supplu | Ready to use |  | Irfan R
9000189 | 2021-12-16 | RS Gunung Maria Tomohon | Delivery Bed | Paramount Bed | PD1000 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
7000379 | 2021-12-15 | RSA UGM | Mesin Anasthesia | Draeger/Primus | ASKE0210 | Survey | O2 tidak akurat, gas sensor tidak akurat | Pending | Disarankan penggantian ilca 3 dan flow power. Sudah diberikan penawaran kontrak Service. RS tidak ada budget | Jarwo
6000468 | 2021-12-15 | RSU Muhammadiyah Siti Aminah | Mesin Anasthesia | Draeger/Fabius Plus | ASND0060 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2022-12-15
6000468 | 2021-12-15 | RSU Muhammadiyah Siti Aminah | Mesin Anasthesia | Draeger/Fabius Plus | ASND0060 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000469 | 2021-12-15 | RSUD Bumiayu | Incubator Transport | Draeger/Isolette TI500 | JY11058 | Survey | Pemasangan baterai | Ready to use |  | Afrizal
6000454 | 2021-12-15 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0073 | Service | O2 sensor failure | Pending | Disarankan penggantian O2 sensor PATO. Sudah diberikan penawaran | Afrizal
15000179 | 2021-12-15 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Maintenance | X-ray tidak keluar | Ready to use |  | Sutan
3000200 | 2021-12-14 | RSUD Kardinah Tegal | X-Ray Analog | Apelem/Camargue Analog | IMA703962 | Service | Rotor fault; Kontaktor dan board dibawa untuk di cek lebih lanjut | Ready to use | Sedang menunggu spare part datang | Imam
16000181 | 2021-12-14 | Klinik Pratama | Dental Unit | Kavo/1058TM | 10206373 | Maintenance | Pengecekan unit | Ready to use |  | Ipan
13000117 | 2021-12-14 | RSU Haji Medan | ESWL | Medispec/E3000 | 0719138 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sofian | 2023-12-14
6000466 | 2021-12-14 | RSUD Dr. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | AH36595 | Survey | Pemasangan part right front hinge acces panel | Ready to use |  | Afrizal
6000467 | 2021-12-14 | RSUD Dr. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | AH36595 | Service | Pintu tidak bisa menutup | Ready to use | Sudah dilakukan perbaikan (6000466) | Afrizal
6000466 | 2021-12-14 | RSUD Dr. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | AH36598 | Service | Pemasangan part right front hinge acces panel | Ready to use |  | Afrizal
19000374 | 2021-12-14 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Survey | Batu tidak pecah | Ready to use |  | Nanda
9000186 | 2021-12-13 | RSUD Provinsi Sulawesi Utara | Bed Patient | Paramount Bed | PK1127A | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
6000464 | 2021-12-13 | RSUD Banyumas | Baby Incubator | Draeger/Caleo | ASHK0029 | Survey | Display blank | Pending | Disarankan penggantian LCD, LCD driver, dan inverter LCD. Sudah diberikan penawaran | Afrizal
6000465 | 2021-12-13 | RSUD Banyumas | Baby Incubator | Draeger/Caleo | ASHK0072 | Survey | Casing belakang display patah | Pending | Disarankan penggantian back display case dan tubing grommet. Sudah diberikan penawaran | Afrizal
8000412 | 2021-12-13 | RSUD Kab. Temanggung | Vaporizer | Draegre/Vapor 3000 | ASPF0281 | Service | Vaporizer tidak bisa disetting | Ready to use |  | Amna
9000348 | 2021-12-13 | RS Manembo-nembo Bitung | Air Matrass Decubitus | Paramount Bed | C900027701 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
8000411 | 2021-12-10 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0014 | Service | VGC bunyi | Ready to use |  | Amna
6000462 | 2021-12-10 | RSUI Harapan Anda | Ventilator | Draeger/Savina 300 | ARZF0049 | Survey | Perbaikan kabel flow sensor sementara dan penggantian expiratory valve | Ready to use |  | Afrizal
6000461 | 2021-12-09 | RSUD Kardinah Tegal | Ventilator Transport | Draeger/Oxylog VE300 | ASNF0197 | Maintenance | Alat nyala namun harus forced shut down | Ready to use | Disarankan ganti tombol power. Sudah dilakukan perbaikan | Afrizal
7000377 | 2021-12-09 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Service | Total care | Ready to use |  | Jarwo
9000347 | 2021-12-09 | RSUP Ratatotok Buyat | X-ray Digital | Apelem/Camargue DR | CPD32474K18 | Maintenance | Penggantian board PCB | Ready to use |  | Irfan R
7000376 | 2021-12-08 | RSUD Bagas Waras | Mesin Anasthesia | Draeger/Primus | ASHL0062 | Survey | Service set 2 tahun | Ready to use |  | Jarwo
6000460 | 2021-12-08 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASHM0015 | Service | Unit tidak membaca sumber listrik | Ready to use | Disarankan penggantian power supply dan ganti maintenance Service kit 3 thn. Sudah diperbaiki (6000471) | Afrizal
8000410 | 2021-12-07 | RSUD Kajen | Patient Monitor | Draeger/Vista 120 | V3SHJ0006 | Maintenance | Alat mati total | Pending | Disarankan penggantian main board 9B45. Sudah diberikan penawaran | Amna
19000373 | 2021-12-07 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0321 | Service | Pemasangan Service set 3 tahun | Ready to use |  | Nanda
7000374 | 2021-12-07 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Service | Error leakage | Ready to use |  | Jarwo
7000375 | 2021-12-07 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Service | Error Ilca dan flow power | Ready to use |  | Jarwo
6000458 | 2021-12-06 | RSUD Banyumas | C-Arm | Apelem/Evo R+ | 55116055585 | Survey | Gambar xray tidak muncul di layar | Ready to use | Disarankan penggantian board Sigma. Sudah diberikan penawaran. Sedang proses perbaikan | Afrizal
6000459 | 2021-12-06 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASFL0120 | Service | Calibration gas supply berwarna merah, low batterai | Ready to use |  | Afrizal
19000371 | 2021-12-06 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0085 | Service | O2 supply error | Ready to use | Disarankan penggantian sensor PATO. Penawaran sudah diberikan | Nanda
15000178 | 2021-12-06 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Survey | Monitor workstation mati; Pengecekan PU Box, Generator OK | Ready to use |  | Sutan
6000457 | 2021-12-04 | RSUD Hj. Anna Lasmanah | Baby Incubator | Draeger/Isolette C2000 | CF32805 | Service | controller error 14 | Ready to use |  | Afrizal
19000370 | 2021-12-04 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0152 | Service | O2 fail, gas measurement fail | Ready to use |  | Nanda
8000457 | 2021-12-03 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | Service | O2 sensor fail | Ready to use |  | Amna
8000457 | 2021-12-03 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Survey | O2 sensor fail | Ready to use |  | Amna
6000455 | 2021-12-03 | RSKBD Panti Nugroho | Baby Incubator | Draeger/Isolette C2000 | EH37103 | Service | Engsel tidak bisa turun | Ready to use |  | Afrizal
6000456 | 2021-12-02 | RSUD Hj. Anna Lasmanah | Baby Incubator | Draeger/Caleo | ASJH0021 | Service | humidity tidak tercapai, motor tifting tidak berfungsi | Ready to use |  | Afrizal
9000345 | 2021-12-02 | RSUD Otohana | Ventilator | Draeger/Evita V800 | ASPL0095 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000345 | 2021-12-02 | RSUD Otohana | Ventilator | Draeger/Evita V800 | ASPL0165 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
70000370 | 2021-12-02 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASHL0032 | Service | Error basis board dan inverter | Ready to use |  | Jarwo
3000199 | 2021-12-01 | RSUD Dr. Adjidarmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Service | Status kalibrasi detector expired | Ready to use |  | Imam
19000369 | 2021-12-01 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Maintenance | meja tidak bisa bergerak | Ready to use |  | Nanda
7000368 | 2021-11-30 | RSUD Bagas Waras | Mesin Anasthesia | Draeger/Primus | ASHL0062 | Service | Total care | Ready to use |  | Jarwo
13000114 | 2021-11-30 | RSUD Dr. Fauziah | ESWL | Medispec/EM1000 | 0716249 | Maintenance | EMH bisa mengisi namun tidak bisa kempes | Ready to use |  | Sofian
9000344 | 2021-11-30 | RSUD KOTAMOBAGU | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0045 | Maintenance | Kontak Service | Ready to use |  | Irfan R
9000190 | 2021-11-29 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120s | VGSNE2468 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000190 | 2021-11-29 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120s | VGSNF2127 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000190 | 2021-11-29 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120s | VGSNF0972 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
19000367 | 2021-11-29 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Caleo | ASJE0010 | Service | air bocor ke laci | Ready to use | Sambungan heater patah. Penawaran sudah diberikan | Nanda
8000456 | 2021-11-29 | RSUD Kab. Temanggung | Resuscitaire | Draeger/Resuscitaire | JY10905 | Service | Error 10 | Pending | Disarankan penggantian main board supercap. Pihak RS belum ada anggaran | Amna
14000163 | 2021-11-28 | RSUD Dr. R. Sosodoro Djatikoesoemo Bojonegoro | Dental Unit | Kavo/E30TM | 10111154 | Pra Instalasi | cap ujung bor (on/off) tidak normal | Ready to use |  | Satria
3000198 | 2021-11-26 | RSUP Dr. Hasan Sadikin | USG | ESAOTE | 04323 | Service | Beberapa menu tidak berfungsi | Ready to use |  | Imam
14000162 | 2021-11-26 | RSUD Dr. Saiful Anwar | Bone Mineral Densitometer | Medilink/Medix DR | J207H0735 | Service | menunggu kalibrasi detektor & filamen tube | Ready to use | Alat sudah dapat befungsi dengan baik | Satria
19000364 | 2021-11-26 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASJC0044 | Service | leak tidak ter deteck, nilai volume tidal tidak mencapai setting | Ready to use |  | Nanda
19000365 | 2021-11-26 | RSUP Dr. Kariadi | Baby Incubator | Draeger/Isolette 8000 | MR07739 | Service | controller failure 4 | Ready to use |  | Nanda
19000363 | 2021-11-25 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Service | O2 sensor fail | Ready to use |  | Nanda
17000135 | 2021-11-25 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Service | EMH habis | Ready to use |  | Bahar
8000455 | 2021-11-24 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Fabius Plus | ASND0059 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-11-24
15000176 | 2021-11-24 | RSUD Dr. M. Yunus | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/056/C1 | Service | Woskstation bermasalah | Ready to use |  | Sutan
19000362 | 2021-11-24 | RSUD Kota Salatiga | ESWL | Medispec/E3000 | 09131174 | Service | Inverter fault | Ready to use |  | Nanda
19000361 | 2021-11-24 | RSUD Kota Salatiga | Bone Mineral Densitometer | Medilink/Medix 90 | L11015M207 | Service | generator fault & scanner error | Pending | disarankan penggantian monoblock, vanbelt dan blok inverter. Penawaran sudah diberikan | Nanda
16000180 | 2021-11-23 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058TM | 10007769 | Service | Penggantian main board | Ready to use |  | Ipan
19000360 | 2021-11-23 | RS Nasional Diponegoro | Incubator Transport | Draeger/Isolette TI500 | AF10588 | Maintenance | Pendampingan kalibrasi | Ready to use |  | Nanda
19000359 | 2021-11-22 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0482 | Maintenance | General check | Ready to use |  | Nanda
19000359 | 2021-11-22 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0483 | Maintenance | General check | Ready to use |  | Nanda
19000359 | 2021-11-22 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0484 | Maintenance | General check | Ready to use |  | Nanda
19000358 | 2021-11-22 | RS JIH | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0125 | Service | General check | Ready to use |  | Nanda
6000453 | 2021-11-22 | RSUD Dr. Soeselo Slawi | Ventilator | Draeger/Savina 300 | ASHD0014 | Service | Penggantian power supply | Ready to use |  | Afrizal
6000463 | 2021-11-22 | RSUI Harapan Anda | Ventilator | Draeger/Evita 4 edition | ASCC0210 | Maintenance | flow measurement inop | Ready to use |  | Afrizal
7000364 | 2021-11-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJF0009 | Survey | Total care | Ready to use |  | Jarwo
7000365 | 2021-11-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Survey | Error system internal fault | Ready to use | Erros sensor mixer. Sudah diperbaiki 7000391 | Jarwo
7000366 | 2021-11-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Service | Leakage pada manual spontan | Ready to use | Alat sudah diperbaiki (7000374) | Jarwo
7000367 | 2021-11-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Service | Error Ilca dan flow power | Ready to use |  | Jarwo
16000179 | 2021-11-22 | RSUD Malingping | Dental Unit | Kavo/1058TM | 10205319 | Survey | Penggantian board dentist element | Ready to use |  | Ipan
6000451 | 2021-11-20 | RSKBD Panti Nugroho | Ventilator | Draeger/Evita V300 | ASLD00168 | Service | Pemasangan baterai PS500, baterai internal | Ready to use |  | Afrizal
6000452 | 2021-11-20 | RSKBD Panti Nugroho | Ventilator | Draeger/Savina 300 | ASLL0399 | Service | Pemasangan spare part PBA CCB II, kit baterai internal | Ready to use |  | Afrizal
3000197 | 2021-11-19 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Survey | Penggantian board Sigma IQ | Ready to use |  | Imam
7000362 | 2021-11-19 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASKJ0043 | Service | Error internal system | Ready to use | Disarankan penggantian board VGC power. Sudah diperbaiki | Jarwo
7000363 | 2021-11-19 | RSUD Prambanan | Jaundice Meter | Draeger/JM 105 | H3500432 | Service | Error time period calibration | Ready to use | Alat perlu dikalibrasi. Sudah dikalibrasi (7000371) | Jarwo
8000454 | 2021-11-19 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0079 | Maintenance | O2 dan air sytem check fail | Ready to use |  | Amna
3000196 | 2021-11-18 | RSUD Malingping | X-ray Digital | Apelem/Camargue DR | IMA703916 | Service | Pengecekan kapasitas data dan sistem komputer | Ready to use |  | Imam
8000453 | 2021-11-18 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0301 | Service | Replace Service 2 years | Ready to use |  | Amna
20000106 | 2021-11-17 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Paramount Bed/PB2100 | 90101210006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000106 | 2021-11-17 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Paramount Bed/PB2100 | 90101210007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000106 | 2021-11-17 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Paramount Bed/PB2100 | 90101210009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000106 | 2021-11-17 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Paramount Bed/PB2100 | 900709700037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000107 | 2021-11-17 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Paramount Bed/PK8031B | 901003900014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000107 | 2021-11-17 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Paramount Bed/PK8031B | 901003900016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000107 | 2021-11-17 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Paramount Bed/PK8031B | 901003900024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
7000361 | 2021-11-16 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | FH36368 | Survey | Humidity low | Pending | Disarankan penggantian humidify. Sudah diberikan penawaran. | Jarwo
17000136 | 2021-11-16 | RSUP Fatmawati | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0493 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2024-11-16
9000338 | 2021-11-15 | RS Mata Sulawesi Utara | Trolly Emergency | Paramount Bed | 00030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000338 | 2021-11-15 | RS Mata Sulawesi Utara | Trolly Emergency | Paramount Bed | 00031 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000338 | 2021-11-15 | RS Mata Sulawesi Utara | Trolly Emergency | Paramount Bed | 00035 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000338 | 2021-11-15 | RS Mata Sulawesi Utara | Trolly Emergency | Paramount Bed | 00036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
19000357 | 2021-11-12 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Maintenance | General check | Ready to use |  | Nanda
20000150 | 2021-11-12 | RSUD Undata Palu | ESWL | Paramount bed/AR200 | 100024210012P | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
15000174 | 2021-11-12 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 0916032 | Maintenance | Error #300#6 | Ready to use |  | Sutan
7000373 | 2021-11-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Maintenance | Service set 2 tahun | Ready to use |  | Jarwo
7000360 | 2021-11-12 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Service | Service set 2 tahun | Ready to use |  | Jarwo
3000195 | 2021-11-11 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Pergerakan meja error | Ready to use |  | Imam
7000372 | 2021-11-11 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Service | error pressure sensor H2O | Ready to use |  | Jarwo
7000358 | 2021-11-10 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina 300 | ASPJ0062 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-11-10
7000358 | 2021-11-10 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina 300 | ASPJ0063 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-11-10
7000358 | 2021-11-10 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina 300 | ASPJ0064 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-11-10
7000358 | 2021-11-10 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Savina 300 | ASPJ0065 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-11-10
5000157 | 2021-11-10 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Survey | Touchsreen tidak berfungsi | Ready to use |  | Bilal
7000356 | 2021-11-08 | RS Pratama | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0048 | Survey | O2 cell tidak terbaca | Pending | Disarankan penggantian O2 cell satu set. Sudah diberikan penawaran namun RS belum PO | Jarwo
7000355 | 2021-11-08 | RS Pratama | Kompressor | Draeger | ASKD000006 | Survey | Kompressor error | Ready to use |  | Jarwo
7000355 | 2021-11-08 | RS Pratama | Kompressor | Draeger | ASKD00008 | Service | Kompressor error | Ready to use |  | Jarwo
9000339 | 2021-11-08 | RSUD Datoe Binangkang | Bubble CPAP | Draeger/MK | QM20215105 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
7000357 | 2021-11-08 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASJL00039 | Service | O2 sensor dan nebulizer tidak berfungsi | Ready to use |  | Jarwo
18000282 | 2021-11-06 | RSUD Sanana | Dental Unit | Kavo/E30TM | 10111150 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-11-06
19000354 | 2021-11-05 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0300 | Maintenance | Pemasangan Service set 2 tahun | Ready to use |  | Nanda
19000354 | 2021-11-05 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0303 | Maintenance | Pemasangan Service set 2 tahun | Ready to use |  | Nanda
19000354 | 2021-11-05 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0322 | Service | Pemasangan Service set 2 tahun | Ready to use |  | Nanda
3000194 | 2021-11-04 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Service | Image workstation kurang bagus hasil exposenya | Ready to use | Board interface dibawa untuk dicek lebih lanjut | Imam
14000161 | 2021-11-04 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Service | Kuncian probe lepas | Ready to use |  | Satria
11000262 | 2021-11-04 | RSUD H. Andi Sulthan Daeng Radja | Ventilator | Draeger/Babylog VN500 | ASMK0030 | Service | kalibrasi gas supply failed | Ready to use |  | Nasrul
9000337 | 2021-11-04 | RSUD Tani dan Nelayan | Ventilator | Draeger/Evita V800 | ASPH0001 | Service | Air supply error | Ready to use |  | Irfan R
9000337 | 2021-11-04 | RSUD Tani dan Nelayan | Ventilator | Draeger/Evita V800 | ASPH0002 | Service | Air supply error | Ready to use |  | Irfan R
19000353 | 2021-11-04 | RSUP Dr. Kariadi | Vaporizer | Draeger/Vapor2000 | ASLL1299 | Service | Nilai tidak sesuai settingan | Ready to use |  | Nanda
19000353 | 2021-11-04 | RSUP Dr. Kariadi | Vaporizer | Draeger/Vapor2000 | ASNE0707 | Survey | Nilai tidak sesuai settingan | Ready to use |  | Nanda
20000105 | 2021-11-04 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Paramount bed/PA93485 | 90100520007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000105 | 2021-11-04 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Paramount bed/PA93485 | 90100520020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000105 | 2021-11-04 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Paramount bed/PA93485 | 901005200024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000105 | 2021-11-04 | RSUP Dr. Wahidin Sudirohusodo | Bed Patient | Paramount bed/PA93485 | 901005200165 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
19000350 | 2021-11-03 | RSUD Kota Salatiga | ESWL | Medispec/E3000 | 09131174 | Survey | tidak muncul tampilan windows | Ready to use |  | Nanda
19000349 | 2021-11-03 | RSUD Kota Salatiga | Bone Mineral Densitometer | Medilink/Medix 90 | L11015M207 | Service | Generator fault | Ready to use | Alat digudangkan | Nanda
19000348 | 2021-11-02 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0079 | Service | Melakukan kalibrasi unit | Ready to use |  | Nanda
5000156 | 2021-11-01 | RS Marinir Cilandak | X-Ray Analog | Apelem/Camargue Analog | CPD26753A17 | Service | Control panel tidak menyala, ada kabel yang putus digigit tikus; Perbaikan kabel | Ready to use |  | Bilal
20000149 | 2021-11-01 | RSUD Banggai Laut | Meja Operasi | Paramount Bed/KM301 | 03382 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
19000347 | 2021-11-01 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0079 | Service | Air supply & O2 supply merah | Ready to use | Alat sudah diperbaiki (19000380) | Nanda
15000173 | 2021-11-01 | RSUP Dr. Mohammad Hoesin | C-Arm | Medispec/CS-2 | 06111187 | Service | Penggantian applicator ESMR | Ready to use |  | Sutan
7000359 | 2021-10-30 | RSUD Wonosari | C-Arm | Apelem/Evo R+ | 55121060911 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2023-10-30
7000353 | 2021-10-29 | Klinik Drg. Lenny | Dental Unit | Kavo/E30TM | 10111158 | Service | Scaller mati | Ready to use |  | Jarwo
7000352 | 2021-10-29 | RSUD Bagas Waras | Incubator Transport | Draeger/Isolette TI500 | JY11049 | Service | Penggantian baterai dan iris port | Ready to use |  | Jarwo
20000148 | 2021-10-29 | RSUD Bombana | Ventilator | Draeger/Babylog VN800 | ASPH0008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2022-10-29
6000402 | 2021-10-28 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0025 | Survey | karet-karet seal sudah getas | Ready to use |  | Afrizal
6000402 | 2021-10-28 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0027 | Survey | spontan pasient idak akurat, karet sudah getas | Ready to use |  | Afrizal
6000402 | 2021-10-28 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0039 | Survey | Perlu penggantian pressure regulator | Ready to use |  | Afrizal
6000402 | 2021-10-28 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0047 | Service | Perlu penggantian pressure regulator | Ready to use |  | Afrizal
19000346 | 2021-10-28 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Fabius GS Premium | ARVC0012 | Survey | O2 fail, leaktest fail | Pending | Disarankan penggantian APL modul, Service set, sensor O2 dan housingnya. Penawaran sudah diberikan | Nanda
7000174 | 2021-10-28 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0025 | Survey | O2 cell failed | Pending | Disarankan penggantian O2 cell karna sudah lemah | Jarwo
16000178 | 2021-10-28 | RSUD Arifin Achmad | Dental Unit | Kavo/E30TM | 10111156 | Survey | Training User | Ready to use |  | Ipan
6000401 | 2021-10-27 | RSUD Kardinah Tegal | Incubator Transport | Draeger/Isolette TI500 | AC13228 | Service | Bau gosong | Ready to use | Dianjurkan penggantian filter udara | Afrizal
6000500 | 2021-10-26 | RSUD Prof. Dr. Margono Soekarjo | Jaundice Meter | Draeger/JM 105 | B3503802 | Service | Pemasangan baterai baru | Ready to use |  | Afrizal
17000134 | 2021-10-25 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Survey | Touchsreen tidak berfungsi | Ready to use |  | Bahar
16000177 | 2021-10-25 | RSUD Samratulangi Tondano | Dental Unit | Kavo/1058TM | 10205322 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan | 2022-10-25
18000278 | 2021-10-23 | RSUD Dr. H. Chasan Boesoirie | Bedside Cabinet | Paramount Bed/PF-1100 | 91090010094 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000279 | 2021-10-23 | RSUD Dr. H. Chasan Boesoirie | Bed Patient | Paramount Bed/PA-6918 | 900915500021 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Sukri
18000279 | 2021-10-23 | RSUD Dr. H. Chasan Boesoirie | Bed Patient | Paramount Bed/PA-6918 | 900915500034 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000279 | 2021-10-23 | RSUD Dr. H. Chasan Boesoirie | Bed Patient | Paramount Bed/PA-6918 | 900915500037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000277 | 2021-10-23 | RSUD Dr. H. Chasan Boesoirie | Over bed table | Paramount Bed/KF282 | 910800400011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000277 | 2021-10-23 | RSUD Dr. H. Chasan Boesoirie | Over bed table | Paramount Bed/KF282 | 910800400024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000277 | 2021-10-23 | RSUD Dr. H. Chasan Boesoirie | Over bed table | Paramount Bed/KF282 | 910800400041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000278 | 2021-10-23 | RSUD Dr. H. Chasan Boesoirie | Bedside Cabinet | Paramount Bed/PF-1100 | 910900100088 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000278 | 2021-10-23 | RSUD Dr. H. Chasan Boesoirie | Bedside Cabinet | Paramount Bed/PF-1100 | 910900100107 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
3000193 | 2021-10-22 | RSUD Tidar Magelang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0496 | Service | DICOM Service error; Konfigurasi ulang DICOM | Ready to use |  | Imam
20000147 | 2021-10-22 | RS UD Lamadukkeleng | Stretcher | Paramount bed/PK5114X | 900901500039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000147 | 2021-10-22 | RS UD Lamadukkeleng | Stretcher | Paramount bed/PK5114X | 900901500053 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000147 | 2021-10-22 | RS UD Lamadukkeleng | Stretcher | Paramount bed/PK5114X | 900901500055 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000147 | 2021-10-22 | RS UD Lamadukkeleng | Stretcher | Paramount bed/PK5114X | 900901500056 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000147 | 2021-10-22 | RS UD Lamadukkeleng | Stretcher | Paramount bed/PK5114X | 900901500057 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
7000172 | 2021-10-21 | RSUD Prambanan | Ventilator | Draeger/Babylog VN500 | ASPD0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-10-21
9000320 | 2021-10-21 | RSUD Bumi Panua | Ventilator | Draeger/Evita 300 Advance | ASPH0001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2023-10-21
20000109 | 2021-10-19 | RSUD Haji Makassar | Ventilator | Draeger/Evita V300 | ASPH0056 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-10-19
20000109 | 2021-10-19 | RSUD Haji Makassar | Ventilator | Draeger/Evita V300 | ASPH0057 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-10-19
20000109 | 2021-10-19 | RSUD Haji Makassar | Ventilator | Draeger/Evita V300 | ASPH0059 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-10-19
20000109 | 2021-10-19 | RSUD Haji Makassar | Ventilator | Draeger/Evita V300 | ASPH0069 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-10-19
14000160 | 2021-10-19 | RSUD Gambiran Kediri | C-Arm | Apelem/Evo R+ | 55121001850 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2023-10-19
7000171 | 2021-10-19 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | VGSNF1838 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-10-19
7000171 | 2021-10-19 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | VGSNF1839 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-10-19
7000171 | 2021-10-19 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | VGSNF1843 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-10-19
15000172 | 2021-10-19 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Service | tube ceilling bocor | Ready to use | Alat masih bisa digunakan | Sutan
6000498 | 2021-10-15 | RSUD Dr. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette 8000 | AH36595 | Survey | Pintu patah | Ready to use | Disarankan penggantian right front hinge access | Afrizal
6000498 | 2021-10-15 | RSUD Dr. Soedirman - Kebumen | Draeger/Isolette C2000 | Draeger/Isolette C2000 | AH36596 | Survey | Pintu patah | Ready to use | Disarankan penggantian right front hinge access | Afrizal
6000498 | 2021-10-15 | RSUD Dr. Soedirman - Kebumen | Incubator Transport | Draeger/Isolette TI500 | AH36597 | Service | Pengecekan alat | Ready to use | Sudah diperbaiki pengecekan selanjutnya (06000466) | Afrizal
6000498 | 2021-10-15 | RSUD Dr. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette 8000 | AH36598 | Service | Penggantian filter udara | Ready to use | Sudah diperbaiki pengecekan selanjutnya (06000466) | Afrizal
15000171 | 2021-10-15 | RSUP Dr. M. Djamil Padang | USG | ESAOTE | 03862 | Survey | pengecekan alat | Ready to use |  | Sutan
13000111 | 2021-10-15 | RSUD Arifin Achmad | Dental Unit | Kavo/E30TM | 10111156 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sofian | 2022-10-15
7000200 | 2021-10-14 | RS Nur Hidayah | Ventilator | Draeger/Savina 300 | ASPJ0033 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-10-14
6000497 | 2021-10-14 | RSUD DR. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEL0081 | Survey | touchscreen error | Ready to use | Disarankan penggantian pressure regulator, touchsreen saja. Sudah diberikan penawaran | Afrizal
6000497 | 2021-10-14 | RSUD DR. R. Goeteng Taroenadibrata | Medical Compressor | Draeger/Medical Compressor | ASJL0056 | Service | kompressor berisik | Ready to use | Disarankan penggantian motor compressor, filter O2 dan air, filter board, dan Service set compreesor. Sudah diberikan penawaran | Afrizal
15000170 | 2021-10-14 | RSUP Dr. M. Djamil Padang | EECP | Lumenair/Vasomedical | 06111188 | Service | pengecekan unit | Ready to use | Alat tipe lama sudah discontinue, budget tidak masuk jika harus upgrade alat | Sutan
15000169 | 2021-10-14 | RSUP Dr. M. Djamil Padang | EECP | Lumenair/Vasomedical | 5TC1093 | Survey | kabel spo2 harus diganti | Pending | Menunggu info RS untuk permintaan penawaran (mereka melihat budget terlebih dahulu) | Sutan
16000176 | 2021-10-14 | RSU Haji Medan | Dental Unit | Kavo/2058 life | 10207192 | Survey | Training User | Ready to use |  | Ipan
6000496 | 2021-10-13 | RSUD Hj. Anna Lasmanah | Baby Incubator | Draeger/Caleo | ASJD0039 | Survey | Display monitor tidak menyala | Ready to use |  | Afrizal
6000495 | 2021-10-13 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0025 | Survey | Leak test gagal | Ready to use |  | Afrizal
6000495 | 2021-10-13 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0082 | Service | EtCo2 kadang terbaca kadang tidak | Pending | Disarankan penggantian waterlock. Sudah diberikan penawaran | Afrizal
19000341 | 2021-10-13 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0305 | Service | Alarm baterai internal | Ready to use |  | Nanda
7000198 | 2021-10-13 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Survey | Error valve MV2, error VGC power board, errpr power supply, error internal battery | Ready to use |  | Jarwo
7000199 | 2021-10-13 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Service | Sensor EtO2 kadang bisa kadang tidak | Ready to use | Disarankan penggantian flow sensor | Jarwo
20000102 | 2021-10-13 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 09111371 | Service | Gear box moor actuator rusak | Ready to use |  | Irfan S
5000155 | 2021-10-12 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Survey | LED sering mati dan muncul cursor | Ready to use |  | Bilal
7000196 | 2021-10-12 | RSUD Bagas Waras | Incubator Transport | Draeger/Isolette TI500 | JY11049 | Service | incubator tidak bisa dibuat transport | Ready to use | Disarankan penggantian baterai internal. Sudah diperbaiki (7000352) | Jarwo
11000261 | 2021-10-12 | RSUD Morowali | Mesin Anasthesia | Draeger/Atlan A350 | ASPH0021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-10-12
6000493 | 2021-10-12 | RSUD Brebes | Resuscitaire | Draeger/Resuscitaire | AR11911 | Maintenance | Pemasangan kit corner block RW | Ready to use |  | Afrizal
19000338 | 2021-10-12 | RSUD dr. Gunawan Mangunkusumo | Incubator Transport | Draeger/Isolette TI500 | KY11400 | Service | General check | Ready to use |  | Nanda
19000339 | 2021-10-12 | RSUD dr. Gunawan Mangunkusumo | Central Monitor | Draeger/Vista 120 CMS | VLSDMF0152 | Service | urutan patient monitor acak | Ready to use |  | Nanda
8000409 | 2021-10-12 | RSUD Kajen | Resuscitaire | Draeger/Resuscitaire | HF10401 | Service | Error 05 & Kit corner block rw | Ready To Use |  | Amna
11000261 | 2021-10-12 | RSUD Morowali | Mesin Anasthesia | Draeger/Atlan A350 | ASPH0022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-10-12
11000263 | 2021-10-12 | RSUD Morowali | CPAP | Draeger/Oxymixer High Flow | QM20205059 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-10-12
7000197 | 2021-10-12 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | VGSNE6312 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-10-12
6000494 | 2021-10-12 | RSUI Harapan Anda | Ventilator | Draeger/Evita 4 edition | ASCC0210 | Service | flow measurement inop | Ready to use | Disarankan pengecekan ulang dengan board flow sensor. Sudah diperbaiki | Afrizal
15000167 | 2021-10-12 | RSUP Dr. M. Djamil Padang | ESMR | Medispec/ED1000 | 0621220 | Service | penggantian applicator | Ready to use |  | Sutan
15000168 | 2021-10-12 | RSUP Dr. M. Djamil Padang | Bone Mineral Densitometer | Medilink/Medix 90 | J13015M270 | Survey | pengecekan alat | Ready to use |  | Sutan
6000491 | 2021-10-11 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHH0153 | Service | Batterai drop | Ready to use | Disarankan penggantian baterai internal. Sudah diperbaiki | Afrizal
6000492 | 2021-10-11 | RSUD Brebes | Baby Incubator | Draeger/Caleo | ASHM0066 | Service | Pemasangan conv humidity | Ready to use |  | Afrizal
6000491 | 2021-10-11 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASNE0151 | Service | Pemasangan back up spare part PBA VGC analog 2 | Ready to use |  | Afrizal
20000146 | 2021-10-09 | RSIA Defina | Medical compressor | Durr/Sicolab Med | DSPC0072 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2022-10-09
14000159 | 2021-10-08 | K8 Dental Care | Dental Unit | Kavo/1058TM | 10010470 | Maintenance | keypad dentist element rusak | Ready to use |  | Satria
7000195 | 2021-10-08 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0039 | Service | Pemasangan Service set 2 tahun | Ready to use |  | Jarwo
20000145 | 2021-10-07 | RSUD Mokoyurli Buol | Ventilator | Draeger/Babylog VN500 | ASPD0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2022-10-07
15000166 | 2021-10-07 | RSUP Dr. Mohammad Hoesin | ESMR | Apelem/Rafale EV30 | 009/15/01095 | Service | X-ray tidak keluar | Ready to use |  | Sutan
15000165 | 2021-10-06 | RSUP Dr. Mohammad Hoesin | Mamografi | Medispec/CS-2 | 06111187 | Service | shock habis | Ready to use |  | Sutan
3000191 | 2021-10-05 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helinthus BYM | HELI/0023/CB0 | Service | Komputer work station tidak menyala | Ready to use |  | Imam
3000190 | 2021-10-01 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | console generator rusak | Ready to use |  | Imam
17000133 | 2021-10-01 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Service | Pergerakan unit terkadang berhenti mendadak | Ready to use |  | Bahar
18000274 | 2021-10-01 | RSUD dr. Zainal Umar Sidiki | Ventilator | Draeger/Savina 300 | APSH0297 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-10-01
18000273 | 2021-10-01 | RSUD dr. Zainal Umar Sidiki | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0080 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-10-01
8000408 | 2021-10-01 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0014 | Service | System check sensor check fail | Ready to use |  | Amna
7000192 | 2021-10-01 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1804 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-10-01
16000175 | 2021-09-30 | RSCM Kencana | oxymixer High Flow | Kavo/E30TM | 1007780 | Maintenance | Pengecekan dental unit berkala | Ready to use |  | Ipan
17000132 | 2021-09-30 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0488 | Service | Pemasangan cover column tube stand yang telah diperbaiki | Ready to use |  | Bahar
18000272 | 2021-09-30 | RSUD Tani dan Nelayan | Ventilator | Draeger/Evita V800 | ASPH0001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-09-30
18000272 | 2021-09-30 | RSUD Tani dan Nelayan | Ventilator | Draeger/Evita V800 | ASPH0002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-09-30
8000407 | 2021-09-30 | RSUD Tidar Magelang | Resuscitaire | Draeger/Resuscitaire | CY11160 | Service | kit cournet block rw patah | Ready to use |  | Amna
7000194 | 2021-09-30 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASHL0031 | Service | Pemasangan flow sensor dan kabelnya | Ready to use |  | Jarwo
7000194 | 2021-09-30 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASHL0032 | Service | Pemasangan flow sensor dan kabelnya | Ready to use |  | Jarwo
7000194 | 2021-09-30 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASJK0038 | Service | Pemasangan flow sensor dan kabelnya | Ready to use |  | Jarwo
7000194 | 2021-09-30 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASKE0007 | Service | Pemasangan flow sensor dan kabelnya | Ready to use |  | Jarwo
7000194 | 2021-09-30 | RSUP Dr. Sardjito | Incubator Transport | Draeger/Globe Trotter | CY11356 | Service | Pemasangan O2 sensor | Ready to use |  | Jarwo
7000194 | 2021-09-30 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | VM500 | Service | Pemasangan neonatal flow sensor | Ready to use |  | Jarwo
16000172 | 2021-09-30 | RSCM Kencana | Dental Unit | Kavo/E30TM | 10000659 | Maintenance | Pengecekan dental unit berkala | Ready to use |  | Ipan
19000336 | 2021-09-29 | RS Nasional Diponegoro | Resuscitaire | Draeger/Resuscitaire | MF10231 | Maintenance | Error 05 | Ready to use |  | Nanda
9000316 | 2021-09-29 | RSUD Noongan | Mesin Anasthesia | Draeger/Perseus A500 | ASJF0093 | Maintenance | kontrak Service total care | Ready to use |  | Irfan R
9000316 | 2021-09-29 | RSUD Noongan | Mesin Anasthesia | Draeger/Perseus A500 | ASJL0006 | Maintenance | kontrak Service total care | Ready to use |  | Irfan R
9000316 | 2021-09-29 | RSUD Noongan | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0066 | Service | kontrak Service total care | Ready to use |  | Irfan R
7000191 | 2021-09-29 | RSUP Dr. Sardjito | Central Monitor | Draeger/Infinity Central Station | 10Q7XL00263 | Service | Beberapa monitor tidak terhubung dan sudah di re-setting | Ready to use |  | Jarwo
19000335 | 2021-09-29 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHC0076 | Service | kabel flow sensor tidak bisa, o2 fail, leak test fail | Ready to use | Disarankan penggantian kabel flow sensor, sudah diberikan penawaran | Nanda
19000334 | 2021-09-29 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Service | O2 sensor fail | Ready to use |  | Nanda
16000172 | 2021-09-28 | Klinik Family Dentistry | Dental Unit | Kavo/E30TM | 1000655 | Service | Selenoid valve, nozzle suction bermasalah | Ready to use |  | Ipan
14000158 | 2021-09-28 | RSUD Provinsi NTB | Mobile X-ray | Apelem/Rafale EV30 | 009/18/01274 | Service | unit tidak menyala | Ready to use |  | Satria
18000249 | 2021-09-28 | RSUD Toto Kabila | Phototherapy | Draeger/Bililux | 1111527012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2022-09-28
18000249 | 2021-09-28 | RSUD Toto Kabila | Phototherapy | Draeger/Bililux | 1111527020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2022-09-28
18000247 | 2021-09-28 | RSUD Toto Kabila | Infant Warmer | Draeger/Babytherm 8010 | ASNL0041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2022-09-28
18000247 | 2021-09-28 | RSUD Toto Kabila | Resuscitaire | Draeger/Resuscitaire | DSPF0014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2022-09-28
18000247 | 2021-09-28 | RSUD Toto Kabila | Resuscitaire | Draeger/Resuscitaire | DSPF0020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2022-09-28
18000248 | 2021-09-28 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | VGSNF1683 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2022-09-28
18000248 | 2021-09-28 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | VGSNF1684 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2022-09-28
18000248 | 2021-09-28 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | VGSNF1685 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2022-09-28
18000248 | 2021-09-28 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | VGSNF1686 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2022-09-28
18000247 | 2021-09-28 | RSUD Toto Kabila | Infant Warmer | Draeger/Babytherm 8010 | ASNL0047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2022-09-28
9000343 | 2021-09-28 | RSUD Toto Kabila | Ventilator | Draeger/Babylog VN500 | ASPA0014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2023-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1509 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1510 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1512 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1587 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1594 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1597 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1598 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1599 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1600 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1716 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1724 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1730 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1732 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1733 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1738 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1803 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1804 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
7000188 | 2021-09-28 | RSUP dr. Soeradji Tirtonegoro | Patient Monitor | Draeger/Vista 120 | VGSNF1805 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-28
19000333 | 2021-09-27 | RS Telogorejo | Mesin Anasthesia | Draeger/Primus | ASDK0223 | Maintenance | pengecekan berkala | Ready to use |  | Nanda
19000333 | 2021-09-27 | RS Telogorejo | Mesin Anasthesia | Draeger/Primus | ASDK0300 | Maintenance | pengecekan berkala | Ready to use |  | Nanda
19000333 | 2021-09-27 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDK0302 | Maintenance | pengecekan berkala | Ready to use |  | Nanda
19000333 | 2021-09-27 | RS Telogorejo | Mesin Anasthesia | Draeger/Primus | ASDL0002 | Maintenance | pengecekan berkala | Ready to use |  | Nanda
19000333 | 2021-09-27 | RS Telogorejo | Mesin Anasthesia | Draeger/Primus | ASDL0003 | Survey | pengecekan berkala | Ready to use |  | Nanda
6000489 | 2021-09-27 | RSUD Dr. Tjitrowardojo Purworejo | Ventilator | Draeger/Savina 300 | ASBM0243 | Survey | volume tidal tidak sesuai | Pending | Disarankan penggantian spare O2 exytrace 2pcs dan Service set 6thn. Sudah diberikan penawaran | Afrizal
6000490 | 2021-09-27 | RSUD Dr. Tjitrowardojo Purworejo | Pendant | Draeger/Movita | ASHE0161 | Service | Pneumatic tidak berfungsi | Pending | Disarankan penggantian pneumatic valve. Sudah diberikan penawaran | Afrizal
8000406 | 2021-09-27 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0034 | Service | Gas analyzer fail & o2 sensor fail | Ready to use |  | Amna
11000264 | 2021-09-26 | RS PIP Makassar | Ventilator | Draeger/Savina 300 | ASPJ0084 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-09-26
3000188 | 2021-09-24 | RSUD KRT. Setjonegoro Wonosobo | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0302 | Service | Komputer workstation error; Reload software workstation | Ready to use |  | Imam
3000189 | 2021-09-24 | RSUD KRT. Setjonegoro Wonosobo | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/044/C0 | Service | lengan alat tidak bisa bergerak | Ready to use |  | Imam
7000186 | 2021-09-24 | RSPAU Harjolukito | Central Monitor | Draeger/Infinity Central Station | DTCJ02B0006 | Service | central monitor mati total | Ready to use | disarankan upgrade unit karna spare part sudah discontinue | Jarwo
19000332 | 2021-09-24 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Caleo | ASKL0013 | Service | pengisian humidify bocor | Ready to use | Alat bisa digunakan tanpa humidify, disarankan penggantian conv water heater | Nanda
20000143 | 2021-09-24 | RSUD Nene Mallomo | Lampu Operasi | Draeger/Polaris 600 | ASNE1104 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-09-24
7000187 | 2021-09-24 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0041 | Service | Selftes (nilai tidal volum rendah) | Ready to use |  | Jarwo
3000187 | 2021-09-23 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | Koneksi monitor operator serign terputus | Ready to use |  | Imam
7000184 | 2021-09-22 | RSUD Wonosari | Ventilator | Draeger/Babylog VN800 | ASPF0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Jarwo | 2022-09-22
8000405 | 2021-09-22 | RSUD Tidar Magelang | Medical compressor | Durr/Sicolab Med | DSPH0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-09-22
19000331 | 2021-09-22 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0180 | Service | Fan failure | Ready to use |  | Nanda
20000144 | 2021-09-22 | RSUD Nene Mallomo | Resuscitaire | Draeger/Resuscitaire | DSNJ0047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-09-22
8000404 | 2021-09-21 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0039 | Maintenance | Penggantian flow power dan distributing valve | Ready to use |  | Amna
7000183 | 2021-09-20 | RSUD Bagas Waras | Mesin Anasthesia | Draeger/Primus | ASHL0062 | Maintenance | Total care | Ready to use |  | Jarwo
7000183 | 2021-09-20 | RSUD Bagas Waras | Mesin Anasthesia | Draeger/Primus | ASHL0062 | Maintenance | Maintenance total care | Ready to use |  | Jarwo
14000157 | 2021-09-20 | RSUD Dr. R. Sosodoro Djatikoesoemo Bojonegoro | Dental Unit | Kavo/E30TM | 10111154 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2022-09-20
9000319 | 2021-09-19 | RSUD Bolaang Mongondow Selatan | Kompressor CPAP | Draeger/Kompressor | DSPC0071 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
20000142 | 2021-09-18 | RSUD Muna | Incubator Transport | Draeger/Isolette TI500 | DSNM0014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2022-09-18
3000186 | 2021-09-17 | RSUD Kab. Temanggung | ESWL | Medispec/EM1000 | 1115056 | Service | Display counter tidak sesuai | Ready to use |  | Imam
4000321 | 2021-09-17 | RS Sentra Medika Cibinong | EECP | Lumenair/Vasomedical | 5TC1064 | Service | Touchsreen tidak sensitif, selang manset bocor, ergantian monitor lama | Ready to use |  | Bayu
7000190 | 2021-09-17 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASHL0032 | Survey | Dilakukan penggantian pressure regulator | Ready to use |  | Jarwo
7000189 | 2021-09-17 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASHL0032 | Service | Lcd monitor mati total dan dilakukan penggantian board inverter dan basis board | Ready to use | Sudah dilakukan perbaikan (70000370) | Jarwo
3000185 | 2021-09-16 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | Gambar pada USG tidak tampil | Ready to use |  | Imam
17000131 | 2021-09-16 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0917012 | Service | Monitor C-arm tidak menyala | Ready to use |  | Bahar
7000180 | 2021-09-15 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0128 | Service | Penggantian flow sensor inspirasi | Ready to use |  | Jarwo
7000179 | 2021-09-15 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0144 | Service | Penggantian flow sensor inspirasi | Ready to use |  | Jarwo
4000320 | 2021-09-15 | RSUD Dr. Drajat Prawiranegara | ESWL | Medispec/EM1000 | 0617048 | Survey | Monitor USG mati | Ready to use |  | Bayu
8000403 | 2021-09-14 | RSJD Dr. Amino Gondohutomo | AGS Absorption system | Draeger/Scowenging system | ASPB0300 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
7000178 | 2021-09-14 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASHN0040 | Service | Pengecekan power supply, pressure regulator | Ready to use | Disarankan penggantian power supply, pressure regulator, proportional valve. Sudah diberikan penawaran namun RS tidak PO karna tidak ada anggaran | Jarwo
7000177 | 2021-09-14 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKE0069 | Service | mengganti gas supply sensor | Ready to use |  | Jarwo
7000176 | 2021-09-14 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKE0070 | Service | Mengganti pressure regulation | Ready to use |  | Jarwo
16000171 | 2021-09-13 | Klinik Budihardja | Dental Unit | Kavo/1058 TM | 10010476 | Service | Bunyi alarm pada sensor buangan air menyala | Ready to use |  | Ipan
18000239 | 2021-09-13 | RSUD Samratulangi Tondano | Ventilator | Draeger/Evita V300 | ASNL0255 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-09-13
18000239 | 2021-09-13 | RSUD Samratulangi Tondano | Ventilator | Draeger/Evita V300 | ASNL0256 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-09-13
8000402 | 2021-09-10 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0039 | Survey | O2 system check fail | Ready to use |  | Amna
14000156 | 2021-09-10 | RSUD dr. H. Moh. Anwar | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/084/C1 | Service | Pendampingan ukes LPFK Solo | Ready to use |  | Satria
4000319 | 2021-09-10 | RSUD Malingping | Dental Unit | Kavo/1058TM | 10205319 | Service | suction saliva tidak bersih, cek scaller | Ready to use |  | Bayu
9000311 | 2021-09-09 | RSUD dr. Zainal Umar Sidiki | Meja Operasi | Paramount/KM 302 | 03423 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
8000452 | 2021-09-09 | RSI Sultan Agung | HFNC | Draeger/Oxymixer High Flow | QMHF20205099 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-09-09
8000452 | 2021-09-09 | RSI Sultan Agung | HFNC | Draeger/Oxymixer High Flow | QMHF20207478 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-09-09
8000452 | 2021-09-09 | RSI Sultan Agung | HFNC | Draeger/Oxymixer High Flow | QMHF20207479 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-09-09
14000155 | 2021-09-09 | RSUD Syarifah Ambami Rato Ebu | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/055/C1 | Service | Pendampingan ukes LPFK Solo | Ready to use |  | Satria
18000269 | 2021-09-09 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 0920036 | Service | Fatal dasabase error | Ready to use | Alat sudah diperbaiki. Alat berfungsi dengan baik | Sukri
11000260 | 2021-09-08 | RSUD Tenriawaru Bone | Mesin Anasthesia | Draeger/Primus | ASNH0255 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-09-08
19000327 | 2021-09-08 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0115 | Service | O2 low dan fail | Ready to use |  | Nanda
19000327 | 2021-09-08 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0148 | Service | O2 low dan fail | Ready to use |  | Nanda
8000451 | 2021-09-08 | RSUD Bumiayu | Resuscitaire | Draeger/Resuscitaire | DSPF0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-09-08
6000450 | 2021-09-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Maintenance | kontrak Service total care | Ready to use |  | Afrizal
6000450 | 2021-09-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Maintenance | pengecekan berkala total care | Ready to use |  | Afrizal
6000450 | 2021-09-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Maintenance | kontrak Service total care | Ready to use |  | Afrizal
6000450 | 2021-09-08 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Maintenance | General check | Ready To Use |  | Afrizal
5000154 | 2021-09-08 | RSCM Kencana | Dental Unit | Kavo/1058TM | 1007780 | Service | Selang angin bocor | Ready to use |  | Bilal
9000323 | 2021-09-07 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120 | VGSNF1511 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000323 | 2021-09-07 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120 | VGSNF1687 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000323 | 2021-09-07 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120 | VGSNF1501 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000323 | 2021-09-07 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120 | VGSNF1508 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000323 | 2021-09-07 | RSUD Ainun Habibie | Patient Monitor | Draeger/Vista 120 | VGSNF1696 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
13000108 | 2021-09-07 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Service | Balon EMH bocor | Ready to use |  | Sofian
19000324 | 2021-09-07 | RSUD Pandan Arang | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASKL0040 | Maintenance | pengecekan unit berkala | Ready to use |  | Nanda
19000326 | 2021-09-07 | RSUD Pandan Arang | Baby Incubator | Draeger/Isolette C2000 | NY33130 | Maintenance | Pengecekan unit berkala | Ready to use |  | Nanda
19000325 | 2021-09-07 | RSUD Pandan Arang | Patient Monitor | Draeger/Vista 120 | VTSKC0022 | Maintenance | Pengecekan unit berkala | Ready to use |  | Nanda
17000130 | 2021-09-06 | RSUD Berkah Pandeglang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0497 | Uji Fungsi | Pelaksanaan uji fungsi dan training | Ready to use |  | Bahar
15000163 | 2021-09-06 | RSUP Dr. Mohammad Hoesin | X-ray Digital | Apelem/Camargue DR | CPD361817K20 | Service | Detektor tidak tersambung | Ready to use |  | Sutan
6000448 | 2021-09-04 | RSUD DR. R. Goeteng Taroenadibrata | Incubator Transport | Draeger/Isolette TI500 | CR12077 | Service | Batterai drop | Ready to use |  | Afrizal
6000447 | 2021-09-04 | RSUD DR. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CU03922 | Service | Penggantian engsel | Ready to use |  | Afrizal
6000446 | 2021-09-04 | RSUD DR. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CU03923 | Service | Sensor inop, engsel patah | Ready to use |  | Afrizal
6000447 | 2021-09-04 | RSUD DR. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CU03924 | Service | Penggantian engsel | Ready to use |  | Afrizal
6000447 | 2021-09-04 | RSUD DR. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CU03925 | Survey | Penggantian engsel | Ready to use |  | Afrizal
20000141 | 2021-09-03 | RSUD Bahteramas | Incubator Transport | Draeger/Isolette TI500 | MH12469 | Service | power failure, baterai tidak bisa mengisi | Ready to use | Sudah dilakukan perbaikan | Irfan S
3000184 | 2021-09-02 | RSUD Sekayu | X-ray Digital | Apelem/Camargue DR | CPD33562E19 | Service | Perbaikan komputer WS | Ready to use |  | Imam
6000445 | 2021-09-02 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASFL0120 | Service | gas supply unit merah | Ready to use |  | Afrizal
4000318 | 2021-09-02 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Maintenance | Pergerakan meja stag (tidak mau bergerak) | Ready to use |  | Bayu
14000154 | 2021-09-02 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 0720015 | Service | kalibrasi probe USG | Ready to use |  | Satria
8000401 | 2021-09-02 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASMD0040 | Service | Penggnatian baterai internal | Ready to use |  | Amna
19000322 | 2021-09-02 | RSUP Dr. Kariadi | Resuscitaire | Draeger/Resuscitaire | MC1338 | Service | Error 05 | Ready to use |  | Nanda
13000113 | 2021-09-02 | RSU Haji Medan | Dental Unit | Kavo/1058TM | 10207192 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sofian | 2022-09-02
6000444 | 2021-09-01 | RSUD DR. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Fabius Tiro | ASLH0077 | Survey | Pemasangan O2 sensor dan housingnya | Ready to use |  | Afrizal
9000309 | 2021-08-31 | RSUD Noongan | Ventilator | Draeger/Savina 300 | ASND0137 | Service | Training User | Ready to use |  | Irfan R
19000318 | 2021-08-31 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0305 | Service | Baterai internal tidak ada | Ready to use | Disarankan ganti baterai atau Service set 2 thn | Nanda
19000317 | 2021-08-30 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Maintenance | O2 pressure low | Ready to use |  | Nanda
8000300 | 2021-08-30 | RSU Permata Blora | Ventilator | Draeger/Savina 300 | ASPF0151 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-30
11000258 | 2021-08-30 | RSKD Dadi | C-Arm | Apelem/Evo R+ | 55120063849 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2023-08-30
19000316 | 2021-08-30 | RSUD Bung Karno | ESMR | Draeger/Primus | ASMA0194 | Service | Unit mati total | Pending | Disarankan ganti baterai, dan Service set 3thn. Sudah diberikan penawaran namun pihak RS belum ada anggaran | Nanda
15000162 | 2021-08-30 | RSUP Dr. Mohammad Hoesin | Dental Unit | Medispec/CS-2 | 06111187 | Service | Penggantian applicator karna shock habis | Ready to use |  | Sutan
16000173 | 2021-08-30 | RSCM Kencana | Dental Unit | Kavo/E30TM | 10000659 | Maintenance | Pengecekan dental unit berkala | Ready to use |  | Ipan
16000164 | 2021-08-27 | Wisma Atlit | Bed Patient | Paramount | 90080650009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
19000315 | 2021-08-27 | RSUD Tugurejo | Incubator Transport | Draeger/Isolette TI500 | CR12079 | Service | alarm power gagal | Pending | Disarankan penggantian baterai internal. Sudah diberikan penawaran | Nanda
16000165 | 2021-08-27 | Wisma Atlit | Bed Patient | Paramount | 90080650011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
16000169 | 2021-08-27 | Wisma Atlit | Bed Patient | Paramount | 900806500002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
16000170 | 2021-08-27 | Wisma Atlit | Bed Patient | Paramount | 900806500037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
16000163 | 2021-08-27 | Wisma Atlit | Bed Patient | Paramount | 900806500049 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
16000168 | 2021-08-27 | Wisma Atlit | Bed Patient | Paramount | 900806500352 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
16000166 | 2021-08-27 | Wisma Atlit | Bed Patient | Paramount | 900806500361 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
16000167 | 2021-08-27 | Wisma Atlit | Bed Patient | Paramount | 900806500366 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
16000162 | 2021-08-27 | Wisma Atlit | Bed Patient | Paramount | 900806500457 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
16000161 | 2021-08-27 | Wisma Atlit | Bed Patient | Paramount | 19100065000737 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
8000299 | 2021-08-27 | RS JIH | Ventilator | Draeger/Babylog VN500 | ASPA0051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-27
5000153 | 2021-08-26 | RS Marinir Cilandak | X-ray Digital | Apelem/Camargue DR | CPD26753A17 | Service | Kabel control panel dimakan tikus dan kolimator kendor | Ready to use |  | Bilal
8000298 | 2021-08-26 | RSUD RA Kartini | Ventilator | Draeger/Babylog VN500 | VSMK0028 | Maintenance | Value volume tidak terdeteksi | Ready to use |  | Amna
4000317 | 2021-08-26 | RSUD Berkah Pandeglang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0497 | Service | Pelaksanaan uji paparan dan uji kesesuaian | Ready to use |  | Bayu
17000129 | 2021-08-26 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0488 | Survey | Detektor dan tube tidak sinkron | Ready to use |  | Bahar
14000153 | 2021-08-26 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Survey | Training User | Ready to use |  | Satria
6000438 | 2021-08-26 | RSUD Majenang | Infant Warmer | Draeger/Babytherm 8010 | ASPH0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2022-08-26
19000314 | 2021-08-26 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0153 | Maintenance | O2 supply lemah | Ready to use |  | Nanda
19000312 | 2021-08-25 | Hospital Siloam Semarang | Ventilator | Draeger/Babylog 2000 | ASKL0038 | Maintenance | Pengecekan berkala | Ready to use |  | Nanda
19000311 | 2021-08-25 | Hospital Siloam Semarang | Ventilator | Draeger/Savina 300 | ASKN0288 | Maintenance | Pengecekan berkala | Ready to use |  | Nanda
19000313 | 2021-08-25 | Hospital Siloam Semarang | Baby Incubator | Draeger/Isolette C2000 | CR12070 | Maintenance | pengecekan berkala | Ready to use |  | Nanda
19000313 | 2021-08-25 | Hospital Siloam Semarang | Baby Incubator | Draeger/Isolette C2000 | CR35565 | Service | pengecekan berkala | Ready to use |  | Nanda
8000297 | 2021-08-25 | RSUD RA Kartini | Ventilator | Draeger/Savina 300 | ASEK0197 | Survey | 4400000 & 4400001 failure device | Pending | Disarankan penggantian PCB power. Sudah diberikan penawaran | Amna
6000418 | 2021-08-24 | RSU PKU Muhammadiyah Bantul | Bed Patient | Paramount Bed/ 3 crank | 900806500169 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000418 | 2021-08-24 | RSU PKU Muhammadiyah Bantul | Bed Patient | Paramount Bed/ 3 crank | 900806500172 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000418 | 2021-08-24 | RSU PKU Muhammadiyah Bantul | Bed Patient | Paramount Bed/ 3 crank | 900806500190 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000418 | 2021-08-24 | RSU PKU Muhammadiyah Bantul | Bed Patient | Paramount Bed/ 3 crank | 900806500191 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000418 | 2021-08-24 | RSU PKU Muhammadiyah Bantul | Bed Patient | Paramount Bed/ 3 crank | 900806500383 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000418 | 2021-08-24 | RSU PKU Muhammadiyah Bantul | Bed Patient | Paramount Bed/ 3 crank | 900806500384 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
19000310 | 2021-08-24 | RSUD Dr. Soeselo Slawi | Ventilator | Draeger/Savina 300 | ASHD0014 | Service | power off | Ready to use |  | Nanda
11000202 | 2021-08-24 | RSUP Dr. Wahidin Sudirohusodo | Vaporizer | Draeger/Vapor2000 | ASMM0226 | Service | Savoflurance berfungsi dengan baik | Ready to use |  | Nasrul
11000202 | 2021-08-24 | RSUP Dr. Wahidin Sudirohusodo | Vaporizer | Draeger/Vapor2000 | ASNC0018 | Service | Isoflurance berfungsi dengan baik | Ready to use |  | Nasrul
19000309 | 2021-08-23 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Service | O2 sensor gagal | Ready to use |  | Nanda
8000296 | 2021-08-23 | RSJD Dr. Amino Gondohutomo | HFNC | Draeger/Oxymixer High Flow | 100820215 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-23
8000296 | 2021-08-23 | RSJD Dr. Amino Gondohutomo | HFNC | Draeger/Oxymixer High Flow | 1000820214 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-23
19000308 | 2021-08-23 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Maintenance | Titik batu di carm dan usg tidak sesuai | Ready to use |  | Nanda
6000436 | 2021-08-23 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Maintenance | kontrak Service total care | Ready to use |  | Afrizal
20000135 | 2021-08-21 | RSUD Haji Makassar | Wheel Chair | Paramount bed/ AR200B | 100017803034V | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
11000256 | 2021-08-21 | RSKD IBU DAN ANAK PERTIWI | Phototherapy | Draeger/Bililux | 1106351006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-08-21
6000435 | 2021-08-21 | RSUD Tidar Magelang | Baby Incubator | Draeger/Caleo | ASJM0063 | Service | Humidiefnot ok | Ready to use |  | Afrizal
20000134 | 2021-08-20 | RSUD Haji Makassar | Meja Operasi | Paramount bed/ KM 303 | 03435 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
11000256 | 2021-08-20 | RSKD IBU DAN ANAK PERTIWI | Phototherapy | Draeger/Bililux | 1111527005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-08-20
18000266 | 2021-08-19 | RSUD Ainun Habibie | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0160 | Service | O2 sensor dan housing rusak | Ready to use |  | Sukri
8000295 | 2021-08-19 | RSUD Limpung | Patient Monitor | Draeger/Vista 120 | VGSNE2463 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-19
8000295 | 2021-08-19 | RSUD Limpung | Patient Monitor | Draeger/Vista 120 | VGSNF0747 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-19
8000295 | 2021-08-19 | RSUD Limpung | Patient Monitor | Draeger/Vista 120 | VGSNF0762 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-19
8000295 | 2021-08-19 | RSUD Limpung | Patient Monitor | Draeger/Vista 120 | VGSNF0765 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-19
8000295 | 2021-08-19 | RSUD Limpung | Patient Monitor | Draeger/Vista 120 | VGSNF0799 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-19
8000295 | 2021-08-19 | RSUD Limpung | Patient Monitor | Draeger/Vista 120 | VGSNF1043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-19
8000295 | 2021-08-19 | RSUD Limpung | Patient Monitor | Draeger/Vista 120 | VGSNF1182 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-19
8000295 | 2021-08-19 | RSUD Limpung | Patient Monitor | Draeger/Vista 120 | VGSNF1217 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-19
11000257 | 2021-08-19 | RSUD Salewangan | Baby Incubator | Draeger/Isolette 8000 | BR07492 | Service | alarm water supply | Pending | Disarankan penggantian exaporator. Sudah diberikan penawaran | Nasrul
19000305 | 2021-08-18 | RSUD Dr. Moewardi | Mesin Anasthesia | Draeger/Fabius Plus | ASEE0202 | Survey | Leak dan O2 sensor gagal | Ready to use |  | Nanda
19000306 | 2021-08-18 | RSUD Dr. Soediran Mangun Sumarso | Resuscitaire | Draeger/Resuscitaire | BC13132 | Survey | Error 05 | Pending | Disarankan penggantian heater. Sudah diberikan penawaran namun RS belum ada anggaran | Nanda
11000255 | 2021-08-18 | RSUD Kota Makassar | Meja Operasi | Paramount Bed/KM502 | 03342 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000255 | 2021-08-18 | RSUD Kota Makassar | Meja Operasi | Paramount Bed/KM502 | 03364 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
14000152 | 2021-08-18 | RSUD Dr. Soetomo | Bed Patient | Paramount/Pa93485 | 900210000006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria
14000151 | 2021-08-18 | RSUD Dr. Soetomo | Bed Patient | Paramount/Pa93485 | 901106400010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria
6000434 | 2021-08-16 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Evita V300 | ASPC0205 | Service | upgrade kit neonatal mode ventilation | Ready to use |  | Afrizal
6000434 | 2021-08-16 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Evita V300 | ASPC0206 | Service | upgrade kit neonatal mode ventilation | Ready to use |  | Afrizal
15000164 | 2021-08-14 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Service | penggantian applicator | Ready to use |  | Sutan
6000433 | 2021-08-13 | RSUD Bagas Waras | Patient Monitor | Draeger/Vista 120 | VGSNE5312 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2022-08-13
6000432 | 2021-08-13 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASNL0106 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2022-08-13
6000432 | 2021-08-13 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASNL0107 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2022-08-13
18000267 | 2021-08-12 | RSUD Otohana | Ventilator | Draeger/Evita V300 | ASAA0007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-08-12
15000161 | 2021-08-11 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Mati total dan ganti power supply | Ready to use |  | Sutan
6000427 | 2021-08-10 | RSGM Prof. Soedomo | Mesin Anasthesia | Draeger/Primus | ASMA0129 | Service | Pemasangan pressure sensor air | Ready to use |  | Afrizal
8000293 | 2021-08-10 | RSUD Dr. R. Soeprapto Cepu | Patient Monitor | Draeger/Vista 120 | VGSNF1054 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-10
8000293 | 2021-08-10 | RSUD Dr. R. Soeprapto Cepu | Patient Monitor | Draeger/Vista 120 | VGSNF1057 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-10
15000160 | 2021-08-10 | RSUD Sekayu | X-ray Digital | Apelem/Camargue DR | CPD33562E19 | Service | CPU mati total | Ready to use |  | Sutan
19000303 | 2021-08-10 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0094 | Service | Leakage tinggi | Ready to use |  | Nanda
19000302 | 2021-08-10 | RSUP Dr. Kariadi | Resuscitaire | Draeger/Resuscitaire | MC13338 | Service | Error 05 | Ready to use | Disarankan penggantian heater | Nanda
19000301 | 2021-08-10 | RSUP Dr. Kariadi | Vaporizer | Draeger/Vapor2000 | ASNL0390 | Service | unit rusak | Ready to use | Unit diganti dengan unit baru | Nanda
19000304 | 2021-08-09 | RS Nasional Diponegoro | Baby Incubator | Draeger/Isolette C2000 | AF32536 | Service | o2 SENSOR fail | Ready to use |  | Nanda
19000304 | 2021-08-09 | RS Nasional Diponegoro | Baby Incubator | Draeger/Isolette C2000 | AF32538 | Service | o2 SENSOR fail | Ready to use |  | Nanda
19000304 | 2021-08-09 | RS Nasional Diponegoro | Baby Incubator | Draeger/Isolette C2000 | AF32539 | Service | o2 SENSOR fail | Ready to use |  | Nanda
19000304 | 2021-08-09 | RS Nasional Diponegoro | Baby Incubator | Draeger/Isolette C2000 | AF32540 | Service | o2 SENSOR fail | Ready to use |  | Nanda
8000292 | 2021-08-09 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 edition | ASEL0143 | Service | o2 fail alarm | Ready to use |  | Amna
8000291 | 2021-08-07 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Perseus A500 | ASNK0040 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-07
8000291 | 2021-08-07 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Perseus A500 | ASNK0043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-07
8000291 | 2021-08-07 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Perseus A500 | ASNM0010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-07
8000291 | 2021-08-07 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Perseus A500 | ASNM009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-08-07
3000182 | 2021-08-06 | RSUD Kab. Temanggung | C-Arm | Apelem/Evo R+ | 55115011469 | Service | Komputer C-arm tidak booting | Ready to use |  | Imam
19000300 | 2021-08-06 | RST Bhakti Wira Tamtama | Ventilator | Draeger/Savina 300 | ASHD0011 | Service | internal battery failure | Ready to use |  | Nanda
18000265 | 2021-08-06 | RSUD Dr. H. Chasan Boesoirie | Bed Patient | Paramount Bed | 900501100057 | Service | Handle dan kopling patah | Ready to use |  | Sukri
16000160 | 2021-08-04 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058TM | 10007769 | Service | Penggantian membran air yang bocor | Ready to use |  | Ipan
4000316 | 2021-08-04 | RS Sentra Medika Cibinong | EECP | Lumenair/Vasomedical | STC1064 | Maintenance | Heart rate tinggi, balon pada kaki bocor | Ready to use |  | Bayu
19000299 | 2021-08-03 | RST Bhakti Wira Tamtama | Ventilator | Draeger/Savina 300 | ASEL0194 | Service | O2 naik turun tidak stabil | Ready to use | Disarankan penggantian mixer valve oksigen. Sudah diberikan penawaran namun RS tidak PO | Nanda
18000263 | 2021-08-03 | RSUD Weda | Bed Patient | Paramount | 900408800005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000262 | 2021-08-03 | RSUD Weda | Emergency Stretcher | Paramount | 900700700094 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
9000305 | 2021-08-03 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/C2000 | AR35152 | Survey | Kerusakan pada FAN Motor | Ready to use | Disarankan penggantian fan motor | Irfan R
9000306 | 2021-08-03 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0183 | Service | Pressure regulator harus diganti | Ready to use |  | Irfan R
16000158 | 2021-08-03 | RSCM Kencana | Dental Unit | Kavo/1058TM | 1007780 | Maintenance | General check | Ready to use |  | Ipan
16000159 | 2021-08-03 | RSCM Kencana | Dental Unit | Kavo/E30TM | 10000659 | Service | General check | Ready to use |  | Ipan
6000423 | 2021-08-02 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Service | sensor check ventilator merah | Ready to use | Sudah dilakukan perbaikan (6000450) | Afrizal
6000424 | 2021-08-02 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Survey | Flow sensor rusak | Ready to use |  | Afrizal
3000183 | 2021-08-01 | RSUD Banyumas | C-Arm | Apelem/Evo R+ | 55116055585 | Service | Connector board sigma failed | Ready to use | Sudah diberikan penawaran. RS sudah PO. Sedang menunggu spare part datang | Imam
18000259 | 2021-07-31 | RSUD Jailolo | Bed Patient | Paramount | 90031080013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
17000128 | 2021-07-30 | RSAU Dr. Esnawan Antariksa | Mobile X-ray | Apelem/Rafale EV30 | 0051801391 | Survey | handle kendor, pergerakan tube keras | Ready to use |  | Bahar
14000150 | 2021-07-30 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Service | karet balon kurang mengembang | Ready to use |  | Satria
19000298 | 2021-07-30 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 S | S2SNM1420 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-30
19000298 | 2021-07-30 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 S | S2SNM1440 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-30
19000298 | 2021-07-30 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 S | S2SNM1441 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-30
19000298 | 2021-07-30 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 S | S2SNM1442 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-30
19000298 | 2021-07-30 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 S | S2SNM1466 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-30
19000298 | 2021-07-30 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 S | S2SNM1483 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-30
6000422 | 2021-07-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Maintenance | Gas analyzer fail | Ready to use |  | Afrizal
14000149 | 2021-07-29 | RSUD dr. H. Moh. Anwar | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/084/C1 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2023-07-29
3000179 | 2021-07-28 | RSUD Tidar Magelang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0496 | Service | Komputer medecom error windows; Penggantian komputer baru | Ready to use |  | Imam
11000254 | 2021-07-28 | RSUD Bahteramas | Mesin Anasthesia | Draeger/Primus | ASNH0249 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-07-28
11000254 | 2021-07-28 | RSUD Bahteramas | Mesin Anasthesia | Draeger/Primus | ASNH0250 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-07-28
6000420 | 2021-07-28 | RS Mahardika | Jaudince Meter | Draeger/JM 105 | H3505880 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2022-07-28
6000420 | 2021-07-28 | RS Mahardika | TCB | Pocking system | H7032304 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2022-07-28
19000297 | 2021-07-28 | RSUD Dr. Soehadi Prijonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-28
19000296 | 2021-07-27 | RSUD KOTA SURAKARTA | Resuscitaire | Draeger/Resuscitaire | DSNL0045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-27
19000296 | 2021-07-27 | RSUD KOTA SURAKARTA | Resuscitaire | Draeger/Resuscitaire | DSNL0046 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-27
19000296 | 2021-07-27 | RSUD KOTA SURAKARTA | Resuscitaire | Draeger/Resuscitaire | DSNL0047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-27
19000296 | 2021-07-27 | RSUD KOTA SURAKARTA | Resuscitaire | Draeger/Resuscitaire | DSNL0048 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-27
18000255 | 2021-07-26 | RSUD Boliyohuto | Bed Patient | Paramount | 9000615100008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
19000295 | 2021-07-26 | RSI Sultan Agung | Oxymixer High Flow | Draeger/HFOT | QMHF2205099 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nanda | 2022-07-26
15000159 | 2021-07-26 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | fixed merah ganti kabel fiber optic | Ready to use |  | Sutan
20000131 | 2021-07-25 | RSUD Labuang Baji | Dental Unit | Kavo/1058TM | 10207190 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2022-07-25
19000294 | 2021-07-23 | RST Bhakti Wira Tamtama | Ventilator | Draeger/Savina 300 | ASEL0194 | Service | Fraksi O2 naik turun dan tidka mencapai settingan | Ready to use | Disarankan penggantian mixer valve oksigen. Sudah diberikan penawaran namun RS tidak PO | Nanda
18000258 | 2021-07-23 | RSUD Toto Kabila | Bed Patient | Paramount | 900610700011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
16000155 | 2021-07-23 | RSUD Malingping | Dental Unit | Kavo/1058TM | 10205319 | Service | Monitor tidak menyala, | Ready to use |  | Ipan
18000257 | 2021-07-22 | RSUD Toto Kabila | C-Arm | Apelem/Evo R+ | 55121004853 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2023-07-22
20000133 | 2021-07-22 | RSUD Haji Makassar | Emergency Stretcher | Paramount Bed/PK8000 | 900509500035 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
5000152 | 2021-07-21 | RSUP Dr. M. Djamil Padang | ESWT | Medispec/ED1000 | 0713914 | Service | Penggantian applicator | Ready to use |  | Bilal
15000157 | 2021-07-19 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 06111187 | Service | Penggantian applicator | Ready to use |  | Sutan
15000158 | 2021-07-19 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Survey | Fixed merah | Ready to use | alat masih bisa digunakan jika ceilling | sutan
18000253 | 2021-07-15 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0480 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-07-15
18000253 | 2021-07-15 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0481 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-07-15
18000252 | 2021-07-15 | RSUD Bumi Panua | Ventilator | Draeger/Babylog VN500 | ASNL0110 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-07-15
16000157 | 2021-07-12 | RSCM Kencana | Dental Unit | Kavo/1058TM | 1007780 | Maintenance | General check | Ready to use |  | Ipan
16000156 | 2021-07-12 | RSCM Kencana | Dental Unit | Kavo/E30TM | 10000659 | Service | General check | Ready to use |  | Ipan
8000290 | 2021-07-08 | RSUD Brebes | Baby Incubator | Draeger/Caleo | ASHM0066 | Survey | alarm water empty | Ready to use | Disarankan penggantian conv water heater 230V. Sudah diperbaiki (6000492) | Amna
13000106 | 2021-07-08 | RSUP Dr. H. Adam Malik Medan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/083/C1 | Maintenance | Kabel LAN ke WS digigit tikus | Ready to use |  | Sofian
13000103 | 2021-07-08 | RSUP Dr. H. Adam Malik Medan | Bone Mineral Densitometer | Medilink/Medix DR | J17016M280 | Survey | Bersihkan board dan konek ulang a;at | Ready to use |  | Sofian
19000293 | 2021-07-08 | RSUP Dr. Kariadi | Mesin Vaporizer | Draeger/Vapor2000 | ASMN0346 | Service | Vaporizer bocor | Ready to use |  | Nanda
6000421 | 2021-07-07 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Fabius Tiro | ASHL0096 | Maintenance | Leak test error, O2 sensor calibration failed, baterai drop | Pending | Disarankan penggantian O2 sensor + housing, maintenance 2 years, fabius Service 3 years | Afrizal
19000288 | 2021-07-06 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHH0122 | Maintenance | General check | Ready To Use |  | Nanda
19000290 | 2021-07-06 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0179 | Maintenance | General check | Ready to use |  | Nanda
19000291 | 2021-07-06 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0180 | Maintenance | General check | Ready to use |  | Nanda
19000292 | 2021-07-06 | RSUD Tugurejo | Baby Incubator | Draeger/Caleo | ASHK0035 | Maintenance | General check | Ready to use |  | Nanda
19000289 | 2021-07-06 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Perseus | ASHL0016 | Maintenance | General check | Ready to use |  | Nanda
9000351 | 2021-07-05 | RS Ratumbuysang Manado | Bed Patient | Paramount | PA3100000127 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
19000287 | 2021-07-05 | RSUD Dr. Soehadi Prijonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASKE0184 | Service | Unit mati total | Ready to use |  | Nanda
18000236 | 2021-07-03 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0081 | Service | Gas monitor failure | Ready to use |  | Sukri
19000285 | 2021-07-03 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Fabius Plus | ARSM0127 | Maintenance | General check | Ready to use |  | Nanda
19000286 | 2021-07-03 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0066 | Maintenance | Penggantian baterai | Ready to use |  | Nanda
19000283 | 2021-07-03 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Maintenance | General check | Ready to use |  | Nanda
19000282 | 2021-07-03 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0152 | Maintenance | General check | Ready to use |  | Nanda
19000284 | 2021-07-03 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0235 | Maintenance | General check | Ready to use |  | Nanda
18000235 | 2021-07-03 | RSUD Toto Kabila | Dental Unit | Kavo/1058TM | 10205320 | Maintenance | software error | Ready to use |  | Sukri
20000128 | 2021-07-02 | RSIA Siti Fatimah | Bed Patient | Paramount Bed | 9000611900029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
4000305 | 2021-07-01 | RSUD Malingping | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0489 | Survey | Pendampingan pengurusan izin pemanfaatan alat | Ready to use |  | Bayu
8000289 | 2021-07-01 | RSUD RAA Soewondo Pati | Baby Incubator | Draeger/Isolette C2000 | DSNK0036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-07-01
8000289 | 2021-07-01 | RSUD RAA Soewondo Pati | Baby Incubator | Draeger/Isolette C2000 | DSNK0037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-07-01
19000281 | 2021-06-30 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0485 | Service | Install circuit breathing | Ready to use |  | Nanda
17000127 | 2021-06-30 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0488 | Service | Covor column tube stand rusak | Ready to use |  | Bahar
8000288 | 2021-06-30 | RSUD Sultan Fatah Demak | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0049 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-06-30
9000304 | 2021-06-30 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0166 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000304 | 2021-06-30 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0172 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000304 | 2021-06-30 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0178 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000304 | 2021-06-30 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0179 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000304 | 2021-06-30 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0165 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000304 | 2021-06-30 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0411 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000304 | 2021-06-30 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0415 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000304 | 2021-06-30 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0417 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000304 | 2021-06-30 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0419 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
20000127 | 2021-06-30 | RSUP Dr. Wahidin Sudirohusodo | Dental Unit | Kavo/1058TM | 10205313 | Service | Suction tiba tiba mati | Ready to use |  | Irfan S
8000287 | 2021-06-29 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0138 | Service | O2 sensor rusak | Ready to use |  | Amna
8000287 | 2021-06-29 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0120 | Service | O2 sensor rusak | Ready to use |  | Amna
8000287 | 2021-06-29 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Primus | URYH0049 | Service | O2 sensor rusak | Ready to use |  | Amna
15000156 | 2021-06-29 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/ED1000 | 06111187 | Service | Penggantian applicator | Ready to use |  | Sutan
9000302 | 2021-06-29 | RSUP Ratatotok Buyat | X-Ray Analog | Apelem/Camargue Analog | CPD36816520 | Service | Error Fault EEPROM S62 | Ready to use |  | Irfan S
8000286 | 2021-06-28 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 edition | ASEL0142 | Service | Alarm O2 measurement inop and neo flow | Ready to use |  | Amna
20000126 | 2021-06-28 | RSUD Labuang Baji | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0494 | Service | Film pada printer selalu tersangkut | Ready to use |  | Irfan S
8000285 | 2021-06-22 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0482 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-06-22
8000285 | 2021-06-22 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0483 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-06-22
8000285 | 2021-06-22 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0484 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-06-22
8000284 | 2021-06-21 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0014 | Service | Sensor check fail | Ready to use | Sudah diperbaiki pada pengecekan selanjutnya (8000408) | Amna
8000283 | 2021-06-19 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0482 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-06-19
8000283 | 2021-06-19 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0483 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-06-19
8000283 | 2021-06-19 | RS Indriati Solo | Ventilator | Draeger/Savina 300 | ASNL0484 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-06-19
8000282 | 2021-06-19 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0136 | Service | O2 fail | Ready to use |  | Amna
15000155 | 2021-06-19 | RSUD Dr. M. Yunus | ESWL | Medispec/EM1000 | 0917011 | Maintenance | Error 11 | Ready to use |  | Sutan
6000299 | 2021-06-17 | RSA UGM | Mesin Anasthesia | Draeger/Fabius Plus | ASEM0135 | Maintenance | Pemasangan O2 sensor | Ready to use |  | Afrizal
6000298 | 2021-06-17 | RSA UGM | Mesin Anasthesia | Draeger/Primus | ASKE0210 | Survey | General check | Ready to use |  | Afrizal
6000297 | 2021-06-17 | RSA UGM | Mesin Anasthesia | Draeger/Primus | ASNE0160 | Survey | alat lama tidak digunakan | Ready to use |  | Afrizal
6000296 | 2021-06-17 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Plus | ASNJ0038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2022-06-17
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0168 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0169 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0171 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0173 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0177 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0180 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0181 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0182 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0183 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0410 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0413 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
9000303 | 2021-06-16 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJJ0414 | Service | Penggantian baterai external unit | Ready to use |  | Irfan R
5000151 | 2021-06-15 | Klinik Budihardja | Dental Unit | Kavo/1058TM | 10010470 | Service | Handpiece tidak mutar | Ready to use |  | Bilal
18000199 | 2021-06-15 | RS Advent Manado | Stretcher | Paramount Bed/K726E | K72140207F | Service | Roda tidak bisa di rem | Ready to use | Strlu penggdventan module control electronic mainboard, valve block brake arm, gk7 connector | Irfan R
8000281 | 2021-06-14 | RSUD Ambarawa | Incubator Transport | Draeger/Isolette TI500 | KY11399 | Survey | Alarm air flow | Ready to use |  | Amna
14000148 | 2021-06-12 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Service | Training User | Ready to use |  | Satria
6000295 | 2021-06-11 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Caleo | ASKL0037 | Maintenance | Humidity error | Ready to use |  | Afrizal
18000198 | 2021-06-10 | RS Gunung Maria Tomohon | Baby Incubator | Draeger/Isolette C2000 | HC37698 | Maintenance | Pengecekan unit | Ready to use |  | Irfan R
18000198 | 2021-06-10 | RS Gunung Maria Tomohon | Baby Incubator | Draeger/Isolette C2000 | HC37708 | Service | Pengecekan unit | Ready to use |  | Irfan R
8000280 | 2021-06-10 | RSUD RA Kartini | Ventilator | Draeger/Savina 300 | ASEK0197 | Service | Alat ke restart saat running | Pending | Disarankan penggantian PCB power. Sudah diberikan penawaran | Amna
15000154 | 2021-06-10 | RSUP Dr. Mohammad Hoesin | EECP | Lumenair/Vasomedical | 06111187 | Service | Penggantian applicator | Ready to use |  | Sutan
5000150 | 2021-06-09 | RS Sentra Medika Cibinong | EECP | Lumenair/Vasomedical | 5TC1064 | Maintenance | Kabel ECG grafik tidak normal | Ready to use |  | Bilal
8000279 | 2021-06-09 | RSUD Kab Batang | Baby Incubator | Draeger/Isolette C2000 | HH36862 | Survey | Ada bunti pada alat | Ready to use |  | Amna
6000294 | 2021-06-09 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASHL0055 | Survey | alat berisik jika dipakai ke pasien | Ready to use | Alat masih bisa berfungsi namun disarankan ganti motor pump | Afrizal
6000293 | 2021-06-09 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASKJ0043 | Service | gas analyzer error | Ready to use | Disarankan penggantian Pba Mixer 2 dan flow power. Sudah diperbaiki | Afrizal
16000153 | 2021-06-09 | RSCM Kencana | Dental Unit | Kavo/1058TM | 1007780 | Maintenance | General check | Ready to use |  | Ipan
16000154 | 2021-06-09 | RSCM Kencana | Dental Unit | Kavo/E30TM | 10000659 | Service | General check | Ready to use |  | Ipan
18000256 | 2021-06-08 | RSUD Boliyohuto | Bed Patient | Paramount | 900409500015 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
17000126 | 2021-06-08 | RSUD Kab Tangerang | ESWL | Medispec/EM1000 | 0710599 | Service | C-arm tidak menyala | Ready to use | C-arm sudah menyala tapi unit masih tidak bisa digunakan untuk keseluruhan | Bahar
6000292 | 2021-06-08 | RSUD Sleman | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0128 | Maintenance | Pasnag spare part sensor PDMIX | Ready to use |  | Afrizal
8000278 | 2021-06-08 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | XRNN1632 | Maintenance | Penggantian O2 sensor | Ready to use |  | Amna
8000278 | 2021-06-08 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | XRNN1633 | Maintenance | Penggantian O2 sensor | Ready to use |  | Amna
8000278 | 2021-06-08 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | XRNN1636 | Maintenance | Penggantian O2 sensor | Ready to use |  | Amna
18000232 | 2021-06-07 | RSUD Ainun Habibie | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0160 | Maintenance | O2 sensor capsule rusak | Ready to use |  | Sukri
18000197 | 2021-06-07 | RSUD Samratulangi Tondano | Ventilator | Draeger/Savina 300 | ASJM0393 | Maintenance | General check | Ready to use |  | Irfan R
18000197 | 2021-06-07 | RSUD Samratulangi Tondano | Resuscitaire | Draeger/Resuscitaire | JV06039 | Service | General check | Ready to use |  | Irfan R
3000180 | 2021-06-04 | RSUD KRT. Setjonegoro Wonosobo | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0302 | Service | Software error; Reload software duet | Ready to use |  | Imam
3000181 | 2021-06-04 | RSUD KRT. Setjonegoro Wonosobo | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/044/C0 | Service | tidak ada respon pada RMLO | Ready to use |  | Imam
5000149 | 2021-06-03 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Survey | penambahan air pada applicator | Ready to use |  | Bilal
6000291 | 2021-06-03 | RSGM Prof. Soedomo | Mesin Anasthesia | Draeger/Primus | ASMA0129 | Service | Pemasangan backup sparepart | Ready to use |  | Afrizal
13000102 | 2021-06-03 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Service | karet membran bocor | Ready to use |  | Sofian
18000196 | 2021-06-03 | RSUD Kepulauan Talaud | Ventilator | Draeger/Babylog VN500 | ASND0072 | Service | pengecekan alat | Ready to use |  | Irfan R
5000148 | 2021-06-02 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058TM | 10007769 | Service | light cure stick tidak menyala lampunya | Ready to use |  | Bilal
6000290 | 2021-06-02 | RSA UGM | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASKH0020 | Service | Pasien flow measurement inoperable | Pending | Disarankan penggantian flow sensor | Afrizal
17000152 | 2021-05-31 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058TM | 10007769 | Service | Threeway syringe tidak berfungsi | Ready to use |  | Bahar
19000280 | 2021-05-31 | RSUP Dr. Kariadi | Baby Incubator | Draeger/Caleo | ARVC0048 | Service | cover layar monitor pecah | Ready To Use |  | Nanda
8000277 | 2021-05-28 | RS ISLAM PATI | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0114 | Survey | Layar lcd tidak menyala | Ready to use |  | Amna
19000279 | 2021-05-28 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002719 | Survey | Survey ruangan uni pindah | Ready to use |  | Nanda
19000279 | 2021-05-28 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002721 | Survey | Survey ruangan uni pindah | Ready to use |  | Nanda
19000279 | 2021-05-28 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002722 | Maintenance | Survey ruangan uni pindah | Ready to use |  | Nanda
8000276 | 2021-05-27 | RS Bhina Bhakti Husada | Phototherapy | Draeger/PT 4000 | ASJL0006 | Service | Penggantian lampu | Ready to use |  | Amna
19000278 | 2021-05-27 | RS JIH | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0125 | Service | Ganti konektor central gas | Ready to use |  | Nanda
6000288 | 2021-05-27 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Service | vaporizer tidak lulus kalibrasi BPFK | Ready to use |  | Afrizal
6000288 | 2021-05-27 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Desflurance 3000 | ASKF0421 | Service | vaporizer tidak lulus kalibrasi BPFK | Ready to use |  | Afrizal
6000289 | 2021-05-27 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLC0343 | Service | vaporizer tidak lulus kalibrasi BPFK | Ready to use |  | Afrizal
6000289 | 2021-05-27 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Service | vaporizer tidak lulus kalibrasi BPFK | Ready to use |  | Afrizal
18000195 | 2021-05-27 | RSUP Ratatotok Buyat | X-Ray Analog | Apelem/Camargue Analog | CPD36816520 | Survey | Lampu kolimator putus | Ready to use |  | Irfan R
6000287 | 2021-05-25 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Primus | ASJE0010 | Service | primus total care 2 years | Ready to use |  | Afrizal
8000275 | 2021-05-25 | RSUP Dr. Kariadi | Resuscitaire | Draeger/Resuscitaire | MC13338 | Maintenance | Skin sensor fail | Ready to use |  | Amna
14000147 | 2021-05-25 | RSUD RA Basoeni | Dental Unit | Kavo/1058TM | 10205460 | Service | selang handpiece mampet | Ready to use |  | Satria
18000194 | 2021-05-24 | RSUD Walanda Maramis | Ventilator | Draeger/Savina 300 | ASHL0291 | Maintenance | Pengecekan alat | Ready to use |  | Irfan R
18000192 | 2021-05-24 | RSUD Walanda Maramis | Mesin Anasthesia | Draeger/Primus | ASMH0120 | Maintenance | Pengecekan alat | Ready to use |  | Irfan R
18000194 | 2021-05-24 | RSUD Walanda Maramis | Ventilator | Draeger/Savina 300 | ASNL0478 | Maintenance | Pengecekan alat | Ready to use |  | Irfan R
18000194 | 2021-05-24 | RSUD Walanda Maramis | Ventilator | Draeger/Savina 300 | ASNL0479 | Maintenance | Pengecekan alat | Ready to use |  | Irfan R
18000192 | 2021-05-24 | RSUD Walanda Maramis | Patient Monitor | Draeger/Vista 120 S | SRNF1482 | Service | Pengecekan alat | Ready to use |  | Irfan R
19000277 | 2021-05-24 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0301 | Survey | mode ventilasi tidak muncul | Ready to use |  | Nanda
3000177 | 2021-05-21 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | monitor workstation mati | Ready to use |  | Imam
3000178 | 2021-05-21 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | Status GRN OUT (motor grid macet) | Ready to use |  | Imam
6000286 | 2021-05-21 | RSGM Prof. Soedomo | Mesin Anasthesia | Draeger/Primus | ASMA0129 | Service | gas delivery system blank putih | Ready to use | Disarankan penggantian sensor PP Air. Sudah diperbaiki (6000427) | Afrizal
| 2021-05-21 | RSUP Ratatotok Buyat | X-Ray Analog | Apelem/Camargue Analog | CPD36816520 | Service | Lampu kolimator amti, fuse putus | Ready to use |  | Sukri
19000275 | 2021-05-20 | RSUD Dr. Soehadi Prijonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASKE0184 | Service | general check | Ready to use | Disarankan ganti O2 sensor | Nanda
19000276 | 2021-05-20 | RSUD Dr. Soehadi Prijonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASKE0187 | Service | Ventilasi tidak bisa | Pending | Disarankan ganti modul control ventilator, sudah diberikan penawaran namun RS belum PO | Nanda
19000274 | 2021-05-19 | RS Telogorejo | Pendant | Draeger/Movita | ASDM0064 | Survey | Pendant seret dan berat ketika digerakan | Ready to use | disarankan ganti modul control pendant. Sudah diberikan penawaran namun RS tidak PO | Nanda
14000146 | 2021-05-19 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 0720015 | Maintenance | muncul error thermal fault | Ready to use |  | Satria
19000272 | 2021-05-18 | RS Colombia Asia Semarang | Baby Incubator | Draeger/Isolette C2000 | CU29567 | Maintenance | General check | Ready to use | Disarankan penggantian sensor skin | Nanda
19000272 | 2021-05-18 | RS Colombia Asia Semarang | Baby Incubator | Draeger/Isolette C2000 | CU29568 | Service | General check | Ready to use | Disarankan penggantian sensor skin | Nanda
19000273 | 2021-05-18 | RS Telogorejo | Pendant | Draeger/Movita | ASDM0064 | Maintenance | Pendant seret dan berat ketika digerakan | Ready to use | disarankan ganti modul control pendant. Sudah diberikan penawaran namun RS tidak PO | Nanda
9000180 | 2021-05-17 | RSU Sofifi | Bed Patient | Paramount Bed | PA99275 | Maintenance | General check | Ready to use |  | Irfan R
20000130 | 2021-05-17 | RSUD Haji Makassar | Bed Patient | Paramount Bed | 901100900082 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000130 | 2021-05-17 | RSUD Haji Makassar | Bed Patient | Paramount Bed | 901100900085 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000130 | 2021-05-17 | RSUD Haji Makassar | Bed Patient | Paramount Bed | 901100900088 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000130 | 2021-05-17 | RSUD Haji Makassar | Bed Patient | Paramount Bed | 901100900090 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000130 | 2021-05-17 | RSUD Haji Makassar | Bed Patient | Paramount Bed | 901100900091 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000130 | 2021-05-17 | RSUD Haji Makassar | Bed Patient | Paramount Bed | 901100900092 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000130 | 2021-05-17 | RSUD Haji Makassar | Bed Patient | Paramount Bed | 901100900093 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000130 | 2021-05-17 | RSUD Haji Makassar | Bed Patient | Paramount Bed | 901100900098 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000130 | 2021-05-17 | RSUD Haji Makassar | Bed Patient | Paramount Bed | 901100900099 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
20000130 | 2021-05-17 | RSUD Haji Makassar | Bed Patient | Paramount Bed | 901100900100 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S
15000153 | 2021-05-17 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Service | penggantian applikator karna shock habis | Ready to use |  | Sutan
6000285 | 2021-05-17 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASJL0039 | Service | kabel neonatal flow sensor putus | Ready to use | Disarankan peganggantian PATO Sensor, pressure regulator,kabel neonatal. Sudah diperbaiki | Afrizal
13000101 | 2021-05-13 | RSUD Type D Perawang | X-ray Digital | Apelem/Camargue DR | CPD30502818 | Service | Gambar kurang bagus; Kalibrasi Detektor | Ready to use |  | Sofian
6000284 | 2021-05-11 | RSUD Sleman | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0128 | Service | fresh gas O2 tidak terbaca | Ready to use | Disarankan penggantian sensor PDMIX. Dan sudah diperbaiki (6000292) | Afrizal
19000271 | 2021-05-11 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0301 | Survey | penggantian modul control | Ready to use |  | Nanda
6000283 | 2021-05-10 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0044 | Service | circle leak alarm | Ready to use | Penggantian flow sensor | Afrizal
19000269 | 2021-05-10 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Service | control eswl di ruang operator ngelag dan lama respon | Ready to use |  | Nanda
8000274 | 2021-05-10 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Maintenance | Penggantian EMH | Ready to use |  | Amna
18000231 | 2021-05-07 | RSUD Dr. H. Chasan Boesoirie | Bed Patient | Paramount Bed | 900601000009 | Maintenance | Pengecekan handling dan kopling | Ready to use | Sudah diperbaiki oleh Teknisi paramount | Sukri
8000273 | 2021-05-06 | RST Bhakti Wira Tamtama | Mesin Anasthesia |  | ARKF0012 | Maintenance | Pengecekan general | Ready to use |  | Amna
8000273 | 2021-05-06 | RST Bhakti Wira Tamtama | Ventilator | Draeger/Savina 300 | ASEL0194 | Service | Pengecekan general | Ready to use |  | Amna
6000274 | 2021-05-06 | RSUD Kab. Temanggung | Resuscitaire | Draeger/Resuscitaire | JY10904 | Service | Pemasangan spare part | Ready to use |  | Afrizal
14000141 | 2021-05-05 | RSGM Saraswati | Dental Unit | Kavo/DPS | - | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2022-05-05
8000272 | 2021-05-05 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHL0168 | Maintenance | Sistem check O2 sensor salah | Ready to use
19000267 | 2021-05-05 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Maintenance | General check | Ready to use |  | Nanda
19000267 | 2021-05-05 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0129 | Maintenance | General check | Ready to use |  | Nanda
6000282 | 2021-05-05 | RS PKU Muhammadiyah Gamping | Ventilator | Draeger/Evita V300 | ASPC0206 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2022-05-05
17000124 | 2021-05-04 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Maintenance | cone cover rusak | Ready to use |  | Bahar
19000266 | 2021-05-04 | RSUD dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0149 | Maintenance | general check | Ready to use |  | Nanda
19000266 | 2021-05-04 | RSUD dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0150 | Service | general check | Ready to use |  | Nanda
8000271 | 2021-05-04 | RSUD Dr. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0041 | Maintenance | general check | Ready to use |  | Amna
8000271 | 2021-05-04 | RSUD Dr. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | ASLJ0092 | Service | general check | Ready to use |  | Amna
8000270 | 2021-05-03 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0014 | Service | General check | Ready to use |  | Amna
19000265 | 2021-04-29 | Charlie Hospital | Ventilator | Draeger/Evita 4 Edition | ARLF0017 | Maintenance | O2 measurement failed | Ready to use |  | Nanda
8000269 | 2021-04-29 | RS Bhina Bhakti Husada | Phototherapy | Draeger/PT 4000 | ASJL0007 | Survey | ganti fluorescent lampu biru | Ready to use |  | Amna
20000125 | 2021-04-29 | RS Universitas Hasanuddin | Incubator Transport | Draeger/Isolette TI500 | KC13432 | Service | indikator baterai tidak terbaca | Ready to use | Sudah dilakukan perbaikan | Irfan S
18000191 | 2021-04-29 | RSUD Kota Manado | Bed Patient | Paramount Bed | PA66250DXD | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000191 | 2021-04-29 | RSUD Kota Manado | Bed Patient | Paramount Bed | PK1125 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000191 | 2021-04-29 | RSUD Kota Manado | Bed Patient | Paramount Bed | PK1125 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
8000268 | 2021-04-28 | RSJD Dr. Amino Gondohutomo | Mesin Anasthesia | Draeger | ASLL0145 | Service | replace Service set 2 tahun | Ready to use |  | Amna
2000139 | 2021-04-28 | RSUD Alimuddin Umar | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0402 | Maintenance | Error naik turun meja xray | Ready to use |  | Eddy
16000151 | 2021-04-27 | RSCM Kencana | Dental Unit | Kavo/1058TM | 1007780 | Maintenance | General check | Ready to use |  | Ipan
16000152 | 2021-04-27 | RSCM Kencana | Dental Unit | Kavo/1058TM | 10000659 | Service | General check | Ready to use |  | Ipan
19000264 | 2021-04-26 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Caleo | ASKL0017 | Service | alarm water empty | Ready to use |  | Nanda
19000264 | 2021-04-26 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Caleo | ASKL0022 | Maintenance | alarm water empty | Ready to use |  | Nanda
6000281 | 2021-04-26 | RSUD Panembahan Senopati Bantul | Mesin Anasthesia | Draeger/Primus | ASHJ0223 | Service | Total care | Ready to use |  | Afrizal
19000263 | 2021-04-23 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE5097 | Maintenance | Layar mati total, indikator power nyala | Ready to use |  | Nanda
4000313 | 2021-04-23 | RSUD DR. AGOESDJAM | Digital Radiografi Fluoroskopi | Apelem/Platinum | E196F0342 | Service | Pelaksanaan uji kesesuaian | Ready to use |  | Bayu
3000176 | 2021-04-22 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Service | Kabel HVG rusak | Ready to use |  | Imam
6000195 | 2021-04-22 | RSA UGM | Bone Mineral Densitometer | Medilink/Medix 90 | 330859 | Service | tube set bermasalah | Ready to use | Sudah diberikan penawaran | Afrizal
19000262 | 2021-04-22 | RSUD Ambarawa | Ventilator | Draeger/Savina 300 | ASLA0071 | Survey | Kalibrasi O2 error | Ready to use |  | Nanda
14000142 | 2021-04-22 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Service | Training User | Ready To Use |  | Satria
15000150 | 2021-04-22 | RSUD Dr. M. Yunus | ESWL | Medispec/EM1000 | 0917011 | Service | Error 11 | Ready to use |  | Sutan
16000150 | 2021-04-21 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058TM | 10007769 | Service | Angin threeway syringe tidak keluar | Ready to use |  | Ipan
8000267 | 2021-04-21 | RSI Sunan Kudus | Mesin Anasthesia | Draeger/Atlan A350 | ASNL0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-04-21
19000261 | 2021-04-21 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Fabius | ASHH0122 | Maintenance | General check | Ready to use |  | Nanda
19000261 | 2021-04-21 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0180 | Maintenance | General check | Ready to use |  | Nanda
19000260 | 2021-04-21 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHKJ0179 | Maintenance | General check | Ready to use |  | Nanda
19000259 | 2021-04-20 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Service | general check | Ready to use |  | Nanda
6000280 | 2021-04-20 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASHL0032 | Service | nebulizer merah | Ready to use |  | Afrizal
6000198 | 2021-04-17 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | ARNM0005 | Service | FiO2 hanya mencapai 30% | Ready to use | Disarankan penggantian spare part. Alat digudangkan karna penawaran terlalu mahal | Afrizal
6000198 | 2021-04-17 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | ARNM0006 | Survey | O2 pressure measurement inop | Pending | Disarankan penggantian spare part | Afrizal
3000175 | 2021-04-16 | RSUD Ulin | Bone Mineral Densitometer | Medilink/Medix 90 | J13015M269 | Service | Detector error | Ready to use |  | Imam
19000257 | 2021-04-16 | Charlie Hospital | Ventilator | Draeger/Evita 4 Edition | ARLF0017 | Survey | General check | Ready to use |  | Nanda
19000257 | 2021-04-16 | Charlie Hospital | Ventilator | Draeger/Evita 4 Edition | ARLF0024 | Survey | General check | Ready to use |  | Nanda
19000257 | 2021-04-16 | Charlie Hospital | Ventilator | Draeger/Evita 4 Edition | ARXB0314 | Service | General check | Ready to use |  | Nanda
2000138 | 2021-04-16 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | motor SID error tidak bisa naik turun | Ready to use |  | Eddy
6000279 | 2021-04-15 | RSJ Prof Dr Soerojo Magelang | Patient Monitor | Draeger/Vista 120 | VTSKF0070 | Service | Cek monitor NIBP inoperable | Ready to use |  | Afrizal
7000158 | 2021-04-15 | RSUD DR. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEM0003 | Service | touch screen tidak bisa di klik | Ready to use |  | Amna
6000200 | 2021-04-14 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0128 | Service | Flow sensor inop | Ready to use |  | Afrizal
3000174 | 2021-04-13 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helinthus BYM | HELI/0023/CB0 | Service | komputer console sering restart | Ready to use |  | Imam
19000256 | 2021-04-13 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDK0300 | Survey | O2 sensor fail | Ready to use |  | Nanda
15000149 | 2021-04-13 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Service | Penggantain applikator karna shock habis | Ready to use |  | Sutan
17000123 | 2021-04-12 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Service | Cover pecah dibawa ke jakarta, penyanggah kabel ceilling lepas | Ready to use |  | Bahar
7000157 | 2021-04-12 | RSUD DR. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | AH36595 | Maintenance | Engsel patah | Ready to use |  | Ariawan
18000223 | 2021-04-12 | RSUD KOTAMOBAGU | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0270 | Service | Hasil ekspose tidak muncul di ws | Ready to use |  | Sukri
6000199 | 2021-04-12 | RSI Klaten | Dental Unit | Kavo/E30TM | 10111152 | Service | Ganti main board (Sn menjadi: 10110731) | Ready to use |  | Afrizal
16000149 | 2021-04-12 | RSUD Karawang | Dental Unit | Kavo/1058TM | 10205314 | Service | general check | Ready to use |  | Ipan
19000255 | 2021-04-09 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDK0300 | Maintenance | Pressure regulator oksigen bocor | Ready to use |  | Nanda
6000197 | 2021-04-08 | RSUD Bagas Waras Klaten | Mesin Anasthesia | Draeger/Primus | ASHL0062 | Service | Total care | Ready to use |  | Afrizal
17000122 | 2021-04-08 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Survey | detektor tidak respon, tidak keluar xray | Ready to use |  | Bahar
6000196 | 2021-04-07 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Survey | Tembakan eswl tidak berfungsi dengan benar | Ready to use |  | Afrizal
14000141 | 2021-04-01 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Service | Training User | Ready to use |  | Satria
19000254 | 2021-04-01 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Maintenance | hasil xray tertutup balon | Ready to use |  | Nanda
19000253 | 2021-03-31 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0118 | Maintenance | general check | Ready to use | Disarankan ganti O2 sensor dan konektornya. Sudah diberikan penawaran. Sudah dilakukan perbaikan | Nanda
19000251 | 2021-03-31 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Primus | ASHM0012 | Maintenance | general check | Ready to use |  | Nanda
19000252 | 2021-03-31 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0136 | Survey | General check | Ready to use |  | Nanda
2000136 | 2021-03-31 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Survey | Training User | Ready to use |  | Eddy
6000175 | 2021-03-31 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Survey | pembersihan filter balon eswl | Ready to use |  | Afrizal
6000217 | 2021-03-31 | RSUD Tidar Magelang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0496 | Survey | windows corrupt | Ready to use |  | Afrizal
3000173 | 2021-03-30 | RSUD Kab. Temanggung | ESWL | Medispec/EM1000 | 1115056 | Service | komputer c-arm wndows tidak berfungsi | Ready to use |  | Imam
15000151 | 2021-03-30 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Survey | kabel elektron applikator putus | Ready to use |  | Sutan
6000247 | 2021-03-30 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Maintenance | Battery drop | Ready to use | Disarankan penggantian baterai internal | Afrizal
19000200 | 2021-03-29 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Service | O2 sensor fail | Ready to use |  | Nanda
6000193 | 2021-03-29 | RSUP dr. Soeradji Tirtonegoro | Jaundice Meter | Draeger/JM 105 | B3502776 | Service | Error 5 (Circuit error) | Ready to use | unit dibawa ke kantor Service draeger surabaya. Sudah diperbaiki dan dikembalikan ke rumah sakit | Afrizal
16000148 | 2021-03-29 | RSCM Kencana | Dental Unit | Kavo/1058TM | 1007780 | Maintenance | general check | Ready to use |  | Ipan
16000147 | 2021-03-29 | RSCM Kencana | Dental Unit | Kavo/E30TM | 10000659 | Service | general check | Ready to use |  | Ipan
7000156 | 2021-03-26 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHH0153 | Service | sensor leak fail | Ready to use |  | Ariawan
16000146 | 2021-03-26 | RSUD Malingping | Dental Unit | Kavo/1058TM | 10205319 | Service | Monitor tidak berfungsi pada tombolnya | Ready to use |  | Ipan
3000172 | 2021-03-25 | RSUD Gunung Jati Cirebon | X-ray Digital | Apelem/Camargue DR | IMA704014 | Service | detektor terkadang lost contact | Ready to use |  | imam
3000172 | 2021-03-25 | RSUD Gunung Jati Cirebon | X-ray Digital | Apelem/Camargue DR | IMA704014 | Maintenance | Detektor lost koneksi, aplikasi ddroc sering keluar sendiri | Ready to use |  | Imam
19000148 | 2021-03-25 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0458 | Maintenance | System check | Ready to use |  | Nanda
19000148 | 2021-03-25 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLM0572 | Maintenance | System check | Ready to use |  | Nanda
19000148 | 2021-03-25 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLN0260 | Maintenance | System check | Ready to use |  | Nanda
19000148 | 2021-03-25 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASMH0031 | Maintenance | System check | Ready to use |  | Nanda
19000149 | 2021-03-25 | RSUD K.R.M.T Wongsonegoro | Baby Incubator | Draeger/Isolette C2000 | CR35554 | Survey | cleaning dan running unit | Ready to use |  | Nanda
14000139 | 2021-03-24 | RSUD dr. H. Moh. Anwar | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/084/C1 | Service | Training User | Ready to use |  | Satria
19000147 | 2021-03-24 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | V3SHK0121 | Maintenance | LCD garis & ngeblok | Ready to use | Disarankan ganti LCD. Alat digudangkan | Nanda
3000171 | 2021-03-23 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helinthus BYM | HELI/0023/CB0 | Service | Komputer WS sering restart | Ready to use |  | Imam
19000146 | 2021-03-23 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Fabius Plus | ASJM0126 | Service | O2 sensor rusak | Ready to use |  | Nanda
19000145 | 2021-03-23 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Fabius Plus | ASJM0264 | Survey | display mati ( garis 2 ) | Ready to use |  | Nanda
6000192 | 2021-03-23 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | VTSKB0057 | Service | Board NBP | Pending | Disarankan diajukan pergantian | Afrizal
6000192 | 2021-03-23 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | VTSKF0022 | Service | mainboard inop dan board NBP | Ready to use | Disarankan diajukan pergantian. Sudah diperbaiki | Afrizal
5000147 | 2021-03-22 | RSUD Tarakan | ESWL | Medispec/EM1000 | 917012 | Service | tidak konek ke konsol | Ready to use |  | Bilal
19000144 | 2021-03-22 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0301 | Survey | Error 42.0002 | Ready to use | Alat sudah diperbaiki (8000453) | Nanda
3000170 | 2021-03-19 | RSUD RA Kartini | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/006/C0 | Survey | Error mammo identiry | Ready to use |  | Imam
19000143 | 2021-03-19 | RSJD Dr. Amino Gondohutomo | Jaundice Meter | Draeger/JM 105 | H3500431 | Service | Unit dikembalikan ke RS | Ready to use |  | Nanda
18000185 | 2021-03-19 | RSUD Liunkendage Tahuna | Mesin Anasthesia | Draeger/Fabius | ASBH0201 | Maintenance | Kerusakan selang pembuangan | Ready to use |  | Irfan R
18000184 | 2021-03-19 | RSUD Liunkendage Tahuna | Bed Patient | Paramount Bed | PK1125 | Service | pemasangan stiker box control | Ready to use |  | Irfan R
3000169 | 2021-03-18 | RSUD Tidar Magelang | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0496 | Service | Unit tidak bisa expose; Backup restrore parameter generator dan kalibrasi detector | Ready to use |  | Imam
8000266 | 2021-03-18 | RS Bhina Bhakti Husada | Phototherapy | Draeger/PT 4000 | ASJL0006 | Maintenance | General check | Ready to use |  | Amna
8000266 | 2021-03-18 | RS Bhina Bhakti Husada | Phototherapy | Draeger/PT 4000 | ASJL0007 | Maintenance | General check | Ready to use |  | Amna
6000190 | 2021-03-18 | RSUD Panembahan Senopati Bantul | Mesin Anasthesia | Draeger/Primus | ASJJ0223 | Maintenance | Total care | Ready to use |  | Afrizal
19000142 | 2021-03-18 | RSUP Dr. Kariadi | Ventilator | Draeger | ASMK0301 | Service | cleaning reconnect main power | Ready to use |  | Nanda
15000148 | 2021-03-18 | RSUP Dr. Mohammad Hoesin | ESMR | Medispec/CS-2 | 06111187 | Service | penggantian applikator karna shock habis | Ready to use |  | Sutan
17000121 | 2021-03-17 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Maintenance | EMH habis | Ready to use |  | Bahar
18000222 | 2021-03-17 | RSUD Kab. Bolaang Mongondow | Ventilator | Draeger/Savina 300 | ASKE0348 | Service | General check | Ready to use |  | Sukri
16000145 | 2021-03-16 | Klinik Budihardja | Dental Unit | Kavo/1058 TM | 10010476 | Service | Air handpiece buntu | Ready to use |  | Ipan
19000141 | 2021-03-16 | RSUD Dr. Soediran Mangun Sumarso | Ventilator | Draeger/Savina 300 | ARWL0203 | Service | Tidak ada internal baterai, tidal volume tidak konek | Pending | Disarankan Service set 6thn, oxytrace, pressure regulator. Sudah diberikan penawaran namun RS belum ada anggaran | Nanda
19000140 | 2021-03-16 | RSUD Dr. Soediran Mangun Sumarso | Resuscitaire | Draeger/Resuscitaire | KC13735 | Service | unit tidak bisa naik turun | Ready to use |  | Nanda
19000140 | 2021-03-16 | RSUD Dr. Soediran Mangun Sumarso | Resuscitaire | Draeger/Resuscitaire | NC12996 | Service | unit tidak bisa naik turun | Ready to use |  | Nanda
7000155 | 2021-03-16 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0042 | Service | Pemasangan Service set 2 tahun | Ready To Use |  | Ariawan
7000155 | 2021-03-16 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0043 | Service | Pemasangan Service set 2 tahun | Ready to use |  | Ariawan
7000155 | 2021-03-16 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASLD0042 | Service | Pemasangan Service set 2 tahun | Ready to use |  | Ariawan
6000191 | 2021-03-16 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | ARNM0005 | Service | error mixer & gas connection valve | Ready to use | Alat digudangkan karna penawaran terlalu mahal | Afrizal
6000191 | 2021-03-16 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | ARNM0006 | Service | error mixer & gas connection valve | Pending |  | Afrizal
6000191 | 2021-03-16 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | ARNM0007 | Maintenance | alat bisa digunakan | Ready to use |  | Afrizal
19000139 | 2021-03-15 | RSUD RAA Soewondo Pati | Baby Incubator | Draeger/Isolette 8000 | HF05849 | Service | Heater korosif | Pending | Sudah diberikan penawaran. RS tidak ada budget | Nanda
19000139 | 2021-03-15 | RSUD RAA Soewondo Pati | Baby Incubator | Draeger/Isolette 8000 | HF05854 | Service | General check | Ready to use |  | Nanda
18000222 | 2021-03-10 | RSUD Ainun Habibie | Resuscitaire | Draeger/Resuscitaire | DY03438 | Service | Settingan flow dan PEEP tidak sesuai | Ready to use |  | Sukri
13000078 | 2021-03-10 | RSUP Dr. H. Adam Malik Medan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/083/C1 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sofian | 2023-03-10
19000138 | 2021-03-10 | RSUD dr. R. Seotijono | Central Monitor | Draeger/Vista 120 CMS | B2JY933 | Service | Patient monitor tidak konek ke central monitor | Ready to use |  | Nanda
14000138 | 2021-03-10 | RSUD dr. H. Moh. Anwar | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/084/C1 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2023-03-10
18000221 | 2021-03-09 | RSUD Ainun Habibie | Resuscitaire | Draeger/Resuscitaire | DY03435 | Maintenance | Settingan flow dan PEEP tidak sesuai | Ready to use |  | Sukri
13000084 | 2021-03-09 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Service | bersihkan alat | Ready to use |  | Sofian
13000083 | 2021-03-09 | RSUD Dr. Fauziah | ESWL | Medispec/EM1000 | 0716249 | Maintenance | cek posisi sudut c-arm | Ready to use |  | Sofian
19000137 | 2021-03-09 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0120 | Service | Change O2 sensor | Ready to use |  | Nanda
8000265 | 2021-03-08 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Maintenance | O2 fail alarm | Ready to use |  | Amna
6000189 | 2021-03-08 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Primus | ASJE0010 | Maintenance | valve merah pada ventilator checklist internal kalibrasi | Ready to use | Back up spare part | Afrizal
18000220 | 2021-03-08 | RSUD Tani dan Nelayan | Mesin Anasthesia | Draeger/Perseus A500 | ASHK0052 | Maintenance | General check | Ready to use |  | Sukri
18000220 | 2021-03-08 | RSUD Tani dan Nelayan | Mesin Anasthesia | Draeger/Fabius Plus | ASJE0011 | Service | General check | Ready to use |  | Sukri
18000218 | 2021-03-08 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASJL0135 | Service | MV low, alat restart sendiri | Ready to use |  | Sukri
18000218 | 2021-03-08 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASJL0136 | Maintenance | MV low, alat restart sendiri | Ready to use |  | Sukri
18000220 | 2021-03-08 | RSUD Tani dan Nelayan | Mesin Anasthesia | Draeger/Perseus A501 | ASKJ0080 | Service | General check | Ready to use |  | Sukri
18000218 | 2021-03-08 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0476 | Survey | MV low, alat restart sendiri | Ready to use |  | Sukri
4000312 | 2021-03-06 | RSUP Ratatotok Buyat | X-Ray Analog | Apelem/Camargue Analog | CPD36816520 | Service | Pelaksanaan Uji Kesesuaian | Ready to use |  | Bayu
16000144 | 2021-03-04 | Klinik Budihardja | Dental Unit | Kavo/1058 TM | 10010476 | Service | Lampu tindakan mati | Ready to use |  | Ipan
17000120 | 2021-03-04 | RSUD Cilegon | CR | VITA CR | 181116901 | Maintenance | alat mati total | Ready to use | Menunggu informasi RS untuk penawaran spare part. Alat sudah dikembalikan ke RS karna penawaran spare part tidak acc | Bahar
19000135 | 2021-03-04 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 edition | ASEL0142 | Maintenance | General check | Ready to use |  | Nanda
19000136 | 2021-03-04 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 edition | ASEL0143 | Service | General check | Ready to use |  | Nanda
9000178 | 2021-03-04 | RSUD Liunkendage Tahuna | Bed Patient | Paramount Bed | 901100900080 | Service | Penggantian board set unit | Ready to use |  | Irfan R
15000146 | 2021-03-04 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Aplikasi tidak tampil | Ready to use |  | Sutan
6000194 | 2021-03-04 | RSUP dr. Soeradji Tirtonegoro | Baby Incubator | Draeger/Caleo | ASKL0037 | Maintenance | Humidity error | Ready to use | Sudah dilakukan perbaikan (6000295) | Afrizal
6000185 | 2021-03-03 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Maintenance | Pemasangan part flow power | Ready to use |  | Afrizal
6000186 | 2021-03-03 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Service | Pemasangan part flow power (ASPA0611) | Ready to use |  | Afrizal
19000199 | 2021-03-03 | RSUD K.R.M.T Wongsonegoro | Dental Unit | Kavo/1058TM | 10205318 | Maintenance | general check | Ready to use |  | Nanda
18000179 | 2021-03-02 | RS Gunung Maria Tomohon | Bed Electric | Paramount Bed | 90070120008 | Service | Remote tidak berfungsi | Ready to use |  | Irfan R
7000154 | 2021-03-02 | RSUD Kardinah Tegal | Ventilator | Draeger/Vista V300 | ASHM0104 | Maintenance | Alarm failed | Ready to use |  | Ariawan
8000264 | 2021-03-01 | RSUD RA Kartini | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/006/C0 | Maintenance | General check | Ready to use |  | Amna
14000136 | 2021-02-26 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Maintenance | Training User | Ready to use |  | Satria
6000184 | 2021-02-26 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0005 | Maintenance | Maintaining & cleaningg filter | Ready to use |  | Afrizal
6000183 | 2021-02-25 | RSKBD Panti Nugroho | Mesin Anasthesia | Draeger/Fabius Plus | ASKK0056 | Service | Bunyi dengung bocor | Ready to use |  | Afrizal
160000142 | 2021-02-25 | RSCM Kencana | Dental Unit | Kavo/E30TM | 1000659 | Maintenance | Pengecekan umum | Ready To Use |  | Ipan
16000143 | 2021-02-25 | RSCM Kencana | Dental Unit | Kavo/E30TM | 1007780 | Survey | Pengecekan umum | Ready To Use |  | Ipan
2000135 | 2021-02-24 | RSUD Alimuddin Umar | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0402 | Service | Error table | Ready to use |  | Eddy
14000137 | 2021-02-24 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 720015 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2023-02-24
19000154 | 2021-02-24 | RSUD Simo | Jaundice Meter | Draeger/JM 105 | H3505875 | Maintenance | instal software jaudience | Ready to use |  | Nanda
16000141 | 2021-02-23 | FKG Universitas Trisakti | Dental Unit | Kavo/E30TM | 10111155 | Service | pengecekan air bocor dan angin yang keluar dari table | Ready To Use |  | Ipan
19000153 | 2021-02-23 | RSUP Dr. Kariadi | Baby Incubator | Draeger/Caleo | ARUC0048 | Service | Heater bocor | Ready to use |  | Nanda
19000152 | 2021-02-22 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Service | display monitor user shutdown | Ready to use |  | Nanda
4000311 | 2021-02-19 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helinthus BYM | HELI/0023/CB0 | Service | RIS belum bisa kirim data | Ready to use |  | Bayu
6000180 | 2021-02-17 | RS PKU Muhammadiyah Yogyakarta | Ventilator | Draeger/Babylog VN500 | ASLF0055 | Service | expiratory valve merah, neonatal flow sensor merah, circuit banyak yang bocor | Ready to use | Dianjurkan membeli sensor neonatal baru | Afrizal
8000263 | 2021-02-17 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0126 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-02-17
8000263 | 2021-02-17 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Atlan A350 | ASNH0409 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-02-17
8000263 | 2021-02-17 | RSI Sultan Agung | Baby Incubator | Draeger/Isolette C2000 | DSNK0040 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-02-17
3000138 | 2021-02-15 | RSUD Gunung Jati Cirebon | X-ray Digital | Apelem/Camargue DR | IMA704014 | Service | Detector lost conncetion | Ready to use |  | Imam
18000217 | 2021-02-12 | Puskesmas Kayoa | Bed Patient | Paramount Bed | 10902710436 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000216 | 2021-02-11 | Puskesmas Gandasuli | Bed Patient | Paramount Bed | 900801800018 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
9000170 | 2021-02-11 | RSUD Ainun Habibie | Ventilator | Draeger/Evita V300 | ASLD0170 | Maintenance | Pengecekan alat | Ready to use |  | Irfan R
9000171 | 2021-02-11 | RSUD Ainun Habibie | Mesin Anasthesia | Draeger/Fabius Atlant | ASLL0160 | Maintenance | Pengecekan alat | Ready to use |  | Irfan R
9000171 | 2021-02-11 | RSUD Ainun Habibie | Mesin Anasthesia | Draeger/Fabius Atlant | ASMM0118 | Survey | Pengecekan alat | Ready to use |  | Irfan R
9000168 | 2021-02-10 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0476 | Service | Training User | Ready to use |  | Irfan R
15000145 | 2021-02-09 | RSUP Dr. Mohammad Hoesin (KSO) | CT Scan | Siemens | 77451 | Service | Ganti fan gantry | Ready to use |  | Sutan
3000167 | 2021-02-08 | RSUD DR. AGOESDJAM | Digital Radiografi Fluoroskopi | Apelem/Platinum | E196F0342 | Survey | Training User | Ready to use |  | Imam
18000213 | 2021-02-08 | Puskesmas Laiwui | Bed Patient | Paramount Bed | 900506900003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
4000310 | 2021-02-08 | RSUD Anutapura Palu | C-Arm | Apelem/Evo R+ | 55120019805 | Survey | Pelaksanaan Uji Kesesuaian | Ready to use |  | Bayu
2000134 | 2021-02-08 | RSUD DR. AGOESDJAM | Digital Radiografi Fluoroskopi | Apelem/Platinum | E196F0342 | Survey | Re-install ke ruangan baru | Ready to use |  | Eddy
4000307 | 2021-02-06 | RSUD Andi Makassau | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/031/C0 | Survey | C-arm motor error | Ready to use |  | Bayu
14000135 | 2021-02-05 | RSUD Provinsi NTB | Digital Radiografi Fluoroskopi | Apelem/Platinum | B196F0329 | Service | Tube ceilling tidak mengeluarkan x-ray | Ready to use |  | Satria
8000262 | 2021-02-04 | RSUD LIMPUNG | Central Monitor | Draeger/Vista 120 CMS | VLSNF0254 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-02-04
7000153 | 2021-02-04 | RSUI Harapan Anda | Ventilator | Draeger/Babylog VN500 | ASJC0050 | Maintenance | Nebulizer failed | Ready to use |  | Ariawan
4000309 | 2021-02-04 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 720017 | Maintenance | Pelaksaan Uji Kesesuaian | Ready to use |  | Bayu
4000309 | 2021-02-04 | RSUP Dr. Wahidin Sudirohusodo | C-Arm | Apelem/Evo R+ | 55120062848 | Survey | Pelaksaan Uji Kesesuaian | Ready to use |  | Bayu
4000308 | 2021-02-04 | RSUP Dr. Wahidin Sudirohusodo | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0495 | Maintenance | Pelaksanaan Uji Kesesuaian | Ready to use |  | Bayu
4000308 | 2021-02-04 | RSUP Dr. Wahidin Sudirohusodo | Mamografi | Metaltronica/ Helianthus DBT | HELI/0032/CBO | Survey | Pelaksanaan Uji Paparan | Ready to use |  | Bayu
18000177 | 2021-02-03 | RS Gunung Maria Tomohon | Dyalisis Chair | Paramount Bed | 103148000016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000177 | 2021-02-03 | RS Gunung Maria Tomohon | Dyalisis Chair | Paramount Bed | 202010220820 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000178 | 2021-02-03 | RS Gunung Maria Tomohon | Delivery Bed | Paramount Bed | 900103600013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000176 | 2021-02-03 | RS Gunung Maria Tomohon | Medical Trolley | Paramount Bed | 00006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000176 | 2021-02-03 | RS Gunung Maria Tomohon | Medical Trolley | Paramount Bed | 00016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000176 | 2021-02-03 | RS Gunung Maria Tomohon | Medical Trolley | Paramount Bed | 00053 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000176 | 2021-02-03 | RS Gunung Maria Tomohon | Medical Trolley | Paramount Bed | 00054 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000176 | 2021-02-03 | RS Gunung Maria Tomohon | Medical Trolley | Paramount Bed | 00055 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000176 | 2021-02-03 | RS Gunung Maria Tomohon | Medical Trolley | Paramount Bed | 00057 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000176 | 2021-02-03 | RS Gunung Maria Tomohon | Medical Trolley | Paramount Bed | 00058 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000176 | 2021-02-03 | RS Gunung Maria Tomohon | Medical Trolley | Paramount Bed | 00059 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000176 | 2021-02-03 | RS Gunung Maria Tomohon | Medical Trolley | Paramount Bed | 00061 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
15000143 | 2021-02-03 | RSUD Dr. M. Yunus | ESWL | Medispec/EM1000 | 0917011 | Service | Cek shock dan jumlah shock | Ready to use |  | Sutan
8000261 | 2021-02-03 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Maintenance | O2 sensor fail | Ready to use |  | Amna
15000142 | 2021-02-02 | RSUD Dr. M. Yunus | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/056/C1 | Service | Menemani ukes bpfk medan | Ready to use |  | Sutan
6000179 | 2021-02-01 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0128 | Service | Pasang part claim warranty | Ready to use |  | Afrizal
7000391 | 2021-01-31 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Survey | Penggantian mixer board | Ready to use |  | Jarwo
17000119 | 2021-01-28 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0488 | Survey | Training User | Ready to use |  | Bahar
4000306 | 2021-01-28 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Service | Pelaksanaan uji kesesuaian | Ready to use |  | Bayu
8000260 | 2021-01-28 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Service | Gas analyzer fail & O2 sensor fail | Ready to use |  | Amna
7000152 | 2021-01-27 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHM0104 | Maintenance | Alarm failure | Ready to use |  | Ariawan
8000259 | 2021-01-26 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASHD0034 | Maintenance | General check | Ready to use |  | Amna
8000259 | 2021-01-26 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASKL0110 | Maintenance | General check | Ready to use |  | Amna
8000259 | 2021-01-26 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASLL0147 | Service | General check | Ready to use |  | Amna
15000140 | 2021-01-26 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Maintenance | suara tembakan tidak normal | Ready to use |  | Sutan
15000139 | 2021-01-25 | RSUD Sekayu | X-ray Digital | Apelem/Camargue DR | CPD33562E19 | Service | kabel selongkong terbela | Ready to use |  | Sutan
18000173 | 2021-01-25 | RSUD Tani dan Nelayan | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0339 | Service | Monitor blue screen | Ready to use |  | Irfan R
6000178 | 2021-01-22 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0039 | Service | O2 sensor fail | Ready to use | Disarankan penggantian O2 sensor. Sudah diperbaiki include Service kontrak | Afrizal
8000258 | 2021-01-22 | RSUD Kota Salatiga | Resuscitaire | Draeger/Resuscitaire | FM09495 | Maintenance | Error 05 | Ready to use | Sudah dilakukan perbaikan | Amna
2000133 | 2021-01-21 | RSUD Malingping | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0489 | Service | Pemasangan UPS | Ready to use |  | Eddy
7000151 | 2021-01-21 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0012 | Maintenance | fresh gas failure | Ready to use | Sudah dilakukan perbaikan | Ariawan
4000304 | 2021-01-21 | RSUP Dr. Mohammad Hoesin | X-ray Digital | Apelem/Camargue DR | CPD36817K20 | Survey | Pelaksanaan Uji Kesesuain | Ready to use |  | Bayu
6000177 | 2021-01-21 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASJN0023 | Service | error expiratory valve | Ready to use |  | Afrizal
8000257 | 2021-01-20 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0154 | Maintenance | general check | Ready to use |  | Amna
15000138 | 2021-01-19 | RS. Dr. ERNALDI BAHAR - PALEMBANG | X-Ray Analog | Apelem/Camargue Analog | CPD30672C18 | Service | Cek fungsi ceilling | Ready to use |  | Sutan
7000250 | 2021-01-19 | RSUD Cilacap | Baby Incubator | Draeger/Isolette C2000 | CF32804 | Service | Engsel patah sebelah kiri depan | Ready to use |  | Ariawan
8000256 | 2021-01-19 | RSUD K.R.M.T Wongsonegoro | Dental Unit | Kavo/1058TM | 10205318 | Service | Lampu tindakan mati | Ready to use |  | Amna
7000249 | 2021-01-18 | RSUD DR. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Fabius Tiro | ASLH0077 | Service | O2 sensor failed | Ready to use |  | Ariawan
8000255 | 2021-01-18 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0033 | Survey | Alarm gas sensor fail | Ready to use |  | Amna
3000166 | 2021-01-15 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0492 | Service | Training User | Ready to use |  | Imam
8000254 | 2021-01-15 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Service | Error 1001 | Ready to use |  | Amna
7000248 | 2021-01-14 | RSUD Banyumas | Baby Incubator | Draeger/Caleo | ASHK0029 | Service | humadifier tidak bisa | Ready to use |  | Ariawan
7000247 | 2021-01-14 | RSUD Prof. Dr. Margono Soekarjo | Jaundice Meter | Draeger/JM 105 | B5502042 | Survey | charger patah | Ready to use |  | Ariawan
2000132 | 2021-01-14 | RSUD Provinsi NTB | Digital Radiografi Fluoroskopi | Apelem/Platinum | B196F0329 | Survey | Training User | Ready to use |  | Eddy
17000118 | 2021-01-14 | RSUP Dr. Mohammad Hoesin | X-ray Digital | Apelem/Camargue DR | CPD36817K20 | Survey | Training User | Ready to use |  | Bahar
14000134 | 2021-01-13 | RSUD Provinsi NTB | Digital Radiografi Fluoroskopi | Apelem/Platinum | B196F0329 | Survey | Training User | Ready to use |  | Satria
16000140 | 2021-01-12 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1086417 | Service | Mikrorotor tidak berfungsi | Ready to use |  | Ipan
14000133 | 2021-01-12 | RSUD Provinsi NTB | Mobile X-ray | Apelem/Rafale EV30 | 009/18/01273 | Service | Error overvoltage battery | Ready to use |  | Satria
20000117 | 2021-01-12 | RSUP Dr. Wahidin Sudirohusodo | C-Arm | Apelem/Evo R+ | 55120062848 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2023-01-12
19000133 | 2021-01-11 | RSUD dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0150 | Maintenance | Training User | Ready to use |  | Danu
3000165 | 2021-01-09 | RS Bhayangkara Jambi | Mobile X-ray | Apelem/Rafale EV30 | 005/19/01531 | Maintenance | Uji kesesuaian | Ready to use |  | Imam
19000132 | 2021-01-08 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0016 | Maintenance | General check | Ready to use |  | Danu
4000303 | 2021-01-07 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0488 | Survey | Pelaksaan Uji Kesesuaian | Ready to use |  | Bayu
6000174 | 2021-01-07 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASJH0127 | Service | Gas supply unit merah, O2 sensor kalibrasi merah, nebulizer merah | Ready to use | Bisa digunakan tanpa menggunakan kompressor | Afrizal
19000131 | 2021-01-06 | RS Indriati Solo | Mesin Anasthesia | Draeger/Primus | ASJF0036 | Maintenance | General check | Ready to use |  | Danu
19000131 | 2021-01-06 | RS Indriati Solo | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0063 | Maintenance | General check | Ready to use |  | Danu
19000130 | 2021-01-05 | RS Telogorejo | Ventilator | Draeger/Carina | ASLM0055 | Maintenance | General check | Ready to use |  | Danu
8000253 | 2021-01-05 | RSUD Ki Ageng Getas Pondowo Grobogan | Resuscitaire | Draeger/Resuscitaire | AC13515 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-01-05
8000253 | 2021-01-05 | RSUD Ki Ageng Getas Pondowo Grobogan | Baby Incubator | Draeger/Isolette C2000 | AC37685 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-01-05
8000252 | 2021-01-03 | RSUD K.R.M.T Wongsonegoro | Dental Unit | Kavo/1058TM | 10205318 | Service | air water block | Ready to use |  | Amna
14000131 | 2020-12-30 | RSUD dr. H. Moh. Anwar | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/084/C1 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2022-12-30
3000164 | 2020-12-30 | RSUD Cilegon | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0488 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam
19000129 | 2020-12-29 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Maintenance | General check | Ready To Use |  | Danu
8000251 | 2020-12-29 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASNH0240 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-12-29
9000163 | 2020-12-29 | RS Pratama Dinkes | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-29
9000166 | 2020-12-29 | RS Pratama Dinkes | Ventilator Transport | Draeger/Oxylog VE300 | ASNF0199 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-29
9000165 | 2020-12-29 | RS Pratama Dinkes | Ventilator | Draeger/Savina 300 | ASNL0477 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-29
9000164 | 2020-12-29 | RS Pratama Dinkes | Patient Monitor | Draeger/Vista 120 S | S2SNE0104 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-29
9000164 | 2020-12-29 | RS Pratama Dinkes | Patient Monitor | Draeger/Vista 120 S | S2SNE0106 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-29
9000164 | 2020-12-29 | RS Pratama Dinkes | Patient Monitor | Draeger/Vista 120 S | S2SNE0117 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-29
9000164 | 2020-12-29 | RS Pratama Dinkes | Patient Monitor | Draeger/Vista 120 S | S2SNE0129 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-29
18000171 | 2020-12-24 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 0920036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2022-12-24
18000214 | 2020-12-24 | RSUD Weda | Dental Unit | Kavo/1058TM | 10205459 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-24
17000116 | 2020-12-23 | RSKD Dadi | Bone Mineral Densitometer | Medilink/Medix 90 | J207H0734 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2022-12-23
16000139 | 2020-12-22 | RS Gigi dan Mulut LADOKGI TNI AL R.E. Martadinata | Dental Unit | Kavo/1058TM | 10010468 | Service | Pergantian spare part | Ready to use |  | Ipan
16000139 | 2020-12-22 | RS Gigi dan Mulut LADOKGI TNI AL R.E. Martadinata | Dental Unit | Kavo/1058TM | 10010478 | Survey | Pergantian spare part | Ready to use |  | Ipan
19000128 | 2020-12-22 | RS Nasional Diponegoro | Central Monitor | Draeger/Vista 120 CMS | VLSNF0259 | Survey | Training User | Ready to use |  | Danu
14000132 | 2020-12-22 | RSUD Syarifah Ambami Rato Ebu | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/055/C1 | Service | Pembersihan unit | Ready to use |  | Satria
19000127 | 2020-12-21 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Service | Error valve ventilator | Ready to use |  | Danu
17000117 | 2020-12-21 | RSUD Labuang Baji | Digital Radiografi Fluoroskopi | Apelem/Platinum | J206F0494 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2022-12-21
18000211 | 2020-12-18 | RSUP Ratatotok Buyat | X-Ray Analog | Apelem/Camargue Analog | CPD36816520 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-12-18
11000218 | 2020-12-18 | RSUD POLEWALI KAB. POLEWALI MANDAR | Mesin Anasthesia | Draeger/Perseus A500 | ASNJ0104 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-18
18000169 | 2020-12-18 | RSUD Pabundayan | Ventilator | Draeger/Savina 300 | ASNL0996 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-18
18000169 | 2020-12-18 | RSUD Pabundayan | Ventilator | Draeger/Savina 300 | ASNL0997 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-18
18000169 | 2020-12-18 | RSUD Pabundayan | Ventilator | Draeger/Savina 300 | ASNL0998 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-18
18000169 | 2020-12-18 | RSUD Pabundayan | Ventilator | Draeger/Savina 300 | ASNL0999 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-18
3000163 | 2020-12-17 | RSUD KOTAMOBAGU | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0270 | Service | Pindah unit | Ready To Use |  | Imam
19000126 | 2020-12-17 | RS Nasional Diponegoro | Central Monitor | Draeger/Vista 120 CMS | VLSNF0259 | Service | Pemasangan central monitor di IGD | Ready to use |  | Danu
17000115 | 2020-12-17 | RSUD Dr. Drajat Prawiranegara | Digital Radiografi Fluoroskopi | Apelem/Platinum | J2060492 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2022-12-17
19000125 | 2020-12-17 | RSUP Dr. Kariadi | Resuscitaire | Draeger/Resuscitaire | AF10368 | Service | air way pressure tidak bisa | Ready to use |  | Danu
11000213 | 2020-12-15 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Perseus A500 | ASNK0060 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-15
11000214 | 2020-12-15 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASNL0247 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-15
11000214 | 2020-12-15 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASNL0248 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-15
11000214 | 2020-12-15 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASNL0249 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-15
19000113 | 2020-12-15 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASNL0108 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-12-15
19000113 | 2020-12-15 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASNL0109 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-12-15
18000210 | 2020-12-14 | RSUD Tani dan Nelayan | Baby Incubator | Draeger/Babyleo TN500 | ASNE0053 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-14
18000209 | 2020-12-14 | RSUD Tani dan Nelayan | Ventilator | Draeger/Savina 300 | ASNL0476 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-14
6000267 | 2020-12-13 | RSUD Panembahan Senopati Bantul | Patient Monitor | Draeger/Vista 120 | V2SHK0007 | Service | Penggantian modul ECG 2.0 | Ready to use |  | Afrizal
13000077 | 2020-12-12 | RSUP Dr. H. Adam Malik Medan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/083/C1 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sofian | 2022-12-12
14000130 | 2020-12-11 | RSUD RA Basoeni | Dental Unit | Kavo/1058TM | 10205460 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2021-12-11
11000215 | 2020-12-11 | RSKD IBU DAN ANAK PERTIWI | Patient Monitor | Draeger/Vista 120 | VGSNE2294 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-11
11000215 | 2020-12-11 | RSKD IBU DAN ANAK PERTIWI | Patient Monitor | Draeger/Vista 120 | VGSNE2295 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-11
11000215 | 2020-12-11 | RSKD IBU DAN ANAK PERTIWI | Patient Monitor | Draeger/Vista 120 | VGSNE2297 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-11
11000215 | 2020-12-11 | RSKD IBU DAN ANAK PERTIWI | Patient Monitor | Draeger/Vista 120 | VGSNE5021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-11
11000215 | 2020-12-11 | RSKD IBU DAN ANAK PERTIWI | Patient Monitor | Draeger/Vista 120 | VGSNE5022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-11
11000215 | 2020-12-11 | RSKD IBU DAN ANAK PERTIWI | Patient Monitor | Draeger/Vista 120 | VGSNE5071 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-11
11000216 | 2020-12-11 | RSKD IBU DAN ANAK PERTIWI | Central Monitor | Draeger/Vista 120 CMS | VLSNK0361 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-11
11000216 | 2020-12-11 | RSKD IBU DAN ANAK PERTIWI | Central Monitor | Draeger/Vista 120 CMS | VLSNK0363 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-12-11
8000250 | 2020-12-10 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0094 | Service | Penggantian cookpit | Ready to use |  | Amna
9000157 | 2020-12-10 | RSUP Prof. Dr. R. D. Kandou Manado | Monitor Transport | Draeger/Infinity M540 | 5614876955 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-10
18000172 | 2020-12-10 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator Transport | Draeger/Oxylog VE300 | ASNJ01B4 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-10
19000124 | 2020-12-08 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0153 | Survey | self test unit, test running | Ready to use |  | Danu
18000207 | 2020-12-08 | RSUP Ratatotok Buyat | Patient Monitor | Draeger/Vista 120 | VGSNE4965 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-08
18000207 | 2020-12-08 | RSUP Ratatotok Buyat | Patient Monitor | Draeger/Vista 120 | VGSNE4968 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-08
18000207 | 2020-12-08 | RSUP Ratatotok Buyat | Patient Monitor | Draeger/Vista 120 | VGSNE4976 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-08
18000207 | 2020-12-08 | RSUP Ratatotok Buyat | Patient Monitor | Draeger/Vista 120 | VGSNE4981 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-08
18000205 | 2020-12-07 | RSUP Prof. Dr. R. D. Kandou Manado | Lampu operasi | Draeger/Polaris 200 | ASNK1069 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-07
17000114 | 2020-12-07 | RSUD Dr. Drajat Prawiranegara | Dental Unit | Kavo/1058TM | 10205461 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2021-12-07
19000123 | 2020-12-05 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Maintenance | General check up, ganti bag manual | Ready to use |  | Danu
19000123 | 2020-12-05 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0152 | Maintenance | General check up | Ready to use |  | Danu
19000123 | 2020-12-05 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0154 | Maintenance | General check up | Ready to use |  | Danu
19000123 | 2020-12-05 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0235 | Maintenance | General check up | Ready to use |  | Danu
3000162 | 2020-12-04 | RSUD Ulin | Bone Mineral Densitometer | Medilink/Medix 90 | J13015M269 | Service | tidak dapat melakukan QC | Ready to use | Sudah dibuatkan penawaran | Imam
7000246 | 2020-12-04 | RSUD Majenang | Ventilator | Draeger/Savina 300 | ASWL1000 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-12-04
18000170 | 2020-12-03 | RSUD Maria Walanda Maramis | Ventilator | Draeger/Savina 300 | ASNL0476 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-03
18000170 | 2020-12-03 | RSUD Maria Walanda Maramis | Ventilator | Draeger/Savina 300 | ASNL0478 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-03
7000245 | 2020-12-03 | RSUD Ajibarang | HFNC | Draeger/HFOT | QMHF20206751 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-12-03
9000156 | 2020-12-02 | RS Mata Sulawesi Utara | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-02
19000121 | 2020-12-01 | RSUD Dr. Soediran Mangun Sumarso | ESWL | Medispec/E3000 | 0319201 | Service | monitor mati total | Ready to use | Alat sudah diperbaiki pada pengecekan selanjutnya (19000152) | Danu
19000122 | 2020-12-01 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Survey | Problem dalam monitoring gas | Ready to use |  | Danu
18000206 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/BabyLeo TN500 | ASNF0001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-01
18000206 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/BabyLeo TN500 | ASNF0002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-01
18000206 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/BabyLeo TN500 | ASNF0004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-01
18000206 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/BabyLeo TN500 | ASNF0005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-01
18000206 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/BabyLeo TN500 | ASNF0006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-12-01
9000155 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Patient Monitor | Draeger/Vista 120 | S2SNE0984 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-01
9000155 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Patient Monitor | Draeger/Vista 120 | S2SNE0986 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-01
9000155 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Patient Monitor | Draeger/Vista 120 | S2SNE1281 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-01
9000155 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Patient Monitor | Draeger/Vista 120 | S2SNE1297 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-01
9000155 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Patient Monitor | Draeger/Vista 120 | S2SNE1300 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-01
9000155 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Patient Monitor | Draeger/Vista 120 | S2SNF1280 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-01
9000155 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Patient Monitor | Draeger/Vista 120 | S2SNF1281 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-01
9000155 | 2020-12-01 | RSUP Prof. Dr. R. D. Kandou Manado | Patient Monitor | Draeger/Vista 120 | S2SNF1297 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-12-01
9000154 | 2020-11-30 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASNH0241 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-11-30
9000154 | 2020-11-30 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASNH0242 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-11-30
9000154 | 2020-11-30 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASNH0243 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-11-30
9000154 | 2020-11-30 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASNH0244 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-11-30
9000154 | 2020-11-30 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASNH0245 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-11-30
9000154 | 2020-11-30 | RSUP Prof. Dr. R. D. Kandou Manado | Mesin Anasthesia | Draeger/Primus | ASNH0246 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-11-30
20000114 | 2020-11-30 | RSUP Dr. Wahidin Sudirohusodo | Dental Unit | Kavo/1058TM | 10205313 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2021-11-30
15000136 | 2020-11-30 | RSUP Dr. Mohammad Hoesin | X-ray Digital | Apelem/Camargue DR | CPD36817K20 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan | 2022-11-30
7000244 | 2020-11-30 | RS Bhakti Asih | Ventilator | Draeger/Savina 300 | ASNL0487 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-11-30
18000202 | 2020-11-29 | RSUP Ratatotok Buyat | Meja Operasi | Paramount Bed | 3369 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000201 | 2020-11-28 | RSUP Ratatotok Buyat | Resuscitaire | Draeger/Resuscitaire | ASND1040 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-11-28
18000201 | 2020-11-28 | RSUP Ratatotok Buyat | Resuscitaire | Draeger/Resuscitaire | JV06034 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-11-28
1000120 | 2020-11-28 | RSUD Sultan Suriansyah | Dental Unit | Kavo/1058 | 10205321 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Rizaldi | 2021-11-28
19000118 | 2020-11-26 | RSUD dr. H. Soewondo Kendal | Medical Compressor | Draeger/Medical Compressor | ASNA0051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-11-26
19000117 | 2020-11-26 | RSUD dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0149 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Danu | 2021-11-26
19000117 | 2020-11-26 | RSUD dr. H. Soewondo Kendal | Ventilator | Draeger/Babylog VN500 | ASNL0150 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Danu | 2021-11-26
19000116 | 2020-11-25 | RS Telogorejo | Mesin Anasthesia | Draeger/Primus | ASDK0223 | Service | Piston kadang macet | Ready to use |  | Danu
9000152 | 2020-11-25 | RSUD Kepulauan Talaud | Ventilator | Draeger/Babylog VN500 | ASND0072 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-11-25
7000243 | 2020-11-24 | RSUI Harapan Anda | Patient Monitor | Draeger/Vista 120 | SISNE1146 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-11-24
20000117 | 2020-11-23 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2021-11-23
20000124 | 2020-11-23 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette C2000 | DSNL0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2021-11-23
20000124 | 2020-11-23 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette C2000 | DSNL0030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2021-11-23
20000124 | 2020-11-23 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette C2000 | DSNL0037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2021-11-23
20000124 | 2020-11-23 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette C2000 | DSNL0038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2021-11-23
20000124 | 2020-11-23 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette C2000 | DSNL0040 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2021-11-23
20000124 | 2020-11-23 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette C2000 | DSNL0049 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2021-11-23
20000116 | 2020-11-22 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0032 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2021-11-22
20000115 | 2020-11-21 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan S | 2021-11-21
8000249 | 2020-11-21 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 | VGSNE5325 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-11-21
8000249 | 2020-11-21 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 | VGSNE5326 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-11-21
8000249 | 2020-11-21 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 | VGSNE5330 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-11-21
8000249 | 2020-11-21 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 | VGSNE5331 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-11-21
8000249 | 2020-11-21 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 | VGSNE5351 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-11-21
8000249 | 2020-11-21 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 | VGSNE5352 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-11-21
8000249 | 2020-11-21 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 | VGSNE5353 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-11-21
8000249 | 2020-11-21 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 | VGSNE5354 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-11-21
6000164 | 2020-11-21 | RSUD Dr. Tjitrowardojo Purworejo | Pendant | Draeger/Movita | ASHE0161 | Service | Rem tidak berfungsi dan suction bocor | Ready to use | Disarankan ganti selenoid switch | Afrizal
8000249 | 2020-11-21 | RSUD dr. R. Soetijono Blora | Patient Monitor | Draeger/Vista 120 | VMSNJ0022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-11-21
16000138 | 2020-11-20 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helinthus BYM | HELI/0023/CB0 | Service | Pindah unit | Ready to use |  | Ipan
3000161 | 2020-11-19 | RSUP Dr. H. Adam Malik Medan | Bone Mineral Densitometer | Medilink/Medix 90 | J17016M280 | Survey | Nilai QC tidak memenuhi standard | Ready to use |  | Imam
8000248 | 2020-11-19 | RSUD PANDAN ARANG | Ventilator Transport | Draeger/Oxylog VE300 | ASWJ0136 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-11-19
19000109 | 2020-11-19 | RSUP Dr. Kariadi | Patient Monitor | Draeger/Vista 120 | S48NH0664 | Service | Training User | Ready to use |  | Danu
3000160 | 2020-11-18 | RSU Haji Medan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/033/C1 | Service | Upgrade software | Ready To Use |  | Imam
19000108 | 2020-11-18 | RS Ken Saras | Mesin Anasthesia | Draeger/Fabius GS Premium | ASBC0054 | Service | Alarm O2 fail, Alarm flow sensor | Ready to use |  | Danu
19000106 | 2020-11-18 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C2000 | DSNL0026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-11-18
19000106 | 2020-11-18 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C2000 | DSNL0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-11-18
19000106 | 2020-11-18 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C2000 | DSNL0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-11-18
19000106 | 2020-11-18 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C2000 | DSNL0047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-11-18
14000129 | 2020-11-18 | RSUD Dr. Soegiri Lamongan | ESWL | Medispec/E3000 | 720015 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2022-11-18
19000107 | 2020-11-18 | RSUD Kota Salatiga | Resuscitaire | Draeger/Resuscitaire | FM09495 | Service | Lampu tindakan mati | Ready to use |  | Danu
8000247 | 2020-11-18 | RSUD PANDAN ARANG | Ventilator Transport | Draeger/Oxylog VE300 | ASNJ0136 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-11-18
6000162 | 2020-11-17 | RS Pratama Dinkes | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0048 | Service | MV tidak muncul | Ready to use |  | Afrizal
6000161 | 2020-11-16 | Klinik Sinduadi | Jaundice Meter | Draeger/JM 105 | B3502386 | Service | Error 1 | Ready to use | Alat sudah diperbaiki oleh Draeger | Afrizal
7000242 | 2020-11-16 | RSKBD Panti Nugroho | Resuscitaire | Draeger/Resuscitaire | BV14104 | Service | Error 05 | Ready to use |  | Ariawan
3000159 | 2020-11-13 | RSUD Ulin | X-ray Digital | Apelem/Camargue DR | IMA703458 | Service | unit tidak bisa expose | Ready to use |  | Imam
9000150 | 2020-11-12 | RS Gunung Maria Tomohon | Bed Icu elektrik | Paramount Bed | 90023900016 | Service | box dan remote rusak | Ready to use |  | Irfan R
6000159 | 2020-11-12 | RSUD Kab. Temanggung | Patient Monitor | Draeger/Primus Delta XL | 6008268270 | Service | NBP blocked | Ready to use |  | Afrizal
6000160 | 2020-11-12 | RSUD KRT. Setjonegoro Wonosobo | Ventilator | Draeger/Savina 300 | ASLC0400 | Survey | baterai internal drop | Ready to use |  | Afrizal
10000119 | 2020-11-12 | RSUD Ulin | Bone Mineral Densitometer | Medilink/Medix 90 | J13015M269 | Survey | Main board dibawa untuk pengecekan lebih lanjut di kantor pusat | Ready to use |  | Rizaldi
6000157 | 2020-11-11 | Klinik Sinduadi | Jaundice Meter | Draeger/JM 105 | B3502386 | Service | Error 1 | Ready to use | Alat sudah diperbaiki oleh Draeger | Afrizal
6000158 | 2020-11-11 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Maintenance | Balon ESWL pecah | Ready to use |  | Afrizal
11000209 | 2020-11-11 | RSUD Undata Palu | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0075 | Maintenance | preventive maintenance | Ready to use |  | Nasrul
11000209 | 2020-11-11 | RSUD Undata Palu | Mesin Anasthesia | Draeger/Fabius Plus | ASFN0081 | Service | preventive maintenance | Ready to use |  | Nasrul
8000246 | 2020-11-10 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | Service | ventiteller pecah | Ready to use | Sudah diperbaiki pada pengecekan selanjutnya (8000457) | Amna
19000104 | 2020-11-10 | RS Nasional Diponegoro | Central Monitor | Draeger/Vista 120 CMS | VLSNF0259 | Service | monitor no.bed ngacak | Ready to use |  | Danu
14000128 | 2020-11-10 | RSUD Provinsi NTB | Digital Radiografi Fluoroskopi | Apelem/Platinum | B196F0329 | Maintenance | General check | Ready to use |  | Satria
14000127 | 2020-11-10 | RSUD Provinsi NTB | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT032C0 | Service | tidak bisa kirim data console unit ke ws | Ready to use |  | Satria
6000156 | 2020-11-10 | RSUP Dr. Sardjito | Baby Incubator | Draeger/BabyLeo TN500 | ASMJ0226 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-11-10
15000135 | 2020-11-10 | RSUP Dr. Mohammad Hoesin | X-ray Digital | Apelem/Camargue DR | CPD36817K20 | Pra Instalasi | Pasang besi WF | Ready to use | Alat sudah bisa berfungsi dengan baik | Sutan
19000105 | 2020-11-10 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0321 | Pra Instalasi | TV tinggi dan RR tinggi | Ready to use |  | Danu
19000101 | 2020-11-09 | RSUD Kota Salatiga | Baby Incubator | Draeger/BabyLeo TN500 | ASLL0022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-11-09
19000102 | 2020-11-09 | RSUD Kota Salatiga | Incubator Transport | Draeger/Isolette TI500 | JV13718 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-11-09
3000158 | 2020-11-07 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Meja tidak bergerak | Ready to use |  | Imam
18000166 | 2020-11-07 | RSUP Prof. Dr. R. D. Kandou Manado | Bed Patient | Bed Paramont | 900706600008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
18000165 | 2020-11-07 | RSUP Prof. Dr. R. D. Kandou Manado | Bed Patient | Paramount Bed | 901005200206 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000163 | 2020-11-07 | RSUP Prof. Dr. R. D. Kandou Manado | Bed Patient | Paramount Bed | 901005200226 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000164 | 2020-11-07 | RSUP Prof. Dr. R. D. Kandou Manado | Bed Patient | Paramount Bed | 901005L00224 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
16000137 | 2020-11-06 | FKG Universitas Trisakti | Dental Unit | Kavo/E30TM | 1011149 | Service | Pemasangan support povitable | Ready to use |  | Ipan
16000137 | 2020-11-06 | FKG Universitas Trisakti | Dental Unit | Kavo/E30TM | 10111155 | Maintenance | Pemasangan support povitable | Ready to use |  | Ipan
19000103 | 2020-11-06 | RS Nasional Diponegoro | Central Monitor | Draeger/Vista 120 CMS | VLSNF0259 | Service | Setting central ke ruang ICU | Ready to use |  | Danu
10000117 | 2020-11-06 | RSUD Ulin | Mamografi | Metaltronica/Helinthus BYM | HELI/0048/CB1 | Service | muncul error rot anode // fuse terbakar | Ready to use |  | Rizaldi
17000113 | 2020-11-06 | RSUP Dr. M. Djamil Padang | ESWL | Medispec/E3000 | 0613835 | Service | Kalibrasi USG | Ready to use |  | Bahar
6000152 | 2020-11-05 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASJH0127 | Service | Penggantian pressure regulator | Ready to use |  | Afrizal
6000152 | 2020-11-05 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASJH0128 | Service | Penggantian pressure regulator | Ready to use |  | Afrizal
10000116 | 2020-11-04 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | X-ray saat expose menggunakan detector wireless tidak keluar | Ready to use |  | Rizaldi
17000112 | 2020-11-04 | RSUP Dr. M. Djamil Padang | ESMR | Medispec/ED1000 | 0713914 | Survey | Pemasangan scissors arm | Ready to use |  | Bahar
10000115 | 2020-11-03 | RSUD Ulin | Mamografi | Metaltronica/Helinthus BYM | HELIDBT/048/CB1 | Service | error pada RDT, anide 92-074 fail dan mammo tidak responding, muncul kode error pada mammo | Ready to use |  | Rizaldi
14000126 | 2020-11-03 | RSUP SANGLAH DENPASAR | Bed Patient | Paramount Bed | 9010005200199 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria
7000241 | 2020-11-02 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ARWT0079 | Service | Penggantian pressure regulator | Ready to use |  | Ariawan
18000208 | 2020-10-30 | RSUD Toto Kabila | UPS | Yoshiga/GPM 10 | 501201094180jc30001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-10-30
18000208 | 2020-10-30 | RSUD Toto Kabila | UPS | Yoshiga/GPM 10 | 5012010941JC300007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-10-30
8000245 | 2020-10-27 | RSUD RA Kartini | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/006/C0 | Service | Uji fungsi plat Detektor | Ready to use |  | Amna
19000112 | 2020-10-26 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C2000 | DSNL0026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-10-26
19000112 | 2020-10-26 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C2000 | DSNL0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-10-26
19000112 | 2020-10-26 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C2000 | DSNL0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-10-26
19000112 | 2020-10-26 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C2000 | DSNL0047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-10-26
30000157 | 2020-10-26 | RSUD RA Kartini | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/006/C0 | Service | detector not ready | Ready to use |  | Imam
18000167 | 2020-10-25 | RSUD Noongan | Pendant | Draeger/Movita | ASND0131 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-10-25
18000167 | 2020-10-25 | RSUD Noongan | Pendant | Draeger/Movita | ASND0132 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-10-25
18000167 | 2020-10-25 | RSUD Noongan | Pendant | Draeger/Movita | ASND0135 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-10-25
18000167 | 2020-10-25 | RSUD Noongan | Pendant | Draeger/Movita | ASND0136 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-10-25
18000167 | 2020-10-25 | RSUD Noongan | Pendant | Draeger/Movita | ASND0140 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-10-25
7000240 | 2020-10-24 | RSIA Bunda Arif | Jaundice Meter | Draeger/JM 105 | B3502042 | Service | back up | Ready to use |  | Ariawan
11000206 | 2020-10-24 | RSUD Kab. Bombana | Mesin Anasthesia | Draeger/Fabius Plus | ASHD0010 | Service | O2 fail, flow sensor fail, ventilation leak | Ready to use |  | Nasrul
11000207 | 2020-10-24 | RSUD Kab. Bombana | Ventilator | Draeger/Savina 300 | ASHD0012 | Service | baterai internal failure | Ready to use |  | Nasrul
130000070 | 2020-10-23 | RSUD Simeulue | Bone Mineral Densitometer | Medilink/Medix DR | F207H0702 | Service | Update software 3D | Ready to use |  | Sofian
10000114 | 2020-10-23 | RSUD Ulin | Bone Mineral Densitometer | Medilink/Medix 90 | J15016M269 | Service | minta upgrade software 3D dan stuck di loading 40% | Ready to use |  | Rizaldi
10000110 | 2020-10-23 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0234 | Service | konsol error | Ready to use |  | Rizaldi
10000112 | 2020-10-23 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0234 | Service | whole leg tube tidak bisa digerakan | Ready to use |  | Rizaldi
10000109 | 2020-10-23 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0234 | Service | ceilling tube tidak presisi | Ready to use |  | Rizaldi
10000111 | 2020-10-23 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | detector kadang konek kadang tidak | Ready to use |  | Rizaldi
17000110 | 2020-10-23 | RSUP Dr. M. Djamil Padang | ESWL | Medispec/E3000 | 0613835 | Service | Install standard electrode & scissors arm | Ready to use |  | Bahar
17000111 | 2020-10-23 | RSUP Dr. M. Djamil Padang | X-ray Digital | Apelem/Camargue DR | IMA7003458 | Service | Pemasangan tensator spring | Ready to use |  | Bahar
6000264 | 2020-10-23 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | V2SHF0244 | Service | NIBP bermasalah | Ready to use |  | Afrizal
8000244 | 2020-10-22 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKL0006 | Service | Keypad tidak berfungsi | Ready to use | Sudah dilakukan perbaikan 15 December 2020 | Amna
19000190 | 2020-10-22 | RSUD Kajen | Mesin Anasthesia | Draeger/Primus | ASLD0023 | Maintenance | muncul indicator soda lime | Ready to use |  | Danu
19000190 | 2020-10-22 | RSUD Kajen | Mesin Anasthesia | Draeger/Primus | ASLD0024 | Maintenance | general check | Ready to use |  | Danu
19000190 | 2020-10-22 | RSUD Kajen | Mesin Anasthesia | Draeger/Primus | ASMA0191 | Service | general check | Ready to use |  | Danu
19000189 | 2020-10-20 | RSUD Ambarawa | Ventilator | Draeger/Savina 300 | ASLA0071 | Service | Error air way pressure low | Ready to use |  | Danu
11000204 | 2020-10-20 | RSUD Konawe | Patient Monitor | Draeger/Vista 120 | VGSNE5056 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-10-20
11000204 | 2020-10-20 | RSUD Konawe | Patient Monitor | Draeger/Vista 120 | VGSNE5065 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-10-20
17000108 | 2020-10-19 | RSUD Dr. M. Yunus | ESWL | Medispec/EM1000 | 0917011 | Service | Generator restart pada energi 9/10 | Ready to use |  | Bahar
17000109 | 2020-10-19 | RSUD Harapan dan Doa | X-ray Digital | Apelem/Camargue DR | CPD33564E19 | Service | Ceiling saat dijalankan longitudinal mengalami seret/pengereman; Mengecangkan baut baut pada rel karna alat menabrak baut | Ready to use |  | Bahar
6000232 | 2020-10-19 | RSUP dr. Soeradji Tirtonegoro | Infant Warmer | Draeger/Babytherm 8010 | ASMM0045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-10-19
6000232 | 2020-10-19 | RSUP dr. Soeradji Tirtonegoro | Infant Warmer | Draeger/Babytherm 8010 | ASMM0046 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-10-19
19000188 | 2020-10-16 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Service | kalibrasi flow ins, eksp, penggantian karpet | Ready to use |  | Danu
9000149 | 2020-10-16 | RSUP Ratatotok Buyat | X-ray Digital | Apelem/Camargue DR | CPD32474K18 | Service | Fuse putus | Ready to use |  | Irfan
10000113 | 2020-10-15 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Maintenance | komputer blue screen | Ready to use |  | Rizaldi
8000243 | 2020-10-14 | RS ISLAM PATI | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0114 | Service | general check | Ready to use |  | Amna
19000187 | 2020-10-14 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Caleo | ASKL0022 | Service | alarm water empty | Ready to use |  | Danu
18000160 | 2020-10-13 | RSUP Prof. Dr. R. D. Kandou Manado | Patient Monitor | Draeger/Vista 120 | S2NE1146 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-10-13
18000159 | 2020-10-13 | RSUP Prof. Dr. R. D. Kandou Manado | Patient Monitor | Draeger/Vista 120 | S2NE1494 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-10-13
11000205 | 2020-10-13 | RSUD Konawe | Central Monitor | Draeger/Vista 120 CMS | VLSNF0258 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-10-13
11000205 | 2020-10-13 | RSUD Konawe | Central Monitor | Draeger/Vista 120 CMS | VLSNF0271 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-10-13
19000186 | 2020-10-12 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Survey | Sensor fail ketika self-test | Ready to use | Sudah diperbaiki pada pengecekan selanjutnya (19000188) | Danu
10000108 | 2020-10-12 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | Komputer delay di aplikasi tidak dapat dimatikan | Ready to use |  | Rizaldi
16000136 | 2020-10-10 | RSK Dharmais | Bed Patient | Paramount Bed | 000286 - 000247 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
16000134 | 2020-10-10 | RSK Dharmais | Bed Patient | Paramount Bed | 00042 - 000287 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
14000125 | 2020-10-09 | RSUD Dr. H. Slamet Martodirjo | ESWL | Medispec/E3000 | 1017902 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2022-10-09
19000185 | 2020-10-09 | RSUP Dr. Kariadi | Baby Incubator | Draeger/Caleo | ARUC0048 | Service | Bocor | Ready to use |  | Danu
16000132 | 2020-10-08 | RSUP Cipto Mangunkusumo | Bed Patient | Paramount Bed | 00082 - 000281 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
16000131 | 2020-10-08 | RSUP Cipto Mangunkusumo | Bed Patient | Paramount Bed | 00120 - 00165 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
16000133 | 2020-10-08 | RSUP Cipto Mangunkusumo | Bed Patient | Paramount Bed | 00153 - 00116 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan
8000242 | 2020-10-07 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASKF0024 | Service | gas analyzer fail | Ready to use | Backup spare part dari Draeger. Menunggu PO dari RS | Amna
7000238 | 2020-10-07 | RSUI Harapan Anda | Ventilator | Draeger/Savina 300 | ARRF0049 | Service | MV tidak muncul | Ready To Use |  | Ariawan
7000238 | 2020-10-07 | RSUI Harapan Anda | Ventilator | Draeger/Savina 300 | ASMM0052 | Survey | O2 measurement failed | Ready to use |  | Ariawan
5000144 | 2020-10-07 | RSUP Persahabatan | Bed Patient | Paramount Bed | 9009024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000143 | 2020-10-07 | RSUP Persahabatan | Bed Patient | Paramount Bed | 9000902400314 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000142 | 2020-10-07 | RSUP Persahabatan | Bed Patient | Paramount Bed | 9000902400314 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
8000241 | 2020-10-06 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHD0187 | Service | battery check dan O2 fail | Ready to use |  | Amna
8000241 | 2020-10-06 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHL0168 | Service | battery check dan O2 fail | Ready to use |  | Amna
7000237 | 2020-10-06 | RSUD Bumiayu | Incubator Transport | Draeger/Isolette TI500 | AY11232 | Service | Battery failure | Ready to use |  | Ariawan
19000183 | 2020-10-06 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Fabius GS Premium | ARUC0012 | Survey | test leak fail | Ready to use |  | danu
8000239 | 2020-10-05 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Primus | ASNE0161 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-10-05
8000239 | 2020-10-05 | RSUD Dr. R. Soetrasno | Baby Incubator | Draeger/Isolette C2000 | DSNF0058 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-10-05
19000177 | 2020-10-05 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Service | penggantian flow power | Ready To Use |  | Danu
19000181 | 2020-10-05 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Survey | Penggantian blok VGC | Ready To Use | Sudah diperbaiki pada pengecekan selanjutnya (19000177) | Danu
19000182 | 2020-10-05 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Service | sensor check merah dan breathing circuit kuning | Ready To Use | Sudah diperbaiki pada pengecekan selanjutnya (19000177) | Danu
7000236 | 2020-10-05 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | DSNF0002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-10-05
7000236 | 2020-10-05 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | DSNF0060 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-10-05
7000236 | 2020-10-05 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | DSNF0061 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-10-05
7000236 | 2020-10-05 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | DSNF0063 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-10-05
7000236 | 2020-10-05 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette C2000 | DSNF0064 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-10-05
7000236 | 2020-10-05 | RSUD Kardinah Tegal | Incubator Transport | Draeger/Isolette TI500 | KC13430 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-10-05
7000236 | 2020-10-05 | RSUD Kardinah Tegal | Incubator Transport | Draeger/Isolette TI500 | KC13431 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-10-05
19000184 | 2020-10-05 | RSUD Kota Salatiga | Resuscitaire | Draeger/Resuscitaire | FM09495 | Service | Blok kit korner patah | Ready to use |  | Danu
6000215 | 2020-10-03 | RSUD Wonosari | Baby Incubator | Draeger/Isolette C2000 | FH36366 | Maintenance | Humidity error | Ready to use |  | Afrizal
19000180 | 2020-10-03 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0066 | Maintenance | General check | Ready to use |  | Danu
19000180 | 2020-10-03 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Maintenance | General check | Ready to use |  | Danu
19000180 | 2020-10-03 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0152 | Maintenance | General check | Ready to use |  | Danu
19000180 | 2020-10-03 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0154 | Service | General check | Ready to use |  | Danu
19000179 | 2020-10-02 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0153 | Service | Penggantian konektor N2O | Ready to use |  | Danu
5000141 | 2020-10-01 | RS Pusat Otak Nasional | Bed Patient | Paramount Bed | 900090241 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000139 | 2020-10-01 | RS Pusat Otak Nasional | Bed Patient | Paramount Bed | 900090241 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000140 | 2020-10-01 | RS Pusat Otak Nasional | Bed Patient | Paramount Bed | 900090241 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
19000178 | 2020-10-01 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASKJ0036 | Service | FIO2 fail, gas sensor fail, error leak | Ready to use | Sudah diperbaiki pada pengecekan selanjutnya (19000177) | Danu
14000124 | 2020-10-01 | RSUD Sumbawa (KSO) | CT Scan | Siemens/Somato 90 Top | 120196 | Survey | Pra-install unit CT-Scan | Ready to use | Alat sudah dapat berfungsi dengan baik | Satria
6000213 | 2020-09-30 | RSU Gunung Sawo | Lampu operasi | Draeger/Polaris 200 | ASNF0031 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-09-30
17000107 | 2020-09-30 | RSUD Dr. Adjidarmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Service | console dan komputer tidak menyala saat pagi hari | Ready to use |  | Bahar
6000212 | 2020-09-30 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Maintenance | filter mampet | Ready to use |  | Afrizal
19000176 | 2020-09-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASJC0044 | Maintenance | General check | Ready to use |  | Danu
19000176 | 2020-09-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASJM0044 | Maintenance | General check | Ready to use |  | Danu
19000176 | 2020-09-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASLN0008 | Service | General check | Ready to use |  | Danu
19000175 | 2020-09-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0079 | Maintenance | monitor touch screen error | Ready to use |  | Danu
19000176 | 2020-09-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0079 | Maintenance | General check | Ready To Use |  | Danu
19000176 | 2020-09-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0085 | Maintenance | General check | Ready to use |  | Danu
19000176 | 2020-09-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0094 | Maintenance | General check | Ready to use |  | Danu
3000155 | 2020-09-29 | RSUD Banyumas | C-Arm | Apelem/Evo R+ | 55116055585 | Service | kamera tidak menampilkan image saat expose | Ready to use |  | Imam
11000203 | 2020-09-29 | RSUD Anutapura Palu | Patient Monitor | Draeger/Vista 120 | VGSNE5024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-29
5000134 | 2020-09-29 | RS universitas Indonesia | Bed Patient | Paramount Bed | 00213 - 229 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000135 | 2020-09-29 | RS universitas Indonesia | Bed Patient | Paramount Bed | 00213 - 229 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000136 | 2020-09-29 | RS universitas Indonesia | Bed Patient | Paramount Bed | 00213 - 229 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000137 | 2020-09-29 | RS universitas Indonesia | Bed Patient | Paramount Bed | 00213 - 229 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
5000138 | 2020-09-29 | RS universitas Indonesia | Bed Patient | Paramount Bed | 00213 - 229 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal
6000211 | 2020-09-29 | RSUP dr. Soeradji Tirtonegoro | Jaundice Meter | Draeger/JM 105 | B3502776 | Maintenance | Pasca kalibrasi | Ready to use |  | Afrizal
19000174 | 2020-09-28 | RS Indriati Solo | Mesin Anasthesia | Draeger/Primus | ASJF0036 | Maintenance | Alarm water trap | Ready to use |  | Danu
19000155 | 2020-09-28 | RS Telogorejo | Ventilator | Draeger/Carina | ASLM0055 | Maintenance | General check | Ready to use |  | Danu
8000240 | 2020-09-28 | RSUD RA Kartini | UPS | Yoshiga/HPT2 | 501101236060KA100009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-09-28
18000158 | 2020-09-25 | RSUD Samratulangi Tondano | Resuscitaire | Draeger/Resuscitaire | JV06039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-09-25
17000106 | 2020-09-25 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0917012 | Service | EMH shock sudah expired | Ready to use |  | Bahar
7000235 | 2020-09-24 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0082 | Maintenance | Gas sensor failure | Ready to use |  | Ariawan
19000157 | 2020-09-24 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0152 | Maintenance | Pemasangan/Penggantian selang gas N2O | Ready to use |  | Danu
19000157 | 2020-09-24 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0154 | Service | Pemasangan/Penggantian selang gas N2O | Ready to use |  | Danu
6000208 | 2020-09-24 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | V2SHE0274 | Service | Pemasangan ECG Limb wires 3 lead, ecg trunk cable 3 lead | Ready to use |  | Afrizal
19000156 | 2020-09-24 | RSUP Dr. Kariadi | Dental Unit | Kavo | 1002721 | Service | air di gelas dan pembuangan tidak keluar | Ready to use |  | Danu
19000158 | 2020-09-23 | RS Indriati Solo | Mesin Anasthesia | Draeger/Primus | ASJF0036 | Service | gas servo bocor | Ready to use |  | Danu
19000158 | 2020-09-23 | RS Indriati Solo | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0063 | Service | gas servo bocor | Ready to use |  | Danu
11000201 | 2020-09-23 | RSUD I Lagaligo | Bed Patient | Paramount Bed | 900708500007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
7000234 | 2020-09-23 | RSUD Banyumas | C-Arm | Apelem/Evo R+ | 55116055585 | Service | board interface sigma | Ready to use | Sudah dilakukan perbaikan (3000155) | Ariawan
13000071 | 2020-09-23 | RSUD Aceh Singki | C-Arm | Apelem/Evo R+ | 55120018804 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Sofian | 2022-09-23
6000202 | 2020-09-23 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | V2SHF0244 | Service | Pemasangan SpO2 finger sensor, ECB cable. NIBP cuff | Ready to use |  | Afrizal
6000201 | 2020-09-23 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | V2SHK006 | Service | Pemasangan part ECG trunk cable | Ready to use |  | Afrizal
18000156 | 2020-09-22 | RSUD Noongan | Battery detector | Trixell | DF01359 | service | Ganti battery detector | Ready to use |  | Sukri
9000148 | 2020-09-21 | RSUD KOTAMOBAGU | Examination Table | Paramount Bed/PK-1115 |  | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
7000233 | 2020-09-21 | RSUD DR. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | AH36567 | Service | Rusak part belakang | Ready to use |  | Ariawan
7000233 | 2020-09-21 | RSUD DR. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | AH36595 | Service | Rusak part belakang | Ready to use |  | Ariawan
18000155 | 2020-09-21 | RSUD Samratulangi Tondano | Mesin Anasthesia | Draeger/Fabius Plus | ASNE0076 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-09-21
9000147 | 2020-09-21 | RSUD Kab. Bolaang Mongondow | Incubator Transport | Draeger/Isolette TI500 | JV13723 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-09-21
19000159 | 2020-09-21 | RSUP Dr. Kariadi | Vacuum Suction | Durr/Tornado 1 | H334579100 | Service | alat macet | Ready to use |  | Danu
6000203 | 2020-09-21 | RSUP Dr. Sardjito | Baby Incubator | Draeger/BabyLeo TN500 | ASHL0031 | Service | Pemasangan membran complete, kit air filter, dan filter pad | Ready to use |  | Afrizal
6000210 | 2020-09-21 | RSUP Dr. Sardjito | Baby Incubator | Draeger/BabyLeo TN500 | ASHL0032 | Maintenance | Pemasangan membran complete, kit air filter, dan filter pad | Ready to use |  | Afrizal
6000214 | 2020-09-21 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASJL0039 | Service | Pemasangan membran complete, kit air filter, dan filter pad | Ready to use |  | Afrizal
11000131 | 2020-09-19 | RSUD Andi Djemma Masamba | Baby Incubator | Draeger/Isolette C2000 | DV38324 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-19
11000131 | 2020-09-19 | RSUD Andi Djemma Masamba | Baby Incubator | Draeger/Isolette C2000 | DV38327 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-19
11000131 | 2020-09-19 | RSUD Andi Djemma Masamba | Baby Incubator | Draeger/Isolette C2000 | DV38330 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-19
11000162 | 2020-09-18 | RSUD SYEKH YUSUP KAB. GOWA | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0024 | Service | ventilator fail | Ready to use |  | Nasrul
11000162 | 2020-09-18 | RSUD SYEKH YUSUP KAB. GOWA | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0028 | Survey | ventilator fail | Ready to use |  | Nasrul
100000104 | 2020-09-18 | RSUD Ulin | Bone Mineral Densitometer | Medilink/Medix 90 | J13016M269 | Service | stuck loading pada monitor | Ready to use |  | Rizaldi
180000154 | 2020-09-17 | RSUD Liunkendage Tahuna | Stretcher | Paramount bed/PK5114X | 5003690027 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
180000153 | 2020-09-17 | RSUD Liunkendage Tahuna | Ventilator Transport | Draeger/Oxylog VE300 | ASNF0195 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-09-17
180000153 | 2020-09-17 | RSUD Liunkendage Tahuna | Ventilator Transport | Draeger/Oxylog VE300 | ASNF0196 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-09-17
11000200 | 2020-09-17 | RSUD Lakipadada | Patient Monitor | Draeger/Vista 120 | VGSNE5031 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-17
11000200 | 2020-09-17 | RSUD Lakipadada | Patient Monitor | Draeger/Vista 120 | VGSNE5035 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-17
7000230 | 2020-09-16 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Perseus A500 | ASHL0011 | Service | Gas sensor failure | Ready to use |  | Ariawan
10000105 | 2020-09-16 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0291 | Service | tombol pada konsol tidak dapat dipencet | Ready to use |  | Rizaldi
6000204 | 2020-09-16 | RSUP dr. Soeradji Tirtonegoro | Charger | Draeger/JM 105 | 004977 | Service | alarm time periodic calibration | Ready to use |  | Afrizal
6000205 | 2020-09-16 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASLL0067 | Service | Expiratory valve blink red | Ready to use |  | Afrizal
6000204 | 2020-09-16 | RSUP dr. Soeradji Tirtonegoro | Jaundice Meter | Draeger/JM 105 | B3502776 | Service | alarm time periodic calibration | Ready to use |  | Afrizal
6000204 | 2020-09-16 | RSUP dr. Soeradji Tirtonegoro | Docking station | Draeger/JM 105 | B7035379 | Test Function | alarm time periodic calibration | Ready to use |  | Afrizal
9000141 | 2020-09-15 | RSUP Prof. Dr. R. D. Kandou Manado | Incubator transport | Draeger/Isolette TI500 | JV13717 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-09-15
4000062 | 2020-09-15 | RSUD Gunung Jati Cirebon | X-ray Digital | Apelem/Camargue DR | 704014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2022-09-15
19000160 | 2020-09-14 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0152 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-14
19000160 | 2020-09-14 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0153 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Danu | 2021-09-14
19000160 | 2020-09-14 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0154 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-14
18000152 | 2020-09-14 | RSUD Weda | UPS | Yoshiga/HPT3 | 5.012e+20 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-09-14
11000199 | 2020-09-12 | RSUD Djafar Harun | Ventilator Transport | Draeger/Oxylog VE300 | ASNF0193 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-12
3000154 | 2020-09-11 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Service | Display USG pada layar C-Arm error | Ready to use |  | Imam
7000232 | 2020-09-11 | RSUI Harapan Anda | HFNC | Draeger/HFOT | QMHF20205096 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Ariawan | 2021-09-11
4000061 | 2020-09-11 | RSUD Gunung Jati Cirebon | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/071/C1 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2022-09-11
14000123 | 2020-09-11 | RSUD Dr. Soetomo | UPS | Yoshiga/GPI 33-80kVA | 501201141680L4800003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2021-09-11
6000206 | 2020-09-11 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Survey | Flow power bermasalah | Ready to use |  | Afrizal
3000152 | 2020-09-10 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helinthus BYM | HELI/0023/CB0 | Service | Komputer WS1 rusak operating sistem windows | Ready to use |  | Imam
3000153 | 2020-09-10 | Laboratorium Klinik Pramita | Bone Mineral Densitometer | Stratos (Non Gratia) | H11015D245 | Maintenance | Nilai QC tidak memenuhi standard | Ready to use | Sudah di follow up ke distributor merk tersebut namun tidak ada respon | Imam
19000162 | 2020-09-10 | RS Triharsi | Mesin Anasthesia | Draeger/Primus | ASMH0121 | Maintenance | general check | Ready to use |  | Danu
19000162 | 2020-09-10 | RS Triharsi | Mesin Anasthesia | Draeger/Fabius Plus | ASML0038 | Service | general check | Ready to use |  | Danu
7000231 | 2020-09-10 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHH0153 | Service | back up | Ready To Use |  | Ariawan
11000197 | 2020-09-10 | RSUD Kabupaten Konawa | Patient Monitor | Draeger/Vista 120 | VGSNE4214 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-10
11000197 | 2020-09-10 | RSUD Kabupaten Konawa | Patient Monitor | Draeger/Vista 120 | VGSNE4592 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-10
11000197 | 2020-09-10 | RSUD Kabupaten Konawa | Patient Monitor | Draeger/Vista 120 | VGSNE5002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-10
11000197 | 2020-09-10 | RSUD Kabupaten Konawa | Patient Monitor | Draeger/Vista 120 | VGSNE5019 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-10
11000197 | 2020-09-10 | RSUD Kabupaten Konawa | Patient Monitor | Draeger/Vista 120 | VGSNE5020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-10
11000197 | 2020-09-10 | RSUD Kabupaten Konawa | Patient Monitor | Draeger/Vista 120 | VGSNE5051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-10
9000143 | 2020-09-10 | RSUD Kab. Bolaang Mongondow | Medical cart | Paramount bed | KY40QR0420G | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
19000161 | 2020-09-10 | RSUD KOTA SURAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0112 | Service | Maintenance mesin anasthesi | Ready to use |  | Danu
6000249 | 2020-09-10 | RSUD Tidar Magelang | Incubator Transport | Draeger/Isolette TI500 | CY11377 | Service | batterai drop | Ready to use |  | Afrizal
19000164 | 2020-09-09 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE2950 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
19000164 | 2020-09-09 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE2989 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
19000164 | 2020-09-09 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE2994 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
19000164 | 2020-09-09 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE5093 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
19000164 | 2020-09-09 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE5094 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
19000164 | 2020-09-09 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE5095 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
19000164 | 2020-09-09 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE5096 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
19000164 | 2020-09-09 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE5097 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
19000164 | 2020-09-09 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE5098 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
19000164 | 2020-09-09 | RS Nasional Diponegoro | Patient Monitor | Draeger/Vista 120 | VGSNE5100 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
19000163 | 2020-09-09 | RS Nasional Diponegoro | Central Monitor | Draeger/Vista 120 CMS | VLSNF0255 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
19000163 | 2020-09-09 | RS Nasional Diponegoro | Central Monitor | Draeger/Vista 120 CMS | VLSNF0259 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-09-09
16000129 | 2020-09-08 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1086410 | Service | Selang air buntu semua | Ready to use |  | Ipan
16000130 | 2020-09-08 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1086417 | Service | bow mampet | Ready to use |  | Ipan
6000250 | 2020-09-08 | RSUD Tidar Magelang | Incubator Transport | Draeger/Isolette TI500 | DSNF0009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-09-08
6000250 | 2020-09-08 | RSUD Tidar Magelang | Incubator Transport | Draeger/Isolette TI500 | DSNF0010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-09-08
8000238 | 2020-09-08 | RSUD RA Kartini | Ventilator | Draeger/Savina 300 | ASNH0788 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-09-08
6000248 | 2020-09-08 | RST dr. Soedjono Magelang | Patient Monitor | Draeger/Vista 120 | VGSNE2999 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-09-08
6000248 | 2020-09-08 | RST dr. Soedjono Magelang | Patient Monitor | Draeger/Vista 120 | VGSNE3000 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-09-08
6000248 | 2020-09-08 | RST dr. Soedjono Magelang | Patient Monitor | Draeger/Vista 120 | VGSNE3149 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-09-08
6000248 | 2020-09-08 | RST dr. Soedjono Magelang | Patient Monitor | Draeger/Vista 120 | VGSNE3157 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-09-08
6000248 | 2020-09-08 | RST dr. Soedjono Magelang | Patient Monitor | Draeger/Vista 120 | VGSNE5036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-09-08
6000248 | 2020-09-08 | RST dr. Soedjono Magelang | Patient Monitor | Draeger/Vista 120 | VGSNE5039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-09-08
6000248 | 2020-09-08 | RST dr. Soedjono Magelang | Patient Monitor | Draeger/Vista 120 | VGSNE5045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-09-08
7000229 | 2020-09-07 | RSIA Bunda Arif | Jaundice Meter | Draeger/JM 105 | B3501432 | Service | Error 05 | Ready to use |  | Ariawan
8000237 | 2020-09-07 | RSUD LIMPUNG | Central Monitor | Draeger/Vista 120 CMS | VLSNF0254 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-09-07
13000068 | 2020-09-05 | RSUD Simeulue | Bone Mineral Densitometer | Medilink/Medix DR | F207H0702 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sofian | 2022-09-05
5000133 | 2020-09-04 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Service | Penggantian applicator | Ready To Use |  | Bilal
19000166 | 2020-09-04 | RSUD PANDAN ARANG | Resuscitaire | Draeger/Resuscitaire | CF10509 | Service | Error 05 | Ready to use |  | Danu
18000151 | 2020-09-04 | RSUD Tani dan Nelayan | Ventilator Transport | Draeger/Oxylog VE300 | ASNF0192 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-09-04
19000165 | 2020-09-04 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Service | Penggantian flow power | Ready to use |  | Danu
3000151 | 2020-09-03 | RSUD Provinsi NTB | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/032/C0 | Service | Operating system window bermasalah | Ready to use |  | Imam
2000131 | 2020-09-03 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Survey | Blue screen pada layar console | Ready to use |  | Eddy
14000122 | 2020-09-03 | RSUD Provinsi NTB | Mobile X-ray | Apelem/Rafale EV30 | 009/18/01273 | Survey | Roda belakang terkelupas | Ready to use |  | Satria
14000122 | 2020-09-03 | RSUD Provinsi NTB | Mobile X-ray | Apelem/Rafale EV30 | 009/18/01274 | Service | Roda belakang terkelupas | Ready to use |  | Satria
16000128 | 2020-09-02 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058TM | 10007769 | Service | Penggantain control element | Ready to use |  | Ipan
11000196 | 2020-09-02 | RSUD Anutapura Palu | Baby Incubator | Draeger/Isolette C2000 | DV38343 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-02
11000196 | 2020-09-02 | RSUD Anutapura Palu | Baby Incubator | Draeger/Isolette C2000 | DV38346 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-02
11000196 | 2020-09-02 | RSUD Anutapura Palu | Baby Incubator | Draeger/Isolette C2000 | DV38349 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-02
11000196 | 2020-09-02 | RSUD Anutapura Palu | Baby Incubator | Draeger/Isolette C2000 | DV38350 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-02
11000196 | 2020-09-02 | RSUD Anutapura Palu | Baby Incubator | Draeger/Isolette C2000 | DV38351 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-02
11000195 | 2020-09-02 | RSU MADANI - PALU | Resuscitaire | Draeger/Resuscitaire | ASNA0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-02
11000194 | 2020-09-02 | RSU MADANI - PALU | Resuscitaire | Draeger/Resuscitaire | AV14204 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-09-02
19000167 | 2020-09-02 | RSUD Dr. Soediran Mangun Sumarso | Baby Incubator | Draeger/BabyLeo TN500 | ASLJ0004 | Service | General unit | Ready to use |  | Danu
8000236 | 2020-09-02 | RSUD Kab Batang | Central Monitor | Draeger/Vista 120 CMS | VLSMF0146 | Service | lost connection | Ready to use |  | Amna
7000227 | 2020-09-01 | RSUD Breber | Mesin Anasthesia | Draeger/Primus | ASHH0153 | Service | gas analyzer fail | Ready to use | Sudah dilakukan perbaikan (7000231) | Ariawan
7000227 | 2020-09-01 | RSUD Breber | Mesin Anasthesia | Draeger/Primus | ASHK0151 | Maintenance | gas analyzer fail | Ready To Use |  | Ariawan
19000168 | 2020-09-01 | RSUD KOTA SURAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0112 | Service | Pemasangan konektor N2O | Ready to use |  | Danu
8000235 | 2020-09-01 | RSUD RA Kartini | C-Arm | Apelem/Evo R+ | 55120027813 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2022-09-01
6000246 | 2020-09-01 | RSUD Panembahan Senopati Bantul | Patient Monitor | Draeger/Vista 120 | VGSNE | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-09-01
2000130 | 2020-09-01 | RSUD RA Kartini | C-Arm | Apelem/Evo R+ | 55120027813 | Service | Training User | Ready To Use |  | Eddy
6000245 | 2020-08-31 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASJL0038 | Service | Pemasangan membran complete, kit air filter, parficle filter, filter pad | Ready to use |  | Afrizal
6000241 | 2020-08-31 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | V2SHK0006 | Service | Pemaasangan sensor O2 | Ready to use |  | Afrizal
16000127 | 2020-08-28 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058TM | 10007769 | Service | Air pada bowl menguap | Ready to use |  | Ipan
6000243 | 2020-08-28 | RSA UGM | Resuscitaire | Draeger/Resuscitaire | CV14028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-08-28
11000193 | 2020-08-27 | RSUD Lasinrang | Medical Compressor | Draeger/Medical Compressor | ASNA0045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-08-27
9000141 | 2020-08-27 | RSUD KOTAMOBAGU | Bed Patient | paramount Bed | 900614700001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000142 | 2020-08-27 | RSUD KOTAMOBAGU | Patient Monitor | Draeger/Delta XL | M518596 | Service | Selang pompa NIBP terlepas | Ready to use |  | Irfan R
11000192 | 2020-08-27 | RSUD Lasinrang | Resuscitaire | Draeger/Resuscitaire | AV14026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-08-27
11000191 | 2020-08-27 | RSUD Lasinrang | Baby Incubator | Draeger/Isolette C2000 | DV38328 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-08-27
11000191 | 2020-08-27 | RSUD Lasinrang | Baby Incubator | Draeger/Isolette C2000 | DV38331 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-08-27
7000225 | 2020-08-27 | RSUD Ajibarang | Patient Monitor | Draeger/Vista 120 | VGSNE5023 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-08-27
8000234 | 2020-08-27 | RSUD LIMPUNG | Patient Monitor | Draeger/Vista 120 | VGSNE2121 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-08-27
8000234 | 2020-08-27 | RSUD LIMPUNG | Patient Monitor | Draeger/Vista 120 | VGSNE4202 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-08-27
8000234 | 2020-08-27 | RSUD LIMPUNG | Patient Monitor | Draeger/Vista 120 | VGSNE4211 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-08-27
8000234 | 2020-08-27 | RSUD LIMPUNG | Patient Monitor | Draeger/Vista 120 | VGSNE4215 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-08-27
8000234 | 2020-08-27 | RSUD LIMPUNG | Patient Monitor | Draeger/Vista 120 | VGSNE4216 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-08-27
8000234 | 2020-08-27 | RSUD LIMPUNG | Patient Monitor | Draeger/Vista 120 | VGSNE5067 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-08-27
16000126 | 2020-08-26 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | E50100 | Maintenance | Air mikrorotor nobel osseoset tidak keluar | Ready to use |  | Ipan
19000170 | 2020-08-26 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASMA0087 | Service | General check | Ready to use |  | Danu
19000171 | 2020-08-26 | RSUP Dr. Kariadi | Resuscitaire | Draeger/Resuscitaire | AF10368 | Service | Manometer tidak akurat | Ready to use |  | Danu
19000171 | 2020-08-26 | RSUP Dr. Kariadi | Resuscitaire | Draeger/Resuscitaire | MC13337 | Maintenance | Manometer tidak akurat | Ready to use |  | Danu
19000170 | 2020-08-26 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0304 | Maintenance | General check | Ready to use |  | Danu
19000172 | 2020-08-25 | RS JIH | Mesin Anasthesia | Draeger/Primus | ASMA0146 | Maintenance | general check | Ready to use |  | Danu
5000132 | 2020-08-25 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Service | Pengecekan applicator | Ready to use |  | Bilal
7000226 | 2020-08-25 | RSUD Kardinah Tegal | Ventilator Transport | Draeger/Oxylog VE300 | ASNF0194 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-08-25
7000226 | 2020-08-25 | RSUD Kardinah Tegal | Ventilator Transport | Draeger/Oxylog VE300 | ASNF0197 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-08-25
7000226 | 2020-08-25 | RSUD Kardinah Tegal | Ventilator Transport | Draeger/Oxylog VE300 | ASNF0198 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-08-25
6000242 | 2020-08-25 | RSUD Panembahan Senopati Bantul | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/027/C0 | Survey | Training User | Ready to use |  | Afrizal
17000105 | 2020-08-25 | RSUP Dr. M. Djamil Padang | Bone Mineral Densitometer | Medilink/Medix 90 | J13015270 | Service | Alat tidak terhubung dengan workstation | Ready To Use |  | Bahar
11000217 | 2020-08-24 | RSUP Dr. Wahidin Sudirohusodo | C-Arm | Apelem/Evo R+ | 55120062848 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Nasrul | 2022-08-24
8000233 | 2020-08-24 | RSUD RA Kartini | Medical Compressor | Draeger/Medical Compressor | ASNA0035 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-08-24
8000233 | 2020-08-24 | RSUD RA Kartini | Medical Compressor | Draeger/Medical Compressor | ASNA0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-08-24
7000224 | 2020-08-24 | RS Dadi Keluarga | Resuscitaire | Draeger/Resuscitaire | GV14027 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-08-24
18000150 | 2020-08-22 | RSU Sofifi | Bed Patient | Paramount Bed | 910700200290 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000149 | 2020-08-22 | RSU Sofifi | Bed Patient | Paramount Bed | 9000804600009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
19000249 | 2020-08-19 | RS Triharsi | Ventilator | Draeger/Evita V300 | ASHJ0064 | Maintenance | General check | Ready to use |  | Danu
18000148 | 2020-08-19 | RSUD Tobelo | Stretcher | Paramount bed/PK5114X | 104118000014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000147 | 2020-08-19 | RSUD Tobelo | Bed Patient | Paramount Bed | KFZ09 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000146 | 2020-08-18 | RSUD Tobelo | Bed Patient | Paramount Bed | 9000701200134 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
19000247 | 2020-08-18 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSNE2727 | Maintenance | Pemasangan breket monitor | Ready to use |  | Danu
19000247 | 2020-08-18 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSNE2728 | Maintenance |  | Ready to use |  | Danu
19000248 | 2020-08-18 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VLSMF0152 | Maintenance | General check | Ready to use |  | Danu
11000189 | 2020-08-18 | RSUD Sayang Rakyat | UPS | Fujielectric/UPS7000HX-T4 | 501201141100L4500001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-08-18
70000222 | 2020-08-18 | RSUD Hj. Anna Lasmanah | Incubator Transport | Draeger/Isolette TI500 | HC13268 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-08-18
70000222 | 2020-08-18 | RSUD Hj. Anna Lasmanah | Incubator Transport | Draeger/Isolette TI500 | JV13698 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-08-18
70000222 | 2020-08-18 | RSUD Hj. Anna Lasmanah | Incubator Transport | Draeger/Isolette TI500 | KC13433 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-08-18
8000232 | 2020-08-18 | RSUD Dr. Soeprapto Cepu | Baby Incubator | Draeger/Isolette C2000 | DV38354 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-08-18
9000140 | 2020-08-18 | RSUD Ainun Habibie | Meja Operasi | Paramount Bed | KM30303254 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
6000241 | 2020-08-13 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASKE0007 | Survey | Pemasangan membran complete, kit air filter, parficle filter, filter pad | Ready to use |  | Afrizal
3000150 | 2020-08-12 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Service | Penggantian hard disk | Ready to use |  | Imam
14000121 | 2020-08-12 | Medicelle Clinic | Breast Biopsy System | BARD/ENCOR ULTRA E5230 | DYCZT514 | Service | Ada suara beep saat tindakan | Ready to use |  | Satria
11000187 | 2020-08-12 | RSUD I Lagaligo | Ventilator | Draeger/Babylog VN500 | ASKJ0005 | Service | baterai internal failure | Ready to use |  | Nasrul
11000188 | 2020-08-12 | RSUD I Lagaligo | Patient Monitor | Draeger/Vista 120 | V3SHN0074 | Service | SPO2 failure | Ready to use |  | Nasrul
11000188 | 2020-08-12 | RSUD I Lagaligo | Patient Monitor | Draeger/Vista 120 | V3SHN0098 | Service | SPO2 failure | Ready to use |  | Nasrul
19000245 | 2020-08-12 | RSUD Kajen | Resuscitaire | Draeger/Resuscitaire | HF10402 | Service | Alarm 05 | Ready to use |  | Danu
7000223 | 2020-08-12 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0072 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-08-12
7000223 | 2020-08-12 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0073 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-08-12
7000221 | 2020-08-10 | RSUD DR. R. Goeteng Taroenadibrata | Medical Compressor | Draeger/Medical Compressor | ASJL0056 | Service | bunyi keras pada kompresor unit | Ready to use |  | Ariawan
18000142 | 2020-08-10 | RSUD Kab. Bolaang Mongondow | Ventilator | Draeger/Savina 300 | ASKE0348 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-08-10
18000142 | 2020-08-10 | RSUD Kab. Bolaang Mongondow | Ventilator Transport | Draeger/Oxylog VE300 | ASKM0090 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-08-10
18000144 | 2020-08-06 | RSUD Dr. M. M. Dunda | Medical Compressor | Draeger/Medical Compressor | ASJK0004 | Service | Pressure air tidak cukup | Ready to use |  | Sukri
6000207 | 2020-08-06 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Service | O2, CO2, volume agent N2O blok kuning | Ready to use |  | Afrizal
18000141 | 2020-08-05 | RSUD Ainun Habibie | Mesin Anasthesia | Draeger/Atlan A350 | ASMM0118 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-08-05
6000239 | 2020-08-05 | RSGM Prof. Soedomo | Mesin Anasthesia | Draeger/Primus | ASMA0129 | Service | connector hilang | Ready to use |  | Afrizal
6000239 | 2020-08-05 | RSGM Prof. Soedomo | Pendant | Draeger/Movita | ASMK0348 | Service | Pendant bocor | Ready to use |  | Afrizal
18000141 | 2020-08-05 | RSUD Ainun Habibie | Medical Compressor | Draeger/Medical Compressor | ASNA0004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-08-05
7000220 | 2020-08-05 | RS Dadi Keluarga | Baby Incubator | Draeger/Isolette C2000 | DV38326 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-08-05
19000243 | 2020-08-05 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSNE2728 | Service | Pemsangan keranjang breket | Ready to use |  | Danu
18000143 | 2020-08-05 | RSUD Dr. M. M. Dunda | C-Arm | Apelem/Evo R+ | 55118013636 | Service | Alat mati total, ganti main board, power supply board, generator board | Ready to use |  | Sukri
15000130 | 2020-08-05 | RSUD Raden Mattaher | X-Ray Analog | Apelem/Camargue Analog | CPD28153F17 | Service | Tidak bisa expose | Ready to use |  | Sutan
14000120 | 2020-08-04 | RSUD Lamongan | UPS | Yoshiga/HPT3 | 501101236060KA100008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2021-08-04
19000242 | 2020-08-04 | RSUD Kota Salatiga | Phototherapy | Draeger/PT 4000 | ASEJ0009 | Service | Penggantian lampu phototherapy | Ready to use |  | Danu
19000241 | 2020-08-04 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette 8000 | CF06001 | Service | door lock rusak | Ready to use |  | Danu
14000118 | 2020-08-04 | RSUD Lamongan | UPS | Yoshiga/HPT3 | 501101236070K6200044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2021-08-04
14000119 | 2020-08-04 | RSUD Dr. Soegiri Lamongan | UPS | Yoshiga/HPT2 | 501101236060KA100007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2021-08-04
3000148 | 2020-07-30 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Service | Error disk image, reload software duet | Ready To Use |  | Imam
11000184 | 2020-07-30 | RSUP Tadjudin Chalid | Mesin Anasthesia | Draeger/Primus | ASJL0023 | Service | self test tidak komplit | Ready to use | Sudah dilakukan perbaikan | Nasrul
11000185 | 2020-07-30 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0138 | Service | batterai internal fail | Ready to use |  | Nasrul
11000185 | 2020-07-30 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0392 | Service | batterai internal fail | Ready to use |  | Nasrul
11000185 | 2020-07-30 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0395 | Service | batterai internal fail | Ready to use |  | Nasrul
11000185 | 2020-07-30 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0396 | Service | batterai internal fail | Ready to use |  | Nasrul
9000136 | 2020-07-30 | RSD Tidore | Stretcher | Paramount bed/PK5114X | 11900122 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
3000147 | 2020-07-29 | RSUD Panembahan Senopati Bantul | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/027/C0 | Service | Setting unit ke PACS dan RIS | Ready To Use |  | Imam
3000149 | 2020-07-29 | RSUD Merah Putih | X-Ray Analog | Apelem/Camargue Analog | IMA703979 | Service | Spinning rotor rusak | Ready to use |  | Imam
4000058 | 2020-07-29 | RSUD Merah Putih | X-Ray Analog | Apelem/Camargue Analog | IMA703979 | Service | Rotor tidak bisa berfungsi | Ready to use |  | Bayu
6000237 | 2020-07-29 | RSUD Panembahan Senopati Bantul | Mesin Anasthesia | Draeger/Perseus A500 | ASND0096 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-07-29
3000146 | 2020-07-28 | RSUD Tidar Magelang | Bone Mineral Densitometer | Medilink/Medix DR | M15016M258 | Service | Scanner error | Ready To Use |  | Imam
15000129 | 2020-07-28 | RS. Dr. ERNALDI BAHAR - PALEMBANG | X-Ray Analog | Apelem/Camargue Analog | CPD30672C18 | Service | Tidak bisa ANG | Ready to use |  | Sutan
19000239 | 2020-07-28 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSNC0131 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-07-28
19000239 | 2020-07-28 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSNE2737 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-07-28
13000069 | 2020-07-28 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Service | pegangan pribe usg rusak | Ready to use |  | Sofian
11000183 | 2020-07-28 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 09111371 | Service | board adaptor dan software corrupt // ganti board | Ready To Use | Alat sudah discontinue. Alat digudangkan | Nasrul
16000125 | 2020-07-27 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1086410 | Service | Penggantian valve house yang pecah | Ready To Use |  | Ipan
6000236 | 2020-07-27 | RSUD Kab. Temanggung | Resuscitaire | Draeger/Resuscitaire | JY10904 | Service | Error 10 | Ready to use | Disarankan penggantian control board. Sudah diperbaiki (6000274) | Afrizal
14000117 | 2020-07-27 | RSUD Kanjuruhan | UPS | Fujielectric/UPS7000HX-T4 | PWS00056545 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2021-07-27
2000129 | 2020-07-24 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Service | Sofware rusak | Ready to use |  | Eddy
8000231 | 2020-07-23 | RSI Sunan Kudus | Ventilator | Draeger/Savina 300 | ARUE0022 | Service | alarm internal battery | Ready to use | Sudah diberikan penawaran namun tidak PO. Karna membeli alat baru | Amna
8000231 | 2020-07-23 | RSI Sunan Kudus | Ventilator | Draeger/Savina 300 | ARUM0308 | Service | alarm internal battery | Ready to use | Sudah diberikan penawaran namun tidak PO. Karna membeli alat baru | Amna
19000238 | 2020-07-23 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSNE0131 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-07-23
19000238 | 2020-07-23 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSNE2727 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-07-23
19000238 | 2020-07-23 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSNE2728 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-07-23
19000238 | 2020-07-23 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSNE2729 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-07-23
19000238 | 2020-07-23 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSNE2731 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-07-23
19000238 | 2020-07-23 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VGSNE2732 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-07-23
18000139 | 2020-07-23 | RSUP Ratatotok Buyat | General X-Ray | Apelem/Varian/Diamond | 215918QH21591 | Service | Pemasangan lampu kolimator | Ready to use |  | Sukri
11000182 | 2020-07-22 | RSUD Kab. Muna | Medical Compressor | Draeger/Medical Compressor | ASNA0025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-07-22
11000181 | 2020-07-22 | RSUD Kab. Muna | Resuscitaire | Draeger/Resuscitaire | BV14105 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-07-22
9000131 | 2020-07-21 | RSUD Weda | Medical cart | Paramount Bed | KY40QR0420C | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Irfan R
7000218 | 2020-07-21 | RSKBD Panti Nugroho | Pendant | Draeger/Movita | ASLL0154 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-07-21
7000218 | 2020-07-21 | RSKBD Panti Nugroho | Pendant | Draeger/Movita | ASMA0061 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-07-21
7000218 | 2020-07-21 | RSKBD Panti Nugroho | Pendant | Draeger/Movita | ASMK0349 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-07-21
15000128 | 2020-07-21 | RSUD Dr. M. Yunus | ESWL | Medispec/EM1000 | 0917011 | Service | Tembakan tidak normal | Ready to use |  | Sutan
10000103 | 2020-07-21 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0291 | Service | Pemasangan detector wireless pada ceilling | Ready to use |  | Rizaldi
7000218 | 2020-07-21 | RSKBD Panti Nugroho | Pendant | Draeger/Movita | ASMK0350 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-07-21
19000237 | 2020-07-21 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002705 | Service | General check | Ready to use |  | Danu
19000236 | 2020-07-20 | Klinik Drg. Lenny | Kompressor dental unit | Durr/Tornado 1 | U004191 | Survey | Kompressor sering mati | Ready to use |  | Danu
19000235 | 2020-07-20 | RSUD Dr. Soediran Mangun Sumarso | Resuscitaire | Draeger/Resuscitaire | KC13737 | Survey | Training User | Ready to use |  | Danu
9000128 | 2020-07-20 | RSUD Weda | Bed Obgyn | paramount bed | 00008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
6000253 | 2020-07-20 | RSUD KOTA YOGYAKARTA | Mesin Anasthesia | Draeger/Primus | ASLD0030 | Maintenance | Service kontrak total care | Ready to use |  | Afrizal
9000128 | 2020-07-20 | RSUD Weda | Bed Obgyn | paramount bed | 00009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000128 | 2020-07-20 | RSUD Weda | Bed Obgyn | paramount bed | 00010 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Irfan R
9000132 | 2020-07-20 | RSUD Weda | Bed Patient | Paramount Bed | PA992852720F | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000133 | 2020-07-20 | RSUD Weda | Bed Patient | Paramount Bed | PA992902720F | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
9000130 | 2020-07-20 | RSUD Weda | Emergency Stretcher | Paramount/PK 2100 | PK8031137020F | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
11000180 | 2020-07-20 | RSUD Kab. Bombana | Mesin Anasthesia | Draeger/Primus | ASMH0119 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-07-20
13000065 | 2020-07-19 | RSUP Dr. H. Adam Malik Medan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/033/C0 | Service | Upgrade software | Ready to use |  | Sofian
18000140 | 2020-07-18 | RSUD Toto Kabila | Lampu Operasi | Draeger/Polaris 600 | ASNC1138 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-07-18
18000140 | 2020-07-18 | RSUD Toto Kabila | Lampu Operasi | Draeger/Polaris 600 | ASNC1139 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-07-18
9000127 | 2020-07-18 | RSD Tidore | Bed Patient | Paramount Bed | 11900122 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
6000256 | 2020-07-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0011 | Maintenance | Kontrak Service | Ready To Use |  | Afrizal
3000145 | 2020-07-16 | RSUD Kab Tangerang | C-Arm | Apelem/Evo R+ | 5511161245 | Service | Unit mati | Ready to use | Alat sudah diperbaiki sendiri oleh pihak RS | Imam
4000057 | 2020-07-16 | RSUD Balaraja | C-Arm | Apelem/Evo R+ | 55118034657 | Survey | Pelaksanaan Instalasi dan uji paparan | Ready to use |  | Bayu
6000249 | 2020-07-16 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | VTSKF0012 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Afrizal | 2021-07-16
17000104 | 2020-07-16 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Survey | Pergerakan meja untuk kemiringan meja tidak berfungsi | Ready to use |  | Bahar
15000127 | 2020-07-16 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Service | PU BOX tidak nyala, tube ceiling bocor | Ready to use | Alat masih bisa digunakan | Sutan
6000249 | 2020-07-16 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | VTSKF0014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-07-16
18000138 | 2020-07-16 | RSUD Datoe Binangkang | Ventilator | Draeger/Savina 300 | ASMK0302 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-07-16
9000126 | 2020-07-16 | RSUP Ratatotok Buyat | X-Ray Analog | Apelem/Camargue Analog | CPD36816520 | Service | Lampu kolimator putus | Ready to use |  | Irfan R
7000217 | 2020-07-16 | RSUI Harapan Anda | Dental Unit | Kavo/1033B | 169 | Service | suction motor mati total | Ready to use |  | Ariawan
8000230 | 2020-07-15 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASKF0024 | Maintenance | General check | Ready to use |  | Amna
8000230 | 2020-07-15 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASKF0025 | Maintenance | General check | Ready to use |  | Amna
19000233 | 2020-07-15 | RSUD KOTA SURAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0112 | Maintenance | general check | Ready To Use |  | Danu
9000121 | 2020-07-15 | RSUD Maria Walanda Maramis | Mesin Anasthesia | Draeger/Primus | ASMH0120 | Maintenance | general check | Ready to use |  | Irfan R
9000121 | 2020-07-15 | RSUD Maria Walanda Maramis | Mesin Anasthesia | Draeger/Primus | ASMH0145 | Service | general check | Ready to use |  | Irfan R
6000258 | 2020-07-14 | Klinik Graha Amanah | Jaundice Meter | Draeger/JM 105 | B3401082 | Service | Baterai diganti | Ready to use |  | Afrizal
7000216 | 2020-07-14 | RSKBD Panti Nugroho | Resuscitaire | Draeger/Resuscitaire | MU20503 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-07-14
6000260 | 2020-07-14 | RSUD Bagas Waras Klaten | Resuscitaire | Draeger/Resuscitaire | JY03386 | Service | Error 05 | Ready to use |  | Afrizal
6000262 | 2020-07-14 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASJN0023 | Service | Pemasangan membran complete, kit air filter, parficle filter, filter pad | Ready to use |  | Afrizal
6000262 | 2020-07-14 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0091 | Service | Pemasangan membran complete, kit air filter, parficle filter, filter pad | Ready to use |  | Afrizal
6000261 | 2020-07-14 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASLL0067 | Service | error expiratory valve | Ready to use | Sudah dilakukan perbaikan (6000205) | Afrizal
6000262 | 2020-07-14 | RSUP dr. Soeradji Tirtonegoro | Ventilator | Draeger/Babylog VN500 | ASLL0067 | Service | Pemasangan membran complete, kit air filter, parficle filter, filter pad | Ready to use |  | Afrizal
11000179 | 2020-07-14 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASEF0068 | Service | lampu alarm dan monitor mati | Ready to use |  | Nasrul
19000232 | 2020-07-13 | Siloam Hospital Semarang | Incubator Transport | Draeger/Isolette TI500 | CR12070 | Service | Baterai error | Ready to use |  | Danu
19000232 | 2020-07-13 | Siloam Hospital Semarang | Baby Incubator | Draeger/Isolette C2000 | CR35565 | Service | Error oxygen fail | Ready to use |  | Danu
6000263 | 2020-07-12 | RSPAU Harjolukito | Baby Incubator | Draeger/Isolette 8000 | ASAC0017 | Service | Pemasangan Service set 2 years dan breathing system | Ready to use |  | Afrizal
6000263 | 2020-07-12 | RSPAU Harjolukito | Baby Incubator | Draeger/Isolette 8000 | ASLD0040 | Service | Pemasangan Service set 2 years dan breathing system | Ready to use |  | Afrizal
9000123 | 2020-07-11 | RSUD KOTAMOBAGU | Patient Monitor | Draeger/Infinity Delta XL | 6009707463 | Survey | Error pada NIBP | Ready to use | Alat sudah diperbaiki dan dikembalikan ke rs | Sukri
3000144 | 2020-07-09 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helinthus BYM | HELI/0023/CB0 | Service | Komputer work station error os window | Ready to use |  | Imam
19000231 | 2020-07-09 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0129 | Service | Training User | Ready to use |  | Bayu
4000052 | 2020-07-09 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Survey | Hasil kualitas citra tidak lolos uji | Ready to use |  | Bayu
18000135 | 2020-07-09 | UPTD Puskesmas Sangkub | Stretcher | Paramount bed/PK5114X | 900200700033 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000132 | 2020-07-09 | UPTD Puskesmas Mokoditek | Stretcher | Paramount bed/PK5114X | 900107100169 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000134 | 2020-07-09 | UPTD Puskesmas Bintauna | Stretcher | Paramount bed/PK5114X | 90010030075 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000133 | 2020-07-09 | UPTD Puskesmas Bahabak | Stretcher | Paramount bed/PK5114X | 900200700032 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
4000055 | 2020-07-08 | RSUD Tidar Magelang | Mobile X-ray | Siemens Mobillet Mira Max | 212337 | Maintenance | Pelaksanaan Uji Kesesuaian | Ready to use |  | Bayu
4000053 | 2020-07-08 | RSUD Tidar Magelang | X-Ray Analog | Shimadzu Ezy Rad Pro | 116154202 | Survey | Pelaksanaan Uji Kesesuaian | Ready to use |  | Bayu
4000054 | 2020-07-08 | RSUD Tidar Magelang | General X-Ray | GE Hualun Medical System | 9209HL6 | Maintenance | Pelaksanaan Uji Kesesuaian | Ready to use |  | Bayu
4000056 | 2020-07-08 | RSUD Tidar Magelang | Mobile X-ray | Hitachi Sirius 130 HP | SX10220612 | Survey | Pelaksanaan Uji Kesesuaian | Ready to use |  | Bayu
18000129 | 2020-07-08 | UPTD Puskesmas Ollot | Stretcher | Paramount bed/PK5114X | 900100300066 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000127 | 2020-07-08 | UPTD Puskesmas Buko | Stretcher | Paramount bed/PK5114X | 9.001e+22 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000131 | 2020-07-08 | UPTD Puskesmas Boroko | Stretcher | Paramount bed/PK5114X | 901201200003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000130 | 2020-07-08 | UPTD Puskesmas Bolongitan | Stretcher | Paramount bed/PK5114X | 90010030007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
19000230 | 2020-07-07 | RS Telogorejo | Ventilator | Draeger/Carina | ASLM0055 | Survey | Alarm baterai indicator tidak berfungsi | Ready to use |  | Danu
9000122 | 2020-07-07 | RSUD KOTAMOBAGU | C-Arm | Apelem/Evo R+ | 55118031654 | Service | Pengecekan data SN tube | Ready to use |  | Irfan R
19000227 | 2020-07-07 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Maintenance | Sensor O2 fail | Ready to use | Sudah dibuatkan penawaran | Danu
19000228 | 2020-07-04 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Service | Maintenance mesin anasthesi | Ready to use |  | Danu
9000116 | 2020-07-04 | RSUP Prof. Dr. R. D. Kandou Manado | Mobile X-ray | Apelem/Rafale EV30 | T8621 | Service | Kerusakan pada board starter | Ready to use |  | Irfan R
11000178 | 2020-07-03 | Politeknik Kesehatan Mamuju | Jaundice Meter | Draeger/JM 105 | H3500449 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-07-03
13000066 | 2020-07-03 | RSUP Dr. H. Adam Malik Medan | Bone Mineral Densitometer | Medilink/Medix DR | J17016M280 | Service | generator fault | Ready To Use |  | Sofian
19000227 | 2020-07-03 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Service | Sensor O2 fail | Ready To Use | Sudah diperbaiki (19000227) | Danu
16000123 | 2020-07-02 | RS Mayapada | Dental Unit | Kavo/E30TM | 10001757 | Service | Selang suction (besar) robek | Ready to use |  | Ipan
9000121 | 2020-07-02 | RSUD Anugrah Tomohon | Bed Patient | paramount bed | 900312700177 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R
2000127 | 2020-07-02 | RSUD Kab. Temanggung | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/011/C0 | Service | Tidak bisa print | Ready To Use |  | Eddy
15000126 | 2020-07-02 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 06111187 | Service | Penggantian track ball USG | Ready to use | Sudah dilakukan perbaikan | Sutan
3000143 | 2020-07-01 | RSUD Sekayu | X-ray Digital | Apelem/Camargue DR | IMA703957 | Service | Detector tidak konek ke workstation | Ready To Use |  | Imam
6000265 | 2020-07-01 | RSPAU Harjolukito | Ventilator | Draeger/Savina 300 | ASDL0128 | Survey | no internal baterai | Ready to use |  | Afrizal
11000177 | 2020-07-01 | RSUD Anutapura Palu | C-Arm | Apelem/Evo R+ | 55120019805 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2022-07-01
2000128 | 2020-07-01 | RSUD Kab. Temanggung | ESWL | Medispec/EM1000 | 1115056 | Service | membran tidak mengembang | Ready To Use |  | Eddy
6000245 | 2020-07-01 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Survey | Self-test tidak selesai | Ready to use |  | Afrizal
19000226 | 2020-07-01 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002719 | Survey | Bocor | Ready to use |  | danu
18000126 | 2020-06-30 | RSUD Toto Kabila | Medical Compressor | Draeger/Medical Compressor | ASNA0021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-06-30
18000126 | 2020-06-30 | RSUD Toto Kabila | Medical Compressor | Draeger/Medical Compressor | ASNA0022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-06-30
18000126 | 2020-06-30 | RSUD Toto Kabila | Medical Compressor | Draeger/Medical Compressor | ASNA0024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-06-30
18000124 | 2020-06-30 | RSUD Toto Kabila | Ventilator | Draeger/Evita V300 | ASND0131 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-06-30
18000124 | 2020-06-30 | RSUD Toto Kabila | Ventilator | Draeger/Evita V300 | ASND0133 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-06-30
19000225 | 2020-06-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASJC0044 | Maintenance | Pengecekan unit, check visual | Ready to use |  | Danu
19000225 | 2020-06-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASJM0044 | Maintenance | Pengecekan unit, check visual | Ready to use |  | Danu
19000225 | 2020-06-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASLN0008 | Maintenance | Pengecekan unit, check visual | Ready to use |  | Danu
19000225 | 2020-06-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0079 | Maintenance | Pengecekan unit, check visual | Ready to use |  | Danu
19000225 | 2020-06-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0085 | Maintenance | Pengecekan unit, check visual | Ready to use |  | Danu
19000225 | 2020-06-30 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0094 | Survey | Pengecekan unit, check visual | Ready to use |  | Danu
16000122 | 2020-06-29 | Klinik dr. J. Sintha T | Dental Unit | Kavo 1042 | 7050020 | Service | Separator tidak menyala | Ready to use |  | Ipan
11000175 | 2020-06-29 | RSUD Anutapura Palu | Phototherapy | Draeger/Bililux | 1101999014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-06-29
11000175 | 2020-06-29 | RSUD Anutapura Palu | Phototherapy | Draeger/Bililux | 1101999019 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-06-29
11000173 | 2020-06-29 | RSUD Anutapura Palu | Ventilator | Draeger/Babylog VN500 | ASML0098 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-06-29
11000174 | 2020-06-29 | RSUD Anutapura Palu | Jaundice Meter | Draeger/JM 105 | H3500328 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-06-29
6000266 | 2020-06-29 | RSUD Kab. Temanggung | Resuscitaire | Draeger/Resuscitaire | JY10904 | Service | cek error 10 | Ready to use | Sudah diperbaiki (6000274) | Afrizal
9000119 | 2020-06-29 | RSUP Prof. Dr. R. D. Kandou Manado | ESMR | Medispec/ED1000 | 7111298 | Service | Mengganti aplicator dan perbaikan saluran air | Ready to use |  | Irfan R
3000142 | 2020-06-27 | RSUD Kab. Temanggung | C-Arm | Apelem/Evo R+ | 55115011409 | Service | Objek x ray mengecil | Ready to use |  | Imam
9000115 | 2020-06-27 | RSUD Datoe Binangkang | X-ray Digital | Apelem/Camargue DR | CPD33569E19 | Service | Kolimator goyang karena bautnya lepas | Ready to use |  | Irfan R
3000140 | 2020-06-26 | RSUD Kab. Temanggung | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/011/C0 | Maintenance | copy master OS Workstation | Ready to use |  | imam
3000141 | 2020-06-26 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Service | setting printer fuji drypix smart | Ready to use |  | Imam
4000140 | 2020-06-26 | RSUD Tualang | X-ray Digital | Apelem/Camargue DR | CPD30502813 | Maintenance | Bukaan kolimasi kurang lebar, mAs error | Ready to use |  | Bayu
4000140 | 2020-06-26 | RSUD Tualang | X-ray Digital | Apelem/Camargue DR | CPD30502813 | Service | Pelaksanaan uji kesesuaian dan uji paparan | Ready to use |  | Bayu
19000224 | 2020-06-25 | RS Ken Saras | Mesin Anasthesia | Draeger/Fabius GS Premium | ASBT0054 | Service | frsh gas tidak lolos kalibrasi dan alarm O2 | Ready to use |  | Danu
6000269 | 2020-06-25 | RSA UGM | Mesin Anasthesia | Draeger/Fabius Plus | ASEN0135 | Survey | ventilator leak failed | Ready to use |  | Afrizal
6000268 | 2020-06-25 | RSUP Dr. Sardjito | Ventilator | Draeger/Babylog VN500 | ASJE0010 | Survey | expiratory failed, pressure sensor kalibrasi valve failed | Ready to use | Alat sudah diperbaiki | Afrizal
6000268 | 2020-06-25 | RSUP Dr. Sardjito | Jaundice Meter | Draeger/JM 105 | B3502011 | Service | waktu kalibrasi | Ready to use | Alat sudah dikalibrasi | Afrizal
16000121 | 2020-06-24 | Family Dentistry | Dental Unit | Kavo/E30TM | 10000655 | Service | Foot control error | Ready to use |  | Ipan
16000120 | 2020-06-24 | Klinik drg. Mirna Hakim | Dental Unit | Kavo | 10064564 | Service | Penggantian bearing handpiece | Ready to use |  | Ipan
8000228 | 2020-06-24 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Service | O2 sensor fail | Ready to use |  | Amna
18000136 | 2020-06-24 | RSUD Noongan | Ventilator | Draeger/Evita V300 | ASND0132 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-06-24
18000136 | 2020-06-24 | RSUD Noongan | Ventilator | Draeger/Evita V300 | ASND0134 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-06-24
18000136 | 2020-06-24 | RSUD Noongan | Ventilator | Draeger/Evita V300 | ASND0137 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan R | 2021-06-24
19000223 | 2020-06-23 | RS Indriati Solo | Mesin Anasthesia | Draeger/Primus | ASJF0036 | Service | cleaning unit, block cozy, self test unit dan cek sistem | Ready to use |  | Danu
6000270 | 2020-06-22 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Service | Filter perlu dibersihkan | Ready To Use |  | Afrizal
8000227 | 2020-06-22 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASNB0020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-06-22
9000114 | 2020-06-22 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 0613710 | Service | Kerusakan pada adaptor | Ready to use |  | Irfan
11000176 | 2020-06-20 | RSU MADANI - PALU | Baby Incubator | Draeger/Isolette 8000 | EC09177 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-06-20
11000176 | 2020-06-20 | RSU MADANI - PALU | Baby Incubator | Draeger/Isolette 8000 | EC09178 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-06-20
11000176 | 2020-06-20 | RSU MADANI - PALU | Baby Incubator | Draeger/Isolette 8000 | HC09151 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-06-20
9000109 | 2020-06-19 | RSUD Noongan | Lampu Operasi | Draeger/Polaris 600 | ASNB1017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
8000226 | 2020-06-18 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHD0188 | Service | error alarm system cehck gas supply low | Ready to use |  | Amna
5000130 | 2020-06-18 | RSUD Dr.(H.C) Ir. Soekarno | ESMR | Medispec/ED1000 | 320034 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal | 2022-06-18
18000122 | 2020-06-18 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Babylog VN500 | ASMF0009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | sukri | 2021-06-18
18000122 | 2020-06-18 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Babylog VN500 | ASMF0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | sukri | 2021-06-18
18000122 | 2020-06-18 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Babylog VN500 | ASML0101 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | sukri | 2021-06-18
18000122 | 2020-06-18 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Babylog VN500 | ASNA0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | sukri | 2021-06-18
19000222 | 2020-06-18 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette 8000 | CF06001 | Maintenance | General check | Ready to use |  | Danu
19000222 | 2020-06-18 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette 8000 | CF06005 | Maintenance | General check | Ready to use |  | Danu
19000222 | 2020-06-18 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette 8000 | HC09152 | Service | General check | Ready to use |  | Danu
15000125 | 2020-06-18 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Survey | Shock habisa, baterai rusak | Ready to use |  | Sutan
9000112 | 2020-06-18 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 0613710 | Maintenance | Adaptor dan fuse rusak | Ready to use |  | Irfan R
8000224 | 2020-06-17 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | Maintenance | General check, system check, running test | Ready to use |  | Amna
8000224 | 2020-06-17 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Maintenance | General check, system check, running test | Ready to use |  | Amna
18000121 | 2020-06-17 | RSUD dr. Zainal Umar Sidiki | Bed Patient | Paramount | 900505100176 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
14000116 | 2020-06-17 | RSUD Bangil | UPS | Yoshiga/GPH 33-120 | 501201140080L4500001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2021-06-17
19000221 | 2020-06-17 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0094 | Service | Alarm system failed dan battery replacement | Ready to use |  | Danu
8000225 | 2020-06-17 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0039 | Service | System check, sensor check fail | Ready to use |  | Amna
6000252 | 2020-06-17 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Service | tidak bisa nembak | Ready to use |  | Afrizal
16000119 | 2020-06-16 | oe | Dental Unit | Kavo/1058 TM | 10010476 | Service | Selang water tank pecah | Ready to use |  | Ipan
13000065 | 2020-06-16 | RSU Haji Medan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/033/C0 | Service | Upgrade software | Ready to use |  | Sofian
8000223 | 2020-06-16 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0011 | Service | power fail | Ready to use |  | Amna
8000223 | 2020-06-16 | RSUD Kraton Pekalongan | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASHN0009 | Survey | Alarm paw error | Ready to use |  | Amna
3000139 | 2020-06-15 | RSUD Harapan dan Doa | X-ray Digital | Apelem/Camargue DR | CPD33564E19 | Service | Posisi tube miring dan kolimator; Perbaikan posisi tube dan perapihan kabel | Ready to use |  | Imam
19000220 | 2020-06-15 | RS Triharsi | Ventilator | Draeger/Evita V300 | ASJH0064 | Service | Training User | Ready to use |  | Danu
19000219 | 2020-06-15 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Caleo | ASKL0017 | Service | Error water empty | Ready to use |  | Danu
7000213 | 2020-06-15 | RSUD DR. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASMH0115 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Ariawan | 2021-06-15
7000213 | 2020-06-15 | RSUD DR. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASNB0071 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Ariawan | 2021-06-15
4000051 | 2020-06-15 | RSUD Harapan dan Doa | X-ray Digital | Apelem/Camargue DR | CPD33564E19 | Survey | Uji Kesesuaian dan Paparan | Ready to use |  | Bayu
8000222 | 2020-06-15 | RSUD LIMPUNG | Baby Incubator | Draeger/Isolette C2000 | AC37687 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-06-15
8000222 | 2020-06-15 | RSUD LIMPUNG | Baby Incubator | Draeger/Isolette C2000 | GV38144 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-06-15
8000222 | 2020-06-15 | RSUD LIMPUNG | Infant Warmer | Draeger/Vista 120 | VGSNC0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-06-15
8000222 | 2020-06-15 | RSUD LIMPUNG | Infant Warmer | Draeger/Vista 120 | VGSNC0254 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-06-15
8000222 | 2020-06-15 | RSUD LIMPUNG | Infant Warmer | Draeger/Vista 120 | VTSLK0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-06-15
3000138 | 2020-06-13 | RSUD Alimuddin Umar | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0402 | Service | Power supply error | Ready to use |  | Imam
4000050 | 2020-06-13 | RSUD Alimuddin Umar | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0402 | Survey | Pelaksanaan uji kesesuaian dan uji paparan | Ready to use |  | Bayu
9000109 | 2020-06-13 | RSUD Noongan | Lampu Operasi | Draeger/Polaris 600 | ASNB1017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-06-13
11000172 | 2020-06-13 | RSUD Lakipadada | Mesin Anasthesia | Draeger/Primus | ASMJ0188 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-06-13
8000229 | 2020-06-13 | RSUD Kayen Pati | Patient Monitor | Draeger/Delta XL | 6009956568 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-06-13
8000229 | 2020-06-13 | RSUD Kayen Pati | Resuscitaire | Draeger/Resuscitaire | BV14106 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-06-13
8000229 | 2020-06-13 | RSUD Kayen Pati | Resuscitaire | Draeger/Resuscitaire | BV14107 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-06-13
8000229 | 2020-06-13 | RSUD Kayen Pati | Resuscitaire | Draeger/Resuscitaire | GV14025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-06-13
17000103 | 2020-06-11 | RSUD Kota Salatiga | ESWL | Medispec/E3000 | 09131174 | Maintenance | HVG rusak | Ready To Use |  | Bahar
5000129 | 2020-06-10 | RS Siloam Asri | ESMR | Medispec/ED1000 | 0611274 | Service | Pengisian air applikator | Ready to use |  | Bilal
18000120 | 2020-06-10 | RSUD Anugrah Tomohon | Bed Patient | Paramount Bed | 9000307300024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
19000216 | 2020-06-10 | RSUP Dr. Kariadi | Resuscitaire | Draeger/Resuscitaire | HC13576 | Maintenance | General check | Ready to use |  | Danu
19000215 | 2020-06-10 | RSUP Dr. Kariadi | Resuscitaire | Draeger/Resuscitaire | KC13734 | Maintenance | General check | Ready to use |  | Danu
19000214 | 2020-06-10 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0300 | Maintenance | General check | Ready to use |  | Danu
19000214 | 2020-06-10 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0301 | Maintenance | General check | Ready to use |  | Danu
19000214 | 2020-06-10 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0303 | Maintenance | General check | Ready to use |  | Danu
19000214 | 2020-06-10 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0305 | Maintenance | General check | Ready to use |  | Danu
19000214 | 2020-06-10 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0322 | Service | General check | Ready to use |  | Danu
11000171 | 2020-06-10 | RSUP Dr. Wahidin Sudirohusodo | Patient Monitor | Draeger/Vista 120 | V3SHK0011 | Service | Power board rusak | Ready to use |  | Nasrul
16000117 | 2020-06-09 | Klinik drg. Shinta | Dental Unit | Kavo werk | 7140170 | Service | Mengganti selang suction yang sudah pendek | Ready to use |  | Ipan
16000116 | 2020-06-08 | Stephanie dental clinic | Dental Unit | Kavo/1042 | 7050020 | Service | Main board meletup | Ready to use |  | Ivan
16000115 | 2020-06-07 | Klinik Budihardja | Dental Unit | Kavo/1058 TM | 10010476 | Maintenance | Low speed tidak mengeluarkan air | Ready to use |  | Ivan
19000210 | 2020-06-06 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0065 | Maintenance | General check | Ready to use |  | Danu
19000210 | 2020-06-06 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Maintenance | General check | Ready to use |  | Danu
11000170 | 2020-06-05 | RSUD Lasingran | Mesin Anasthesia | Draeger/Primus | ASMH0116 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-06-05
5000128 | 2020-06-04 | Dr. Mustofa | ESMR | Medispec/ED1000 | 1114007 | Service | Aplikator kekurangan air | Ready to use |  | Bilal
8000221 | 2020-06-04 | RSUD RA Kartini | Lampu Operasi | Draeger/Polaris 600 | ASMA0361 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
110000168 | 2020-06-04 | RSUD LA TEMMAMALA | Ventilator | Draeger/Savina 300 | ASMA0016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-06-04
110000168 | 2020-06-04 | RSUD LA TEMMAMALA | Ventilator | Draeger/Babylog VN500 | ASML0095 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-06-04
16000114 | 2020-06-03 | Klinik Suryatenggara | Dental Unit | Kavo/1058TM | 10010472 | Service | Lampu unit mati | Ready to use |  | Ivan
11000167 | 2020-05-30 | RSUD H. Andi Sulthan Daeng Radja | Ventilator | Draeger/Babylog VN500 | ASMK0030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-05-30
15000124 | 2020-05-30 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Monitor, table dan joy stick tidak nyala | Ready to use |  | Sutan
9000107 | 2020-05-29 | RSUD Ainun Habibie | Bed Patient | Paramount Bed/KY40QR | 900502100026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
9000108 | 2020-05-29 | RSUD Ainun Habibie | Bed Patient | Paramount Bed/DA6315DC5003 | 9.0041e+22 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
18000117 | 2020-05-29 | RSUD Ainun Habibie | Bed Patient | Paramount Bed/PIA11210 | 1220E900504200002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000117 | 2020-05-29 | RSUD Ainun Habibie | Bed Patient | Paramount Bed/PIA11210 | 1220E900504200004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000117 | 2020-05-29 | RSUD Ainun Habibie | Bed Patient | Paramount Bed/PIA11210 | 1220E900542000063 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000118 | 2020-05-29 | RSUD Ainun Habibie | Bed Patient | Paramount Bed/PF1100 | PF1100A | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000118 | 2020-05-29 | RSUD Ainun Habibie | Bed Patient | Paramount Bed/KF282 | Y0020D91040000373 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000118 | 2020-05-29 | RSUD Ainun Habibie | Bed Patient | Paramount Bed/KF282 | Y0020D910400100673 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
7000211 | 2020-05-29 | RSUD Hj. Anna Lasmanah | Baby Incubator | Draeger/Caleo | ASJH0021 | Maintenance | check humadity dan bersihkan unit | Ready to use |  | Ariawan
19000209 | 2020-05-29 | RSUD KOTA SURAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0112 | Maintenance | Pemasangan konektor gas O2 dan air | Ready to use |  | Danu
19000208 | 2020-05-28 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0300 | Maintenance | Test fungsional unit, test leaked, kalibrasi O2 dan flow sensor | Ready to use |  | Danu
19000208 | 2020-05-28 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0301 | Maintenance | Test fungsional unit, test leaked, kalibrasi O2 dan flow sensor | Ready to use |  | Danu
19000208 | 2020-05-28 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0303 | Maintenance | Test fungsional unit, test leaked, kalibrasi O2 dan flow sensor | Ready to use |  | Danu
19000208 | 2020-05-28 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0305 | Maintenance | Test fungsional unit, test leaked, kalibrasi O2 dan flow sensor | Ready to use |  | Danu
19000208 | 2020-05-28 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0322 | Service | Test fungsional unit, test leaked, kalibrasi O2 dan flow sensor | Ready to use |  | Danu
10000102 | 2020-05-27 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | blue screen pada komputer | Ready to use |  | Rizaldi
18000193 | 2020-05-27 | RSUP Ratatotok Buyat | X-ray Digital | Apelem/Camargue DR | CPD32474K18 | Maintenance | Lampu kolimator putus | Ready To Use |  | Irfan R
5000127 | 2020-05-26 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Service | Pengecekan aplicator ED 1000 karena kekurangan air | Ready to use |  | Bilal
15000122 | 2020-05-24 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Joy stick platinum tidak nyala | Ready to use |  | Sutan
19000206 | 2020-05-22 | RSUP Dr. Kariadi | Jaundice Meter | Draeger/JM 105 | 3502043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu
19000207 | 2020-05-22 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASML0087 | Service | Cleaning Filter | Ready to use |  | Danu
16000113 | 2020-05-22 | Klinik Lanud Haluoleo | Mobile X-ray | Apelem/Rafale EV30 | 1802236 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ivan | 2022-05-22
19000207 | 2020-05-22 | RSUP Dr. Kariadi | Ventilator | Draeger/Savina 300 | ASMK0304 | Survey | Cleaning Filter | Ready to use |  | Danu
4000049 | 2020-05-21 | RSAU Lanud Eltari Kupang | Mobile X-ray | Apelem/Rafale EV30 | 005/18/01384 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2022-05-21
3000137 | 2020-05-20 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helinthus BYM | HELI/0023/CB0 | Maintenance | Setting PACS | Ready to use |  | Imam
6000143 | 2020-05-20 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Service | Cek software,hadrware, dan pengecekan regulator O2 | Ready to use |  | Afrizal
6000144 | 2020-05-20 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | V2SHK0006 | Service | No module activated | Ready to use |  | Afrizal
17000102 | 2020-05-18 | RSUD Alimuddin Umar | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0402 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bahar | 2022-05-18
8000218 | 2020-05-15 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASKF0024 | Service | General check, system check, running test | Ready to use |  | Amna
19000204 | 2020-05-15 | RSUP Dr. Kariadi | Baby Incubator | Draeger/Caleo | ASJM0048 | Maintenance | Alarm water empty | Ready to use |  | Danu
6000142 | 2020-05-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASJE0009 | Service | Cek software,hadrware dan ganti breathing circuit | Ready to use |  | Afrizal
19000203 | 2020-05-14 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Primus | ASNC0129 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu
19000202 | 2020-05-13 | RSUD KOTA SURAKARTA | Mesin Anasthesia | Draeger/Primus | ASMF0112 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Danu | 2021-05-13
18000116 | 2020-05-13 | RSUD Kab. Bolaang Mongondow | Mesin Anasthesia | Draeger/Primus | ASMH0117 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-05-13
6000141 | 2020-05-13 | RSUD KOTA YOGYAKARTA | Charger | Draeger | 6654 | Service | Tidak bisa kalibrasi internal | Ready to use |  | Afrizal
6000141 | 2020-05-13 | RSUD KOTA YOGYAKARTA | Jaundice Meter | Draeger/JM 105 | B3503376 | Service | Tidak bisa kalibrasi internal | Ready to use |  | Afrizal
6000141 | 2020-05-13 | RSUD KOTA YOGYAKARTA | Docking station | Draeger | B7038224 | Survey | Tidak bisa kalibrasi internal | Ready to use |  | Afrizal
6000139 | 2020-05-11 | RSUD Sleman | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0128 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-05-11
19000201 | 2020-05-11 | RSUD Dr. Soediran Mangun Sumarso | Mesin Anasthesia | Draeger/Perseus A500 | ASNC0129 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Danu | 2021-05-11
18000115 | 2020-05-11 | RSUD Dr. M. M. Dunda | Lampu Operasi | Draeger/Polaris 600 | ASMA0358 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-05-11
15000121 | 2020-05-11 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Service | Error filament fault | Ready to use |  | Sutan
6000137 | 2020-05-09 | RSUD Prambanan | Jaundice Meter | Draeger/JM 105 | H3500432 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-05-09
8000219 | 2020-05-06 | RSJD Dr. Amino Gondohutomo | Jaundice Meter | Draeger/JM 105 | HC3500431 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-05-06
9000249 | 2020-05-04 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Fabius Plus | ASKK0057 | Service | Selang O2 putus | Ready To Use |  | Irfan
18000162 | 2020-04-30 | RSUD Toto Kabila | Central Monitor | Draeger/Vista 120 CMS | VLSNF0251 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-30
8000217 | 2020-04-30 | RSUD RA Kartini | Ventilator | Draeger/Babylog VN500 | ASMK0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-30
8000217 | 2020-04-30 | RSUD RA Kartini | Incubator Transport | Draeger/Isolette TI500 | ASMK0029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-30
8000217 | 2020-04-30 | RSUD RA Kartini | Lampu Operasi | Draeger/Polaris 600 | HC13276 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-30
7000210 | 2020-04-30 | RSKBD Panti Nugroho | Mesin Anasthesia | Draeger/Primus | ASMF0111 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-04-30
6000134 | 2020-04-30 | RSUP Dr. Sardjito | Incubator Transport | Draeger/GT-500 IH | CY02409 | Service | Baterai error | Ready to use |  | Afrizal
4000146 | 2020-04-29 | RSUD Kab. Bekasi | C-Arm | Apelem/Evo R+ | 55118026649 | Service | Printer tidak berfungsi, break longitudinal loss | Ready to use |  | Bayu
9000212 | 2020-04-29 | RSUD MARIA WALANDA MARAMIS | Mesin Anasthesia | Draeger/Primus | ASMH0120 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-29
9000212 | 2020-04-29 | RSUD MARIA WALANDA MARAMIS | Mesin Anasthesia | Draeger/Primus | ASMH0145 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-29
18000110 | 2020-04-29 | RSUD Toto Kabila | Lampu Operasi | Draeger/Polaris 600 | ASML1180 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-29
8000216 | 2020-04-28 | RSUD RAA Soewondo Pati | Resuscitaire | Draeger/Resuscitaire | AC13489 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-28
8000216 | 2020-04-28 | RSUD RAA Soewondo Pati | Incubator Transport | Draeger/Isolette TI500 | HC13272 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-28
8000216 | 2020-04-28 | RSUD RAA Soewondo Pati | Incubator Transport | Draeger/Isolette TI500 | HC13274 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-28
8000216 | 2020-04-28 | RSUD RAA Soewondo Pati | Resuscitaire | Draeger/Resuscitaire | KC13738 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-28
8000216 | 2020-04-28 | RSUD RAA Soewondo Pati | Resuscitaire | Draeger/Resuscitaire | KC13741 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-28
11000163 | 2020-04-28 | Politeknik Kesehatan Makassar | Jaundice Meter | Draeger/JM 105 | H3500418 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-28
6000302 | 2020-04-27 | RSUD KOTA YOGYAKARTA | Jaundice Meter | Draeger/JM 105 | B3503376 | Survey | Dibawa ke Draeger Surabaya untuk di kalibrasi | Ready to use | Sudah dilakukan perbaikan (6000141) | Afrizal
14000115 | 2020-04-27 | RSUD LAWANG | UPS | FUJI ELECTRIC | 501201114280JC800003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Satria | 2021-04-27
2000126 | 2020-04-25 | RSUD Sayang Rakyat | Digital Radiografi Fluoroskopi | Apelem/Platinum | M196F0403 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Eddy | 2022-04-25
90002447 | 2020-04-24 | RSUD Kab. Bolaang Mongondow | X-ray Digital | Apelem/Camargue DR | CPD30503B18 | Survey | Fuse putus pada generator | Ready to use |  | Irfan
9000213 | 2020-04-23 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | 55115033491 | Survey | Pengecekan alat | Ready to use | Sudah diberikan penawaran. RS belum ada budget | Irfan R
18000123 | 2020-04-22 | RSUD Toto Kabila | Wheel chair | Paramount Bed | 100017447017V | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000123 | 2020-04-22 | RSUD Toto Kabila | Wheel chair | Paramount Bed | 100017447031V | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000123 | 2020-04-22 | RSUD Toto Kabila | Wheel chair | Paramount Bed | 100017447036V | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
18000111 | 2020-04-22 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | VGSLN0092 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-22
18000111 | 2020-04-22 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | VGSLN0100 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-22
18000111 | 2020-04-22 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | VGSLN0109 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-22
18000111 | 2020-04-22 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | VTSLJ0027 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-22
9000106 | 2020-04-22 | RSUD Noongan | Ventilator | Draeger/Savina 300 | ASND0012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-22
9000106 | 2020-04-22 | RSUD Noongan | Ventilator | Draeger/Evita V300 | ASNJ0184 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-22
9000104 | 2020-04-22 | RSUD Noongan | Bed Patient | Paramount | PB2200A-232-C900311500039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
9000101 | 2020-04-21 | RSUD Noongan | Baby Incubator | Draeger/BabyLeo TN500 | ASLL0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-21
9000101 | 2020-04-21 | RSUD Noongan | Baby Incubator | Draeger/BabyLeo TN500 | ASLL0013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-21
9000102 | 2020-04-21 | RSUD Noongan | Ventilator | Draeger/Babylog VN500 | ASMF0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-21
9000101 | 2020-04-21 | RSUD Noongan | Baby Incubator | Draeger/BabyLeo TN500 | ASMJ0025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-21
9000102 | 2020-04-21 | RSUD Noongan | Ventilator | Draeger/Babylog VN500 | ASMK0221 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-21
9000103 | 2020-04-21 | RSUD Noongan | Resuscitaire | Draeger/Resuscitaire | HC05591 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-21
9000103 | 2020-04-21 | RSUD Noongan | Resuscitaire | Draeger/Resuscitaire | HC05591 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-21
9000103 | 2020-04-21 | RSUD Noongan | Resuscitaire | Draeger/Resuscitaire | KC05713 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-21
9000103 | 2020-04-21 | RSUD Noongan | Resuscitaire | Draeger/Resuscitaire | KC05713 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-21
9000103 | 2020-04-21 | RSUD Noongan | Resuscitaire | Draeger/Resuscitaire | KC05715 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-21
9000103 | 2020-04-21 | RSUD Noongan | Resuscitaire | Draeger/Resuscitaire | KC05715 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan | 2021-04-21
9000105 | 2020-04-21 | RSUD Noongan | Bed Patient | Paramount | Pa66250DXS | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
11000157 | 2020-04-20 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Evita V300 | ASMF0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-20
8000200 | 2020-04-20 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0011 | Maintenance | Alarm power fail | Ready To Use |  | Amna
4000145 | 2020-04-20 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Service | Uji Kesesuaian | Ready to use |  | Bayu
4000145 | 2020-04-20 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Service | Perbaikan kolimasi | Ready to use |  | Bayu
6000132 | 2020-04-20 | RS Panti Rapih | Ventilator | Draeger/Savina 300 | ASHA0096 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-04-20
18000109 | 2020-04-17 | RSUD Toto Kabila | Phototherapy | Draeger/Bililux | 1106895050 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-17
18000108 | 2020-04-17 | RSUD Toto Kabila | Mobile X-ray | Apelem/Rafale EV30 | 009/19/01305 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-04-17
18000109 | 2020-04-17 | RSUD Toto Kabila | Resuscitaire | Draeger/Resuscitaire | AC05520 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-17
18000107 | 2020-04-17 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0081 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-17
18000107 | 2020-04-17 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Perseus A500 | ASKK0126 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-17
18000107 | 2020-04-17 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Perseus A500 | ASMM0020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-17
11000164 | 2020-04-17 | RSKD Dadi | Patient Monitor | Draeger/Infinity Delta XL | 6009823462 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-17
11000164 | 2020-04-17 | RSKD Dadi | Patient Monitor | Draeger/Infinity Delta XL | 6009842361 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-17
11000164 | 2020-04-17 | RSKD Dadi | Patient Monitor | Draeger/Infinity Delta XL | 6009853866 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-17
11000164 | 2020-04-17 | RSKD Dadi | Patient Monitor | Draeger/Infinity Delta XL | 6009854072 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-17
8000212 | 2020-04-16 | RSUD Ki Ageng Getas Pondowo Grobogan | Resuscitaire | Draeger/Resuscitaire | AC13515 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-16
8000212 | 2020-04-16 | RSUD Ki Ageng Getas Pondowo Grobogan | Baby Incubator | Draeger/Isolette C2000 | AC37685 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-16
8000212 | 2020-04-16 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-16
8000212 | 2020-04-16 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | ASML0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-16
11000165 | 2020-04-16 | RSKD Dadi | Mesin Anasthesia | Draeger/Primus | ASMJ0189 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-16
14000114 | 2020-04-14 | Laboratorium Klinik Pramita | Bone Mineral Densitometer | J08015D134 | DMS/Stratos DXA | Service | Lampu indikator unit tidak menyala dan unit tidak masuk ke interface Menu | Ready To Use |  | Satria
18000113 | 2020-04-14 | RSUD Toto Kabila | Mobile X-ray | Apelem/Rafale EV30 | 009/19/01305 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2022-04-14
6000131 | 2020-04-13 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger | ASJE0009 | Survey | Kontrak Service | Ready To Use |  | Afrizal
6000135 | 2020-04-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Service | Sensor ILCA tidak berfungsi | Ready to use | Sudah dilakukan perbaikan (6000185) | Afrizal
4000148 | 2020-04-09 | Seksi Kesehatan Lanud Pattimura Ambon | Mobile X-ray | Apelem/Rafale EV30 | 005/18/01387 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2022-04-09
6000302 | 2020-04-09 | RSUD Kab. Temanggung | Mobile X-ray | Apelem/Rafale EV30 | 009/19/01306 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2022-04-09
4000148 | 2020-04-09 | Seksi Kesehatan Lanud Pattimura Ambon | Mobile X-ray | Apelem/Rafale EV30 | 005/18/01387 | Service | Uji fungsi dengan tim | Ready to use |  | Bayu | 2022-04-09
8000301 | 2020-04-09 | Unisia Medika Farma (RS. JIH) | Mesin Anasthesia | x | x | Service | Membongkar breathing circuit untuk Pasien Covid-19 | Ready to use |  | Afrizal
8000215 | 2020-04-08 | RSJD Dr. Amino Gondohutomo | Ventilator | Draeger/Savina 300 | ASMK0306 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-08
18000106 | 2020-04-08 | Dinas Kesehatan Kabupaten Bone Bolango | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASLE0016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-08
11000186 | 2020-04-07 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0138 | Survey | Baterai internal failure | Ready to use |  | Nasrul
11000186 | 2020-04-07 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0395 | Survey | Baterai internal failure | Ready to use |  | Nasrul
11000186 | 2020-04-07 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0396 | Service | Baterai internal failure | Ready to use |  | Nasrul
9000244 | 2020-04-07 | RSUP Prof. Dr. R. D. Kandou Manado | Stretcher | Paramount bed/PK5114X | 176881000003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
9000244 | 2020-04-07 | RSUP Prof. Dr. R. D. Kandou Manado | Stretcher | Paramount bed/PK5114X | 176881000005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
9000244 | 2020-04-07 | RSUP Prof. Dr. R. D. Kandou Manado | Stretcher | Paramount bed/PK5114X | 176881000017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
9000244 | 2020-04-07 | RSUP Prof. Dr. R. D. Kandou Manado | Stretcher | Paramount bed/PK5114X | 176881000023 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
16000112 | 2020-04-06 | Klinik drg. Mirna Hakim | Dental Unit | Kavo | Kavo lux 637 B | Service | Penggantian catridge handpiece | Ready To Use |  | Ipan
6000305 | 2020-04-04 | RSUD Wonosari | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASLE0013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-04-04
6000305 | 2020-04-04 | RSUD Wonosari | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASLE0014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2021-04-04
18000105 | 2020-04-03 | RSUD dr. Zainal Umar Sidiki | Ventilator | Draeger/Savina 300 | ASMK0323 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri | 2021-04-03
8000220 | 2020-04-03 | RSUD RAA Soewondo Pati | Resuscitaire | Draeger/Resuscitaire | DC13253 | Service | Error 05 | Ready to use |  | Amna
8000211 | 2020-04-02 | RSUD K.R.M.T Wongsonegoro | Dental Unit | Kavo/1058TM | 10205318 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-04-02
11000166 | 2020-04-02 | RSKD IBU DAN ANAK PERTIWI | Phototherapy | Draeger/Bililux | 1101999013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-02
11000159 | 2020-04-02 | RSKD IBU DAN ANAK PERTIWI | Baby Incubator | Draeger/Isolette C2000 | AC37686 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-02
11000160 | 2020-04-02 | RSKD IBU DAN ANAK PERTIWI | Infant Warmer | Draeger/Babytherm 1080 | ASRA0036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-02
11000158 | 2020-04-02 | RSKD IBU DAN ANAK PERTIWI | Incubator Transport | Draeger/Isolette TI500 | HC13279 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-02
11000161 | 2020-04-02 | RSKD IBU DAN ANAK PERTIWI | Resuscitaire | Draeger/Resuscitaire | HC13575 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2021-04-02
7000149 | 2020-04-01 | RSUI Harapan Anda | Bubble CPAP | Draeger/Oxymixer | QM20190239 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-04-01
7000148 | 2020-03-31 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASHA0095 | Service | Back up monitor cookpit SN:44243306 | Ready To Use |  | Ariawan
18000104 | 2020-03-31 | RSUD Toto Kabila | Stretcher | Paramount Bed/PK2100 | PK-2100-6320B90020440015 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
8000210 | 2020-03-30 | RSUD Dr. Soeratno Gemolong | Ventilator | Draeger/Carina | ASLM0057 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-03-30
9000243 | 2020-03-30 | RSUP Ratatotok Buyat | X-ray Digital | Apelem/Camargue DR | CPD32474K18 | Service | Kerusakan pada lampu Colimator | Ready to use |  | Irfan R
15000120 | 2020-03-27 | RSUD Harapan dan Doa | X-ray Digital | Apelem/Camargue DR | CPD33564E19 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sutan | 2022-03-27
9000240 | 2020-03-26 | RSUP Ratatotok Buyat | Mesin Anasthesia | Draeger/Fabius Plus | ARWA0005 | Service | Pengecekan berkala | Ready to use |  | Irfan R
13000056 | 2020-03-24 | RSUD Dr. Fauziah | ESWL | Medispec/EM1000 | 0716249 | Service | Plat EMH rusak | Ready to use |  | Sofian
7000147 | 2020-03-23 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Savina 300 | ARWB0347 | Survey | Battery internal fault | Ready To Use |  | Ariawan
6000308 | 2020-03-23 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Service | Gas sensor failure | Ready To Use |  | Afrizal
9000242 | 2020-03-20 | RS Gunung Maria Tomohon | Bed Patient | Paramount | PA932900414E | Survey | Penggantian remote yang rusak | Ready to use |  | Irfan R
6000310 | 2020-03-20 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0002 | Service | Saat selftest, error monitorinh Co2 sensor | Ready to use | Sudah dilakukan perbaikan (6000185) | Afrizal
9000241 | 2020-03-19 | RS Gunung Maria Tomohon | Baby Incubator | Draeger/Isolette C2000 | HC37708 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Irfan | 2021-03-19
9000241 | 2020-03-19 | RS Gunung Maria Tomohon | Baby Incubator | Draeger/Isolette C2000 | NC37698 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Irfan | 2021-03-19
3000136 | 2020-03-18 | RSUD Ulin | Bone Mineral Densitometer | Medilink/Medix 90 | J11015M269 | Service | Error quality control | Ready to use |  | Imam
8000209 | 2020-03-18 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0136 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Amna
8000209 | 2020-03-18 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0138 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Amna
4000147 | 2020-03-18 | RSUD Datoe Binangkang | X-ray Digital | Apelem/Camargue DR | CPD33569E19 | Survey | Pelaksanaan uji kesesuaian dan uji paparan | Ready To Use |  | Bayu
14000113 | 2020-03-17 | RSUD Provinsi NTB | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/032/C0 | Service | Update sofware windows | Ready to use |  | Satria
18000103 | 2020-03-17 | RSUD Toto Kabila | Baby Incubator | Draeger/Isolette C2000 | AC37682 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Sukri | 2021-03-17
18000103 | 2020-03-17 | RSUD Toto Kabila | Baby Incubator | Draeger/Isolette C2000 | AC37683 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Sukri | 2021-03-17
18000103 | 2020-03-17 | RSUD Toto Kabila | Baby Incubator | Draeger/Isolette C2000 | AC37684 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Sukri | 2021-03-17
18000103 | 2020-03-17 | RSUD Toto Kabila | Baby Incubator | Draeger/BabyLeo TN500 | ASLJ0006 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Sukri | 2021-03-17
18000103 | 2020-03-17 | RSUD Toto Kabila | Ventilator | Draeger/Babylog VN500 | ASMK0031 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Sukri | 2021-03-17
6000313 | 2020-03-17 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Delta XL | 6008123373 | Service | Monitor tidak bisa berfungsi (bootloop) | Ready to use |  | Afrizal
8000211 | 2020-03-16 | RSUD Ki Ageng Getas Pondowo Grobogan | Resuscitaire | Draeger/Resuscitaire | AC13515 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-03-16
8000211 | 2020-03-16 | RSUD Ki Ageng Getas Pondowo Grobogan | Baby Incubator | Draeger/Isolette C2000 | AC37685 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-03-16
8000211 | 2020-03-16 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-03-16
8000211 | 2020-03-16 | RSUD Ki Ageng Getas Pondowo Grobogan | Mesin Anasthesia | Draeger/Fabius Plus | ASML0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-03-16
6000314 | 2020-03-16 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Service | Membran tembak ESWL bocor | Ready to use |  | Afrizal
6000312 | 2020-03-16 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0077 | Maintenance | Gas monitoring failure | Ready to use | Sudah dilakukan perbaikan (6000308) | Afrizal
13000064 | 2020-03-13 | RSU Haji Medan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/033/C0 | Service | Update sofware windows | Ready to use |  | Sofian
6000316 | 2020-03-13 | RSU PKU Muhammadiyah Bantul | Mesin Anasthesia | Draeger/Fabius GS Premium | ASEH0119 | Service | O2 sensor fail | Ready to use |  | Afrizal
6000315 | 2020-03-13 | RSUD Panembahan Senopati Bantul | Patient Monitor | Draeger/Vista 120 | V2SHK0007 | Service | ECG communication failed | Ready to use | Alat sudah diperbaiki (6000267) | Afrizal
6000317 | 2020-03-12 | RSUD Prambanan | Patient Monitor | Draeger/Vista 120 | VTSKL0002 | Service | Keypad bermasalah | Ready to use |  | Afrizal
2000125 | 2020-03-12 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Kabel fiber optik digigit tikus | Ready to use |  | Eddy
3000135 | 2020-03-10 | RSUD Dr. Fauziah | ESWL | Medispec/EM1000 | 0716249 | Service | Keluar api pada EMH | ready to use |  | Imam
14000112 | 2020-03-10 | drg. Susana Yasaputera | Dental Unit | Kavo Kerr/1058 | 9889401 | Survey | Lampu unit mati | Ready to use |  | Satria
16000111 | 2020-03-09 | Klinik Drg. Arius Arief | Dental Unit | Kavo | 1002324 | Maintenance | Tumler filter untuk air kumur patah | Ready To Use |  | Ivan
18000101 | 2020-03-09 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0166 | Service | Alat restart sendiri saat device check | Ready to use | Sudah dilakukan perbaikan | Sukri
11000155 | 2020-03-08 | RSUD Kabupaten Kolaka Utara | X-ray Digital | Apelem/Camargue DR | 845889 | Survey | Bucky error karna kabel bucky putus; Perbaikan kabel bucky dan pembersihan generator | Ready to use |  | Nasrul
3000134 | 2020-03-06 | J&J Dental Care | Dental X-Ray | Soredex/Cranex 3D | SE1200223 | Service | 3D App Error | Ready to use |  | Imam
7000150 | 2020-03-06 | RSUD Majenang | Baby Incubator | Draeger/Isolette 8000 | AC09094 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan
7000150 | 2020-03-06 | RSUD Majenang | Baby Incubator | Draeger/Isolette 8000 | AC09135 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2021-03-06
7000150 | 2020-03-06 | RSUD Majenang | Baby Incubator | Draeger/Isolette 8000 | AC09163 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan
6000318 | 2020-03-06 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | ARNM0006 | Service | O2 pressure low | Pending |  | Afrizal
16000110 | 2020-03-05 | Lakes Gikult TNI AU | Dental Unit | Kavo/1058TM | 1007784 | Service | Indikator dentist element tidak menyala | Ready to use |  | Ipan
7000146 | 2020-03-05 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Tiro | ASBM0248 | Service | Tidak bisa settign O2 | Ready To Use |  | Ariawan
7000146 | 2020-03-05 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0025 | Service | Tidak bisa settign O2 | Ready To Use |  | Ariawan
7000146 | 2020-03-05 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0039 | Survey | Tidak bisa settign O2 | Ready To Use |  | Ariawan
13000063 | 2020-03-05 | RSUD Kota Langsa | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/057/C1 | Survey | Menemani tim BPFK Jakarta untuk Ukes | Ready to use |  | Sofian
9000239 | 2020-03-04 | RS Gunung Maria Tomohon | Bed Patient | Paramount | PA93290---0414G900502300002 | Survey | Remot rusak | Ready to use | Sudah diperbaiki oleh Teknisi paramount | Irfan R
9000239 | 2020-03-04 | RS Gunung Maria Tomohon | Bed Patient | Paramount | PA93290---0414G900701200058 | Maintenance | Remot rusak | Ready to use | Sudah diperbaiki oleh Teknisi paramount | Irfan R
8000208 | 2020-03-04 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASJE0014 | Maintenance | General check, system check, running test | Ready to use |  | Amna
8000208 | 2020-03-04 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0033 | Maintenance | General check, system check, running test | Ready to use |  | Amna
8000208 | 2020-03-04 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASKL0036 | Maintenance | General check, system check, running test | Ready to use |  | Amna
8000207 | 2020-03-03 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0136 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-03-03
8000207 | 2020-03-03 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASMK0138 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-03-03
11000156 | 2020-03-02 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 09111371 | Service | EMH tidak maksimal | Ready to use |  | Nasrul
9000047 | 2020-03-02 | RSUP Ratatotok Buyat | Bed Patient | Paramount/2Crank | PL10230A----0420A900102400036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Irfan
5000126 | 2020-02-28 | RS Sentra Medika Cibinong | EECP | Lumenair/Vasomedical | 5TC1064 | Service | Pompa pada bagian kaki tidak berfungsi | Ready to use |  | Bilal
11000154 | 2020-02-27 | RSUD Mamuju | Transport Stratchart | Paramount Bed/K276E | 1684760-00001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000154 | 2020-02-27 | RSUD Mamuju | Transport Stratchart | Paramount Bed/K276E | 1684760-00005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000154 | 2020-02-27 | RSUD Mamuju | Transport Stratchart | Paramount Bed/K276E | 1684760-00016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000153 | 2020-02-27 | RSUD Mamuju | Bed Electric 3 motor | Paramount Bed/3 motor | 9002014-00026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000153 | 2020-02-27 | RSUD Mamuju | Bed Electric 3 motor | Paramount Bed/3 motor | 9002014-00030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000153 | 2020-02-27 | RSUD Mamuju | Bed Electric 3 motor | Paramount Bed/3 motor | 9002014-00035 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000153 | 2020-02-27 | RSUD Mamuju | Bed Electric 3 motor | Paramount Bed/3 motor | 9002014-00036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000153 | 2020-02-27 | RSUD Mamuju | Bed Electric 3 motor | Paramount Bed/3 motor | 9002014-00037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000153 | 2020-02-27 | RSUD Mamuju | Bed Electric 3 motor | Paramount Bed/3 motor | 9002014-00038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
7000145 | 2020-02-27 | RSUI Harapan Anda | Dental Unit | Kavo/1033B | 169 | Survey | Pemasangan motor suction unit | Ready To Use |  | Ariawan
8000299 | 2020-02-26 | RSUD Dr. M. M. Dunda | C-Arm | Apelem/Evo R+ | 55118013636 | Service | Alat mati total | Ready to use | Alat sudah diperbaiki (18000143) | Sukri
6000320 | 2020-02-26 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | V2SHF0244 | Service | Install Manset NIBP | Ready to use |  | Afrizal
6000320 | 2020-02-26 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Vista 120 | V2SHF0274 | Service | Install Manset NIBP | Ready to use |  | Afrizal
8000206 | 2020-02-25 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHD0188 | Service | Gas supply failure and expiratory valve fail | Ready to use |  | Amna
11000152 | 2020-02-25 | RSUD Kota Makassar | ESWL | Medispec/E3000 | 814134 | Maintenance | Mati total, fuse rusak | Ready to use |  | Nasrul
8000205 | 2020-02-24 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ARYJ0111 | Maintenance | General Check, system check, cleaning valve, running test | Ready to use |  | Amna
8000205 | 2020-02-24 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDC0183 | Maintenance | General Check, system check, cleaning valve, running test | Ready to use |  | Amna
8000205 | 2020-02-24 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDC0184 | Maintenance | General Check, system check, cleaning valve, running test | Ready to use |  | Amna
8000205 | 2020-02-24 | RSUD Dr. R. Soetrasno | Ventilator | Draeger/Evita 4 edition | ASEL0138 | Maintenance | General Check, system check, cleaning valve, running test | Ready to use |  | Amna
8000205 | 2020-02-24 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0011 | Maintenance | General Check, system check, cleaning valve, running test | Ready to use |  | Amna
8000205 | 2020-02-24 | RSUD Dr. R. Soetrasno | Ventilator | Draeger/Evita 4 edition | ASKL0120 | Maintenance | General Check, system check, cleaning valve, running test | Ready to use |  | Amna
8000201 | 2020-02-21 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHA0114 | Maintenance | General Check, system check, cleaning valve, running test | Ready to use |  | Amna
8000201 | 2020-02-21 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHC0076 | Maintenance | General Check, system check, cleaning valve, running test | Ready to use |  | Amna
8000201 | 2020-02-21 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Maintenance | General Check, system check, cleaning valve, running test | Ready to use |  | Amna
2000124 | 2020-02-20 | RSUD DR. AGOESDJAM | Digital Radiografi Fluoroskopi | Apelem/Platinum | E196F0342 | Service | Masih kendala pada listrik | Ready to use |  | Eddy | 2022-02-20
8000203 | 2020-02-20 | RSUD Kab Batang | Central Monitor | Draeger/Vista 120 CMS | VLSMF0146 | Service | Blue screen | Ready to use |  | Amna
8000202 | 2020-02-19 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 | ASMN0162 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-02-19
2000123 | 2020-02-18 | RSUD Sultan Syarif Al-Idrus | X-Ray Analog | Apelem/Camargue Analog | CPD30841D18 | Service | Hasil X-ray tidak sempurna; Perbaikan letak kolimasi | Ready to use |  | Eddy
9000297 | 2020-02-17 | RSUD Dr. M. M. Dunda | Mesin Anasthesia | Draeger/Primus | ASLD0027 | Service | No agent | Ready to use |  | Sukri
9000297 | 2020-02-17 | RSUD Dr. M. M. Dunda | Mesin Anasthesia | Draeger/Primus | ASLD0029 | Service | No agent | Ready to use |  | Sukri
7000143 | 2020-02-17 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASKM0028 | Service | Penggantian outlet gas DIN menjadi BS | Ready to use | Disarankan penggjan module control electronic mainboard, valve block brake arm, gas connector | Ariawan
7000143 | 2020-02-17 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASKM0030 | Service | Penggantian outlet gas DIN menjadi BS | Ready to use | Disarankan penggjan module control electronic mainboard, valve block brake arm, gas connector | Ariawan
7000143 | 2020-02-17 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASKM0032 | Service | Penggantian outlet gas DIN menjadi BS | Ready to use | Disarankan penggjan module control electronic mainboard, valve block brake arm, gas connector | Ariawan
7000143 | 2020-02-17 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASKM0033 | Service | Penggantian outlet gas DIN menjadi BS | Ready to use | Disarankan penggjan module control electronic mainboard, valve block brake arm, gas connector | Ariawan
15000110 | 2020-02-15 | RSUD Sekayu | X-ray Digital | Apelem/Camargue DR | CPD33562E19 | Service | Lampu buckey pasien mati | Ready To Use |  | Sutan
15000113 | 2020-02-14 | RSUD Dr. M. Yunus | ESWL | Medispec/EM1000 | 0917011 | Service | Air EMH kotor dan plat EMH berkarat | Ready to use |  | Sutan
5000125 | 2020-02-14 | Suryatenggara Dental Clinic | Dental Unit | Kavo/1058TM | 10010472 | Service | Unit mati total | Ready to use |  | Bilal
7000142 | 2020-02-13 | RSIA Bunda Arif | Jaundice Meter | Draeger/JM 105 | B3501432 | Maintenance | Kalibrasi | Ready To Use |  | Ariawan
9000295 | 2020-02-12 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Babylog VN500 | ASHL0175 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
9000295 | 2020-02-12 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Evita V300 | ASLD0167 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
9000295 | 2020-02-12 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Evita V300 | ASLD0169 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
9000295 | 2020-02-12 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Evita V300 | ASLD0171 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
11000150 | 2020-02-12 | RSUD SYEKH YUSUP KAB. GOWA | Mesin Anasthesia | Draeger/Primus | ASHI0057 | Maintenance | Internal system error | Ready to use |  | Nasrul
15000115 | 2020-02-06 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Maintenance | Cek pergerakan arm, kv/mAs, lampu kolimator | Ready to use |  | Sutan
15000114 | 2020-02-06 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Cek pergerakan table, console joystik, workstation | Ready to use |  | Sutan
13000061 | 2020-02-05 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Service | Gambar di monitor C-arm tidak sebagus di monitor usg | Ready to use | Alat sudah dilakukan perbaikan | Sofian
13000060 | 2020-02-05 | RSUD Dr. Fauziah | ESWL | Medispec/EM1000 | 0716249 | Service | Kurang air dalam EMH | Ready to use |  | Sofian
15000112 | 2020-02-05 | RSUD Sekayu | X-ray Digital | Apelem/Camargue DR | CPD33562E19 | Survey | Lampu buckey putus | Ready To Use |  | Sutan
6000350 | 2020-02-05 | RSUP dr. Soeradji Tirtonegoro | Jaundice Meter | Draeger/JM 105 | B3502776 | Survey | Hasil tidak sesuai AGD | Ready to use |  | Afrizal
6000348 | 2020-02-04 | RSA UGM | Mesin Anasthesia | Draeger/Primus | ASKE0210 | Service | General Check | Ready to use |  | Afrizal
7000141 | 2020-02-04 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASHA0095 | Service | Unit tidak bisa booting dan pressure regulator rusak | Ready to use | Alat digudangkan | Ariawan
16000118 | 2020-02-03 | Klinik Budihardja | Dental Unit | Kavo/1058 TM | 10010476 | Survey | Selang suction (besar) robek | Ready To Use |  | Ipan
6000347 | 2020-01-24 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | ARNM0006 | Service | Machine fault 809, O2 pressure low | Ready to use |  | Afrizal
3000133 | 2020-01-23 | RSUD Kab. Temanggung | C-Arm | Apelem/Evo R+ | 55115011469 | Maintenance | Cant connect to the imager | Ready to use |  | Imam
4000142 | 2020-01-23 | RSUD Kab. Temanggung | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/011/C0 | Service | Update software calibration 2D | Ready to use |  | Bayu
15000111 | 2020-01-22 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Maintenance | Display muncul titik titik | Ready to use |  | Sutan
2000122 | 2020-01-22 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | J166F0184 | Service | Sering muncul error table | Ready to use |  | Eddy
2000121 | 2020-01-21 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 613710 | Service | CCD tidak menampilkan gambar saat expose | Ready to use |  | Eddy
3000132 | 2020-01-17 | Laboratorium Klinik Pramita | Bone Mineral Densitometer | Stratos | H11015D245 | Service | quality control error | Ready to use |  | Imam
5000123 | 2020-01-17 | RS Sentra Medika Cibinong | EECP | Lumenair/Vasomedical | 5TC1064 | Survey | Manset pada kaki tidak mengembang dengan baik | Ready to use |  | Bilal
6000346 | 2020-01-17 | RSUP Dr. Sardjito | Ventilator | Draeger/Evita 4 Edition | ARTK0043 | Service | Vte tidak sesuai setting, batterai drop | Ready to use | Alat digunakan. RS memilih membeli unit baru | Afrizal
6000344 | 2020-01-15 | RSUD Panembahan Senopati Bantul | Medical Compressor | Draeger/Medical Compressor | ASFK0014 | Service | Pressure tidak naik | Ready to use | Alat digudangkan karna penawaran terlalu mahal | Afrizal
6000344 | 2020-01-15 | RSUD Panembahan Senopati Bantul | Medical Compressor | Draeger/Medical Compressor | ASJK0001 | Survey | Pressure tidak naik | Ready to use | Alat digudangkan karna penawaran terlalu mahal | Afrizal
6000345 | 2020-01-15 | RSUP Dr. Sardjito | Ventilator | Draeger/Evita 4 Edition | ARXJ0246 | Service | Pengecekan alat lama tidak terpakai | Ready To Use | Alat digudangkan karna penawaran terlalu mahal | Afrizal
3000131 | 2020-01-14 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helinthus BYM | HELI/0023/CB0 | Survey | Detector error | Ready to use |  | Imam
8000199 | 2020-01-14 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Service | LCD blank | Ready to use |  | Amna
8000199 | 2020-01-14 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Survey | LCD blank | Ready to use |  | Amna
6000343 | 2020-01-14 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | AC09112 | Service | Humidity tidak sesuai setting | Ready to use |  | Afrizal
5000122 | 2020-01-13 | RS Abdi Waluyo | EECP | Vasomedical/Lumenair | 5TC1041 | Service | Penggantain manset pasien dan part yang error | Ready to use |  | Bilal
15000109 | 2020-01-13 | RS. Dr. ERNALDI BAHAR - PALEMBANG | X-Ray Analog | Apelem/Camargue Analog | CPD30672C18 | Service | Dudukan hand switch lepas | Ready to use |  | Sutan
8000198 | 2020-01-13 | RSUD Kab Batang | Baby Incubator | Draeger/Isolette C2000 | HH36864 | Service | alarm oxygen cal different | Ready to use |  | Amna
19000134 | 2020-01-13 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASNE0154 | Service | O2 sensor failure | Ready to use |  | Danu
6000341 | 2020-01-10 | RSUD Tidar Magelang | Medical Compressor | Draeger/Medical Compressor | ASKF0059 | Service | Alarm menyala dan buzzer berbunyi | Ready to use | Alat diperbaiki oleh teknisi RS | Afrizal
17000101 | 2020-01-10 | RSUP Dr. Mohammad Hoesin | Bone Mineral Densitometer | Medilink/Medix 90 | I11015D248 | Service | Generator fault | Ready to use |  | Bahar
14000111 | 2020-01-09 | RSUD Provinsi NTB | Mobile X-ray | Apelem/Rafale EV30 | 009/18/01274 | Service | Roda rusak | Ready to use |  | Satria
4000143 | 2020-01-09 | RSUP Dr. Mohammad Hoesin | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Service | Display muncul titik titik | Ready to use |  | Bayu
4000144 | 2020-01-09 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Survey | Pelaksaaan Uji Kesesuaian dan Uji Paparan | Ready to use |  | Bayu | 2022-01-09
8000197 | 2020-01-08 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita V300 | ASEL0413 | Service | General check, cleaning filter sistem udara, running test | Ready to use |  | Amna
15000108 | 2020-01-08 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Survey | Membran EMH kurang | Ready to use |  | Sutan
6000339 | 2020-01-08 | RSUP Dr. Sardjito | Ventilator | Draeger/Savina 300 | ARTK0043 | Maintenance | Battery drop, humidifer, vt tidak tercapai dan ban roda gesas | Ready to use | Alat digunakan. RS memilih membeli unit baru | Afrizal
8000196 | 2020-01-07 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0094 | Maintenance | General check, cleaning filter sistem udara, running test | Ready To Use |  | Amna
8000196 | 2020-01-07 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLN0009 | Maintenance | General check, cleaning filter sistem udara, running test | Ready To Use |  | Amna
8000196 | 2020-01-07 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLN0011 | Maintenance | General check, cleaning filter sistem udara, running test | Ready To Use |  | Amna
8000196 | 2020-01-07 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASMA0047 | Maintenance | General check, cleaning filter sistem udara, running test | Ready To Use |  | Amna
8000196 | 2020-01-07 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASMD0040 | Maintenance | General check, cleaning filter sistem udara, running test | Ready To Use |  | Amna
8000194 | 2020-01-06 | RSI Sunan Kudus | Ventilator | Draeger/Savina 300 | ARUE0022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-01-06
8000194 | 2020-01-06 | RSI Sunan Kudus | Ventilator | Draeger/Savina 300 | ARUM0308 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2021-01-06
8000195 | 2020-01-06 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Service | O2 inacurate, gas agent tidak akurat | Ready to use | Sudah diperbaiki, claim warranty 31 Mei 2021 | Amna
5000121 | 2019-12-31 | RS Sentra Medika Cibinong | EECP | Lumenair/Vasomedical | 5TC1064 | Service | Tidak menampilkan menu pada monitor | Ready to use |  | Bilal
11000149 | 2019-12-30 | RSIA Defina | Medical Compressor | Draeger/Medical Compressor | ASMJ0007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-12-30
11000143 | 2019-12-27 | RSUD Salewangan | Baby Incubator | Draeger/Isolette C2000 | MC37367 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-12-27
11000144 | 2019-12-27 | Dinas Kesehatan Topoyo | Bed 2 crank | Paramount Bed/2 crank | 1019E90050300011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000144 | 2019-12-27 | Dinas Kesehatan Topoyo | Bed 2 crank | Paramount Bed/2 crank | 1019E90050320010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000144 | 2019-12-27 | Dinas Kesehatan Topoyo | Bed 2 crank | Paramount Bed/2 crank | 1019E90050320012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000144 | 2019-12-27 | Dinas Kesehatan Topoyo | Bed 2 crank | Paramount Bed/2 crank | 1019E90050320014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
8000193 | 2019-12-27 | RSUD Bagas Waras Klaten | Resuscitaire | Draeger/Resuscitaire | JY10902 | Service | Error 05 | Ready to use |  | Amna
11000144 | 2019-12-27 | Dinas Kesehatan Topoyo | Bed 2 crank | Paramount Bed/2 crank | 1019E90053200033 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000147 | 2019-12-23 | RS Universitas Hasanuddin | Baby Incubator | Draeger/Isolette 8000 | HC09164 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-12-23
11000146 | 2019-12-23 | RS Universitas Hasanuddin | Incubator Transport | Draeger/Isolette TI500 | HC13263 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-12-23
11000146 | 2019-12-23 | RS Universitas Hasanuddin | Incubator Transport | Draeger/Isolette TI500 | KC13432 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-12-23
15000107 | 2019-12-20 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0228 | Service | Tidak bisa expose | Ready To Use |  | Sutan
16000107 | 2019-12-18 | Drg. Seriana Lie (Family Dental Care) | Dental Unit | Kavo | 20121210000655 | Service | Handpiece triple bocor | Ready To Use |  | Ivan
8000191 | 2019-12-18 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASLD0028 | Service | sesnor check fail, air supply not detected | Ready to use |  | Amna
7000140 | 2019-12-17 | RSUD Dr. Soeselo Slawi | Ventilator | Draeger/Savina 300 | ASHD0014 | Service | Grafik tidak muncul dan ada kebocoran pada breathing circuit | Ready to use |  | Ariawan
8000190 | 2019-12-17 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASJE0014 | Service | Alarm gas sensor fail | Ready to use |  | Amna
7000139 | 2019-12-16 | RSUI Harapan Anda | Ventilator | Draeger/Savina 300 | ASMM0051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-12-16
7000139 | 2019-12-16 | RSUI Harapan Anda | Ventilator | Draeger/Savina 300 | ASMM0052 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-12-16
9000233 | 2019-12-13 | RSUD Sulawesi Utara | Emergency Troly | Paramount Bed/KY-40QR | 500136600023 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky
9000233 | 2019-12-13 | RSUD Sulawesi Utara | Emergency Troly | Paramount Bed/KY-40QR | 500136600054 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky
9000232 | 2019-12-13 | RSUD Sulawesi Utara | Delivery Bed | Paramount Bed/PD-200 | PD200CB-1119L901113800017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky
9000232 | 2019-12-13 | RSUD Sulawesi Utara | Delivery Bed | Paramount Bed/PD-200 | PD220CB-1119L901113800016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky
7000137 | 2019-12-12 | RSIA Bunda Arif | Jaundice Meter | Draeger/JM 105 | H3500308 | Service | Back up unit | Ready to use |  | Ariawan
7000138 | 2019-12-12 | RSUD DR. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | AH36597 | Maintenance | Hinge access patah | Ready to use |  | Ariawan
4000138 | 2019-12-12 | RSUD Kota Salatiga | General X-Ray | GMM | CPD15937J13 | Survey | Melakukan Uji Paparan & Uji Fungsi | Ready to use |  | Bayu
11000140 | 2019-12-11 | RS Universitas Hasanuddin | Transport Stratchart | Paramount Bed/KK-788E | 145932000013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000140 | 2019-12-11 | RS Universitas Hasanuddin | Transport Stratchart | Paramount Bed/KK-788E | 150450000011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000140 | 2019-12-11 | RS Universitas Hasanuddin | Transport Stratchart | Paramount Bed/KK-788E | 150450000017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000140 | 2019-12-11 | RS Universitas Hasanuddin | Transport Stratchart | Paramount Bed/KK-788E | 156314000007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000140 | 2019-12-11 | RS Universitas Hasanuddin | Transport Stratchart | Paramount Bed/KK-788E | 156314000008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000140 | 2019-12-11 | RS Universitas Hasanuddin | Transport Stratchart | Paramount Bed/KK-788E | 156314000009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000140 | 2019-12-11 | RS Universitas Hasanuddin | Transport Stratchart | Paramount Bed/KK-788E | 156314000010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000141 | 2019-12-11 | RS Universitas Hasanuddin | Delivery Bed | Paramount Bed/PD200BB | 901110000003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000141 | 2019-12-11 | RS Universitas Hasanuddin | Delivery Bed | Paramount Bed/PD200BB | 911110000003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
7000207 | 2019-12-11 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0042 | Service | Pemasangan Service set 2 years dan breathing system | Ready to use |  | Ariawan
7000208 | 2019-12-11 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASKM0032 | Maintenance | Tombol break biru tidak aktif | Pending | Disarankan penggjan module control electronic mainboard, valve block brake arm, gas connector | Ariawan
5000119 | 2019-12-10 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Service | Applicator mengempes dan diisi ulang | Ready to use |  | Bilal
4000139 | 2019-12-10 | RSUD Sekayu | X-ray Digital | Apelem/Camargue DR | CPD33562E19 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2020-12-10
8000189 | 2019-12-09 | RSUD Dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0109 | Pra Instalasi | General check, system check, running test | Ready to use |  | Amna
15000105 | 2019-12-07 | RSUD Sekayu | X-ray Digital | Apelem/Camargue DR | CPD33562E19 | Pra Instalasi | Pemasangan komputer DR, generator, bucky walls, printer trimax | Ready To Use |  | Sutan
8000188 | 2019-12-06 | RSI Sultan Hadlirin | Ventilator | Draeger/Savina 300 | ASMB0172 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-12-06
7000206 | 2019-12-06 | RSUD Banyumas | C-Arm | Apelem/Evo R+ | 55116055585 | Service | Hasil gambar kurang jelas di work station | Ready to use |  | Ariawan
8000187 | 2019-12-05 | RS Panti Wilasa | Ventilator | Draeger/Savina 300 | ASMM0026 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Amna | 2020-12-05
11000139 | 2019-12-05 | Politeknik Kesehatan Makassar | Jaundice Meter | Draeger/JM 105 | HA3500330 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-12-05
9000202 | 2019-12-05 | RSUD Ainun Habibie | Ventilator | Draeger/Evita V300 | ASLD0172 | Service | VT tidak bisa di setting 700 dan kurang dari 10% | Ready to use |  | M. Rizky
9000220 | 2019-12-04 | RSUD Dr. M. M. Dunda | Patient Monitor | Draeger/Vista 120s Advance | S2SLMA0107 | Service | etCO2 tidak berfungsi | Ready to use |  | M. Rizky
4000137 | 2019-12-04 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2020-12-04
5000120 | 2019-12-04 | RSUD Jend. Ahmad Yani | ESWL | Medispec/EM1000 | 1017907 | Service | Tembakan tidak normal | Ready to use |  | Bilal
7000205 | 2019-12-02 | RSUD Banyumas | C-Arm | Apelem/Evo R+ | 55116055585 | Service | Gambar tipis muncul di work station | Ready To Use |  | Ariawan
7000136 | 2019-12-02 | RSUD DR. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASFL0023 | Service | Nebulizer dan pressure regulator tidak berfungsi | Ready to use |  | Ariawan
15000103 | 2019-12-01 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Bucky tidak presisi | Ready To Use |  | Sutan
7000135 | 2019-11-29 | RSUD Brebes | Mesin Anasthesia | Draeger/Primus | ASHM0013 | Service | Sensor check fail | Ready to use |  | Ariawan
15000104 | 2019-11-29 | RSUD Sekayu | X-ray Digital | Apelem/Camargue DR | CPD33562E19 | Pra Instalasi | Pemasangan rel ceilling dan kabel listrik | Ready to use |  | Sutan
6000036 | 2019-11-29 | RSUD KRT. Setjonegoro Wonosobo | Ventilator | Draeger/Savina 300 | ASLC0400 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-11-29
3000129 | 2019-11-28 | RSUP Prof. Dr. R. D. Kandou Manado | C-Arm | Apelem/Evo R+ | 55115033491 | Survey | Display Error code " unavailable inverter power supply" | Ready to use | Alat digudangkan | Imam
4000136 | 2019-11-28 | RSUD Kota Langsa | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/057/C1 | Survey | Melakukan Uji Kesesuaian dan Uji Paparan | Ready to use |  | Bayu
2000129 | 2019-11-28 | RSUP Prof. Dr. R. D. Kandou Manado | C-Arm | Apelem/Evo R+ | 55115033491 | Service | Display error code "unavailable inverter power suppy" | Ready to use | Alat digudangkan | Eddy
3000127 | 2019-11-27 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | J166F0184 | Service | Console Error ET87 | Ready to use |  | Imam
3000128 | 2019-11-27 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 10131182 | Service | Inverter fault | Ready to use |  | Imam
8000186 | 2019-11-27 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASKF0024 | Maintenance | Sensor check fail | Ready to use | Butuh penggantian pressure sensor dan silinder pressure | Amna
4000135 | 2019-11-27 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Survey | Pelaksanaan Uji Kesesuaian & uji paparan | Ready to use |  | Bayu
15000102 | 2019-11-27 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Fluoros tidak bisa | Ready To Use |  | Sutan
11000137 | 2019-11-27 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASAM0044 | Service | Leak system failed | Ready to use |  | Nasrul
7000134 | 2019-11-26 | RSUI Harapan Anda | Ventilator | Draeger/Babylog VN500 | ASLE0048 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-11-26
7000134 | 2019-11-26 | RSUI Harapan Anda | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASMK0055 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-11-26
7000134 | 2019-11-26 | RSUI Harapan Anda | Incubator Transport | Draeger/Isolette TI500 | HC1329 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-11-26
7000204 | 2019-11-25 | RSUD Banyumas | C-Arm | Apelem/Evo R+ | 55116055585 | Service | Hasil tidak muncul lagi | Ready to use |  | Ariawan
2000119 | 2019-11-24 | RSUP Prof. Dr. R. D. Kandou Manado | Laser Surgical | Jena Surgical/MultiPulse HoPlus | JI90100103 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2021-11-24
8000184 | 2019-11-22 | RSUD Kab Batang | Central Monitor | Draeger/Vista 120 CMS | VLSMF0146 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Amna
7000202 | 2019-11-21 | RSUD Banyumas | C-Arm | Apelem/Evo R+ | 55116055585 | Service | Gambar yg muncul tipis dan kurang terang | Ready to use |  | Ariawan
4000134 | 2019-11-21 | RSUD Toto Kabila | X-ray Digital | Apelem/Camargue DR | CPD327474K18 | Survey | Pelaksanaan uji kesesuaian dan uji paparan | Ready to use |  | Bayu | 2020-11-21
8000183 | 2019-11-20 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0011 | Maintenance | General check, device check, system check | Ready to use |  | Amna
8000183 | 2019-11-20 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0093 | Maintenance | General check, device check, system check | Ready to use |  | Amna
8000183 | 2019-11-20 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0094 | Maintenance | General check, device check, system check | Ready to use |  | Amna
8000183 | 2019-11-20 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLN0009 | Maintenance | General check, device check, system check | Ready to use |  | Amna
8000183 | 2019-11-20 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLN0010 | Maintenance | General check, device check, system check | Ready to use |  | Amna
8000183 | 2019-11-20 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASMA0047 | Maintenance | General check, device check, system check | Ready to use |  | Amna
8000183 | 2019-11-20 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASMD0040 | Maintenance | General check, device check, system check | Ready to use |  | Amna
4000133 | 2019-11-19 | RSUD Tani dan Nelayan | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0339 | Survey | Pelaksanaan uji kesesuaian dan uji paparan | Ready to use |  | Bayu | 2020-11-19
15000101 | 2019-11-19 | RSUP Dr. Mohammad Hoesin | ESWL | Medispec/EM1000 | 1116095 | Service | EMH habis | Ready To Use |  | Sutan
11000138 | 2019-11-18 | RSUP Dr. Wahidin Sudirohusodo | Medical Cart | Paramount Bed/KY40QR | 050013660006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000138 | 2019-11-18 | RSUP Dr. Wahidin Sudirohusodo | Medical Cart | Paramount Bed/KY40QR | 050013660008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000138 | 2019-11-18 | RSUP Dr. Wahidin Sudirohusodo | Medical Cart | Paramount Bed/KY40QR | 050013660011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
7000133 | 2019-11-18 | RSUD DR. R. Goeteng Taroenadibrata | Medical Compressor | Draeger/Medical Compressor | ASMJ0011 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Ariawan | 2020-11-18
11000142 | 2019-11-18 | RSKD IBU DAN ANAK PERTIWI | Baby Incubator | Draeger/Isolette C2000 | HC37760 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-11-18
11000142 | 2019-11-18 | RSKD IBU DAN ANAK PERTIWI | Baby Incubator | Draeger/Isolette C2000 | HC37761 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-11-18
11000142 | 2019-11-18 | RSKD IBU DAN ANAK PERTIWI | Baby Incubator | Draeger/Isolette C2000 | HC37762 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-11-18
7000064 | 2019-11-15 | RSUD PREMBUN KAB. KEBUMEN | Resuscitaire | Draeger/Resuscitaire | HH12693 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-11-15
7000132 | 2019-11-15 | RSUD Ajibarang | Ventilator | Draeger/Babylog VN500 | ASMK0054 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Ariawan | 2020-11-15
11000136 | 2019-11-14 | RSU MADANI - PALU | Incubator Transport | Draeger/Isolette TI500 | MH12471 | Service | Mengganti air filter | Ready to use |  | Nasrul
7000202 | 2019-11-14 | RSUD Banyumas | C-Arm | Apelem/Evo R+ | 55116055585 | Service | hasil gambar tidak muncul di work station | Ready to use | Sudah dilakukan perbaikan (7000202) | Ariawan
7000131 | 2019-11-14 | RSUD Banyumas | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0075 | Service | etCO2 tidak berfungsi | Ready To Use |  | Ariawan
2000118 | 2019-11-14 | RSUD Kab. Bolaang Mongondow | X-ray Digital | Apelem/Camargue DR | CPD30503B18 | Service | Grid not moving | Ready to use |  | Eddy
8000182 | 2019-11-14 | RSUD Ki Ageng Getas Pondowo Grobogan | Resuscitaire | Draeger/Resuscitaire | AC13516 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Amna | 2020-11-14
8000182 | 2019-11-14 | RSUD Ki Ageng Getas Pondowo Grobogan | Resuscitaire | Draeger/Resuscitaire | AC13517 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Amna | 2020-11-14
11000134 | 2019-11-13 | RSUD Kabelota | Baby Incubator | Draeger/Isolette C2000 | MU28684 | Service | Motor kipas tidak bekerja | Ready to use |  | Nasrul
7000130 | 2019-11-13 | RSUI Harapan Anda | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASLE0048 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Ariawan | 2020-11-13
7000130 | 2019-11-13 | RSUI Harapan Anda | Ventilator | Draeger/Babylog VN500 | ASMK0055 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Ariawan | 2020-11-13
7000130 | 2019-11-13 | RSUI Harapan Anda | Incubator Transport | Draeger/Isolette TI500 | HC13259 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Ariawan | 2020-11-13
13000058 | 2019-11-13 | RSUD Kota Langsa | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/057/C1 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sofian | 2020-11-13
4000132 | 2019-11-13 | RSUP Prof. Dr. R. D. Kandou Manado | C-Arm | Apelem/Evo R+ | 55115033491 | Service | Hasil kalibrasi tidak masuk, Inverter fault | Ready to use | Alat digudangkan | Bayu
2000117 | 2019-11-13 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | 55115033491 | Service | error data base consule | Ready to use | Sudah diberikan penawaran | Eddy
11000133 | 2019-11-12 | RSIA PARAMOUNT | Delivery Bed | Paramount Bed | 90110020001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000133 | 2019-11-12 | RSIA PARAMOUNT | Delivery Bed | Paramount Bed | 90111420002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
9000218 | 2019-11-12 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | BY33326 | Survey | Keypad tidak berfungsi | Ready to use |  | M. Rizky
9000218 | 2019-11-12 | RSUP Prof. Dr. R. D. Kandou Manado | Baby Incubator | Draeger/Isolette C2000 | BY33331 | Service | Keypad tidak berfungsi | Ready to use |  | M. Rizky
7000201 | 2019-11-11 | RSUD Banyumas | C-Arm | Apelem/Evo R+ | 55116055585 | Service | hasil gambar tidak muncul di work station | Ready to use | Sudah dilakukan perbaikan (7000202) | Ariawan
4000131 | 2019-11-11 | RSUD Noongan | X-ray Digital | Apelem/Camargue DR | CPD26087L16 | Survey | Pelaksanaan uji kesesuaian dan uji paparan | Ready to use |  | Bayu | 2020-11-11
2000116 | 2019-11-11 | RSUD Noongan | X-ray Digital | Apelem/Camargue DR | CPD26087216 | Service | Hasil gambar kurang bagus; Kalibrasi detektor | Ready to use |  | Eddy
80000180 | 2019-11-06 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Service | o2 sensor fail | Ready to use |  | Amna
80000179 | 2019-11-06 | RSUD Dr. Loekmono Hadi | Baby Incubator | Draeger/Isolette 8000 | CU03951 | Service | Hinged asces rusak | Ready to use |  | Amna
11000132 | 2019-11-04 | RSUP Dr. Wahidin Sudirohusodo | Jaundice Meter | Draeger/JM 105 | B3503804 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-11-04
9000217 | 2019-11-02 | RSUD Dr. M. M. Dunda | Medical Compressor | Draeger/Medical Compressor | ASMF0009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-11-02
9000217 | 2019-11-02 | RSUD Dr. M. M. Dunda | Medical Compressor | Draeger/Medical Compressor | ASMF0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-11-02
9000215 | 2019-11-02 | RSUD Dr. M. M. Dunda | Infant Warmer | Draeger/Babytherm 1080 | ASMJ0048 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-11-02
9000215 | 2019-11-02 | RSUD Dr. M. M. Dunda | Infant Warmer | Draeger/Babytherm 1080 | ASMJ0091 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-11-02
2000115 | 2019-11-02 | RSUD Kab. Temanggung | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/011/C0 | Service | Error inverter | Ready to use |  | Eddy
5000116 | 2019-11-01 | RS Siloam Asri | Mesin Anasthesia | Medispec/ED1000 | 0611274 | Service | LCD mati | Ready to use |  | Bilal
9000289 | 2019-11-01 | RSUD Monompiu | Mesin Anasthesia | Draeger/Fabius Plus | ASEM0134 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky
5000118 | 2019-10-31 | RS Sentra Medika Cibinong | EECP | Lumenair/Vasomedical | 5TC1064 | Service | Heart rate selalu tinggi | Ready to use |  | Bilal
9000288 | 2019-10-31 | RSUD KOTAMOBAGU | Mesin Anasthesia | Draeger/Fabius Plus | ASMJ0123 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-31
80000178 | 2019-10-31 | RSUD RA Kartini | Medical Compressor | Draeger/Medical Compressor | ASMJ0012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-10-31
80000177 | 2019-10-30 | RSUD Dr. Soeprapto Cepu | Ventilator | Draeger/Savina 300 | ASJK0211 | Service | General Check | Ready to use |  | Amna
16000106 | 2019-10-29 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1086417 | Service | Key mouth sobek | Ready To Use |  | Ivan
9000290 | 2019-10-29 | RSUD Datoe Binangkang | X-ray Digital | Apelem/Camargue DR | CPD33569E19 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-29
80000174 | 2019-10-29 | RSUD dr. H. Soewondo Kendal | Ventilator | Draeger/Savina 300 | ARTE0292 | Maintenance | General Check | Ready to use |  | Amna
80000174 | 2019-10-29 | RSUD dr. H. Soewondo Kendal | Ventilator | Draeger/Savina 300 | ARXL0178 | Maintenance | General Check | Ready to use |  | Amna
80000176 | 2019-10-29 | RSUD dr. H. Soewondo Kendal | Phototherapy | Draeger/PT 4000 | ASHL0004 | Maintenance | General Check | Ready to use |  | Amna
80000176 | 2019-10-29 | RSUD dr. H. Soewondo Kendal | Phototherapy | Draeger/PT 4000 | ASHL0006 | Maintenance | General Check | Ready to use |  | Amna
80000175 | 2019-10-29 | RSUD dr. H. Soewondo Kendal | Incubator Transport | Draeger/Isolette TI500 | KF10742 | Maintenance | General Check | Ready to use |  | Amna
80000175 | 2019-10-29 | RSUD dr. H. Soewondo Kendal | Incubator Transport | Draeger/Isolette TI500 | KF10748 | Maintenance | General Check | Ready to use |  | Amna
16000105 | 2019-10-28 | FKG Universitas Trisakti | Dental Unit | Kavo/E30TM | 10110730 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan | 2020-10-28
14000110 | 2019-10-25 | Laboratorium Klinik Pramita | Bone Mineral Densitometer | DMS/Stratos DXA | J08015D134 | Service | Unit tidak berfungsi dan indikator mati | Ready To Use |  | Satria
80000173 | 2019-10-25 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0011 | Service | o2 sensor fail | Ready to use |  | Amna
80000173 | 2019-10-25 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius GS Premium | URYH0049 | Maintenance | o2 sensor fail | Ready to use |  | Amna
80000172 | 2019-10-24 | RSJD Dr. Amino Gondohutomo | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0145 | Survey | General Check | Ready to use |  | Amna
9000276 | 2019-10-23 | Siloam International Hospitals Tbk Manado | Ventilator | Draeger/Savina 300 | ASLA0070 | Survey | Training User | Ready to use |  | M. Rizky
9000283 | 2019-10-21 | RS Gunung Maria Tomohon | Bed Patient | Paramount Nbed/KM-303 | 02857 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
16000105 | 2019-10-20 | FKG Universitas Trisakti | Dental Unit | Kavo/E30TM | 10110730 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ipan | 2020-10-20
9000275 | 2019-10-18 | RSUD Ainun Habibie | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0160 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-18
4000130 | 2019-10-18 | RSUD Provinsi NTB | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/032/C0 | Survey | Pengambilan data Acceptance Test Mammo BDT | Ready to use |  | Bayu
70000129 | 2019-10-17 | RSUD Dr. M. Yunus | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/056/C1 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-10-17
9000285 | 2019-10-16 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Babylog VN500 | ASMJ0001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-16
13000055 | 2019-10-16 | RSUD Cut Meutia Lhoksomawe | ESWL | Medispec/EM1000 | 0119127 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Sofian | 2021-10-16
11000130 | 2019-10-15 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 09111371 | Service | EMH tidak maskimal | Ready to use |  | Nasrul
11000129 | 2019-10-15 | RSUP Dr. Wahidin Sudirohusodo | Mamografi | Metaltronica/Helinthus BYM | HELI/0032/CBO | Survey | Alat tidak bisa menyala | Ready to use | Alat digudangkan | Nasrul
14000109 | 2019-10-14 | RSUD Syarifah Ambami Rato Ebu | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/055/C1 | Survey | Training User | Ready To Use |  | Satria
9000287 | 2019-10-12 | RSUD KOTAMOBAGU | Patient Monitor | Draeger/Delta XL | 6009696162 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-12
9000287 | 2019-10-12 | RSUD KOTAMOBAGU | Patient Monitor | Draeger/Delta XL | 6009707463 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-12
9000287 | 2019-10-12 | RSUD KOTAMOBAGU | Patient Monitor | Draeger/Delta XL | 6009713465 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-12
9000287 | 2019-10-12 | RSUD KOTAMOBAGU | Patient Monitor | Draeger/Delta XL | 6009715962 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-12
9000287 | 2019-10-12 | RSUD KOTAMOBAGU | Patient Monitor | Draeger/Delta XL | 6009718068 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-12
9000287 | 2019-10-12 | RSUD KOTAMOBAGU | Patient Monitor | Draeger/Delta XL | 6009722669 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-12
9000287 | 2019-10-12 | RSUD KOTAMOBAGU | Patient Monitor | Draeger/Delta XL | 6009771072 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-12
9000287 | 2019-10-12 | RSUD KOTAMOBAGU | Patient Monitor | Draeger/Delta XL | 6009771469 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-10-12
4000129 | 2019-10-11 | RSUD DR. AGOESDJAM | Digital Radiografi Fluoroskopi | Apelem/Platinum | E196F0342 | Service | Layar console redup | Ready to use |  | Bayu
2000114 | 2019-10-11 | RSUD Kab. Bekasi | C-Arm | Apelem/Evo R+ | 55118026649 | Service | Input AE tittle | Ready to use |  | Eddy
11000127 | 2019-10-10 | Rsia Ananda | Bassinet Carry | Paramount/PB111OC | 0019E90051080001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000127 | 2019-10-10 | Rsia Ananda | Bassinet Carry | Paramount/PB111OC | 0019E90051080002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000127 | 2019-10-10 | Rsia Ananda | Bassinet Carry | Paramount/PB111OC | 0019E90051080003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000127 | 2019-10-10 | Rsia Ananda | Bassinet Carry | Paramount/PB111OC | 0019E90051080004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000127 | 2019-10-10 | Rsia Ananda | Bassinet Carry | Paramount/PB111OC | 0019E90051080005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000127 | 2019-10-10 | Rsia Ananda | Bassinet Carry | Paramount/PB111OC | 0019E90051080006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000127 | 2019-10-10 | Rsia Ananda | Bassinet Carry | Paramount/PB111OC | 0019E90051080009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000127 | 2019-10-10 | Rsia Ananda | Bassinet Carry | Paramount/PB111OC | 0019J90090920001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000127 | 2019-10-10 | Rsia Ananda | Bassinet Carry | Paramount/PB111OC | 0019J90090920005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
4000128 | 2019-10-09 | RSUD Sultan Syarif Al-Idrus | X-Ray Analog | Apelem/Camargue Analog | CPD30841D18 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2020-10-09
80000171 | 2019-10-08 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHD0187 | Service | Breathing circuit check fail | Ready to use |  | Amna
80000171 | 2019-10-08 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHD0187 | Service | Breathing circuit check fail | Ready to use |  | Amna
80000170 | 2019-10-07 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDK0301 | Service | Alarm exp port leakage | Ready to use |  | Amna
80000170 | 2019-10-07 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDK0301 | Service | Alarm exp port leakage | Ready to use |  | Amna
6000400 | 2019-10-04 | RSUD Tidar Magelang | Medical Compressor | Draeger/Medical Compressor | ASKF0059 | Service | motor berisik | Ready to use |  | Afrizal
6000400 | 2019-10-04 | RSUD Tidar Magelang | Resuscitaire | Draeger/Resuscitaire | CY11161 | Service | heater putus, probe putus | Ready to use |  | Afrizal
70000127 | 2019-10-03 | RSUD Ajibarang | Baby Incubator | Draeger/Isolette 8000 | AC09097 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-10-03
70000127 | 2019-10-03 | RSUD Ajibarang | Baby Incubator | Draeger/Isolette 8000 | AC09098 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-10-03
70000126 | 2019-10-01 | RSUD Kardinah Tegal | Incubator Transport | Draeger/Isolette TI500 | AC13228 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-10-01
70000126 | 2019-10-01 | RSUD Kardinah Tegal | Incubator Transport | Draeger/Isolette TI500 | AC13229 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-10-01
70000126 | 2019-10-01 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0082 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-10-01
80000168 | 2019-09-30 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Service | Sensor O2 error | Ready to use |  | Amna
80000169 | 2019-09-30 | RSUD LIMPUNG | Mesin Anasthesia | Draeger/Fabius Plus | AC37622 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-09-30
80000169 | 2019-09-30 | RSUD LIMPUNG | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0146 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-09-30
80000169 | 2019-09-30 | RSUD LIMPUNG | Mesin Anasthesia | Draeger/Fabius Plus | FC37527 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-09-30
70000125 | 2019-09-28 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASKM0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-09-28
70000125 | 2019-09-28 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASKM0030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-09-28
70000125 | 2019-09-28 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASMH0036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-09-28
70000125 | 2019-09-28 | RSUD Hj. Anna Lasmanah | Pendant | Draeger/Movita | ASMH0037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-09-28
80000167 | 2019-09-27 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASMA0130 | Service | Kontras LCD bergerak | Ready to use |  | Amna
6000399 | 2019-09-25 | RSKBD Panti Nugroho | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0139 | Survey | Safety relief valve failed | Ready to use |  | Afrizal
2000113 | 2019-09-25 | RSUD Malingping | X-ray Digital | Apelem/Camargue DR | CPD30501B18 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2021-09-25
9000282 | 2019-09-24 | RSUD Liunkendage Tahuna | Bed Patient | Paramount Bed/PAR-12210 | 1219C 900710300266 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Sukri
70000124 | 2019-09-23 | RST Wijayakusuma | Ventilator | Draeger/Savina 300 | ASEL0195 | Service | Main switch error | Pending |  | Ariawan
70000124 | 2019-09-23 | RST Wijayakusuma | Ventilator | Draeger/Savina 300 | ASEL0196 | Service | Battery internal fault | Pending |  | Ariawan
70000124 | 2019-09-23 | RST Wijayakusuma | Ventilator | Draeger/Savina 300 | ASEL0202 | Service | Main switch error | Pending |  | Ariawan
80000166 | 2019-09-23 | RSUD Kajen | Patient Monitor | Draeger/Vista 120 | V3SHJ0006 | Service | NIBP tidak bekerja | Ready to use |  | Amna
80000166 | 2019-09-23 | RSUD Kajen | Patient Monitor | Draeger/Vista 120 | X3SHJ0022 | Service | NIBP tidak bekerja | Ready to use |  | Amna
1000010 | 2019-09-23 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | Error software | Ready to use |  | Sudirman
80000165 | 2019-09-20 | RS ISLAM PATI | Mesin Anasthesia | Draeger/Fabius | ASLK0114 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-09-20
6000398 | 2019-09-20 | RSUD Panembahan Senopati Bantul | ESWL | Medispec/E3000 | 0719139 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-09-20
3000124 | 2019-09-18 | RSUD Jend. Ahmad Yani | C-Arm | Apelem/Evo R+ | 55116054584 | Survey | Training User | Ready to use |  | Imam
3000125 | 2019-09-18 | RSD Mayjen HM Ryacudu | X-ray Digital | Apelem/Camargue DR | IMA703939 | Survey | Hasilexpose tidak baik; Kalibrasi detektor dan kalibrasi tube | Ready to use |  | Imam
16000104 | 2019-09-17 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058TM | 10007769 | Service | Alat kumur netes | Ready To Use |  | Ivan
11000125 | 2019-09-17 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Draeger/Isolette C2000 | WE2236 | Service | Motor kipas berbunyi bising | Ready to use | Sudah dilakukan perbaikan | Nasrul
2000112 | 2019-09-16 | RS Dr. Suyoto | X-Ray Analog | Apelem/Camargue Analog | CPD27102B17 | Service | Hasil expose rendah kualitas | Ready to use |  | Eddy
6000396 | 2019-09-16 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette C2000 | WE22383 | Service | Upgrade humidity kit | Ready to use |  | Afrizal
11000124 | 2019-09-13 | RSUD Konawe | Patient Monitor | Draeger/Vista 120 | VTSLC0085 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-09-13
80000164 | 2019-09-13 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Service | Notify internal system error | Ready to use |  | Amna
3000123 | 2019-09-12 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186S0271 | Service | Software crash | Ready to use |  | Imam
11000126 | 2019-09-11 | RSUD Andi Makassau | Incubator Transport | Draeger/Isolette TI500 | FC13145 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-09-11
11000126 | 2019-09-11 | RSUD Andi Makassau | Incubator Transport | Draeger/Isolette TI500 | FC13147 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-09-11
80000163 | 2019-09-11 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLN0011 | Survey | LCD Display M 14.5 fail | Ready to use |  | Amna
11000123 | 2019-09-10 | RSUD Anutalako | Ventilator | Draeger/Evita V300 | ASNA0097 | Service | Training User | Ready to use |  | Nasrul
14000108 | 2019-09-09 | Medicelle Clinic | Breast Biopsy System | BARD/ENCOR ULTRA | DYCZT514 | Maintenance | alarm vacuum | Ready To Use |  | Satria
4000123 | 2019-09-06 | RSUP Dr. M. Djamil Padang | ESWL | Medispec/E3000 | 0613835 | Survey | Pelaksaaan uji kesesuaian | Ready to use |  | Bayu
4000124 | 2019-09-06 | RSUP Dr. M. Djamil Padang | EECP | Medispec | 06111188 | Survey | Aplicator rusak | Ready to use | Alat sudah discontinue | Bayu
4000124 | 2019-09-06 | RSUP Dr. M. Djamil Padang | EECP | Lumenair/Vasomedical | STC1093 | Service | Probe Spo2 rusak | Ready To Use |  | Bayu
9000262 | 2019-09-06 | Siloam International Hospitals Tbk Manado | Ventilator | Draeger/Savina 300 | ASLA0070 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-09-06
9000269 | 2019-09-05 | RSUD Noongan | Patient Monitor | Draeger/Delta XL | 6009749453 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-09-05
9000269 | 2019-09-05 | RSUD Noongan | Patient Monitor | Draeger/Delta XL | 6009764668 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-09-05
9000269 | 2019-09-05 | RSUD Noongan | Patient Monitor | Draeger/Delta XL | 6009766363 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-09-05
9000269 | 2019-09-05 | RSUD Noongan | Patient Monitor | Draeger/Delta XL | 6009767166 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-09-05
9000269 | 2019-09-05 | RSUD Noongan | Patient Monitor | Draeger/Delta XL | 6009768263 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-09-05
9000269 | 2019-09-05 | RSUD Noongan | Patient Monitor | Draeger/Delta XL | 6009768763 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-09-05
9000269 | 2019-09-05 | RSUD Noongan | Patient Monitor | Draeger/Delta XL | 60009767362 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-09-05
9000270 | 2019-09-05 | RSUD Noongan | Ventilator | Draeger/Evita V300 | ASME0151 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-09-05
9000270 | 2019-09-05 | RSUD Noongan | Ventilator | Draeger/Evita V300 | ASME0152 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-09-05
1000009 | 2019-09-05 | RSUD Ulin | ESWL | Medispec/E3000 | 09131173 | Maintenance | Alat tidak mengeluarkan shock | Ready To Use |  | Sudirman
80000162 | 2019-09-03 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius GS Premium | ARYJ 0111 | Maintenance | General Check | Ready to use |  | Amna
80000162 | 2019-09-03 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0138 | Maintenance | General Check | Ready to use |  | Amna
80000162 | 2019-09-03 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0011 | Maintenance | General Check | Ready to use |  | Amna
80000162 | 2019-09-03 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0120 | Maintenance | General Check | Ready to use |  | Amna
1000008 | 2019-09-03 | RSUD TARAKAN | Mamografi | Metaltronica/Lilyum | ILILHF/0709/CB2 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Sudirman | 2021-09-03
3000122 | 2019-09-02 | RSUD Kab Tangerang | C-arm | Apelem/DSA 500 | 5511161245 | Service | Software spash | Ready to use |  | Imam
16000103 | 2019-08-30 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1610046179 | Service | Mengganti valve | Ready To Use |  | Ivan
11000122 | 2019-08-30 | RSUD Lakipadada | Stretcher | Paramount Bed/K726E | 12785300005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000122 | 2019-08-30 | RSUD Lakipadada | Stretcher | Paramount Bed/K726E | 12785300008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000122 | 2019-08-30 | RSUD Lakipadada | Stretcher | Paramount Bed/K726E | 13764100013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000122 | 2019-08-30 | RSUD Lakipadada | Stretcher | Paramount Bed/K726E | 13764100014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
14000107 | 2019-08-30 | RSUD Provinsi NTB | C-Arm | Apelem/Evo R+ | 55112078337 | Service | Filamen error | Ready To Use |  | Satria
1000007 | 2019-08-29 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | Error software pada ws platinum | Ready to use |  | Sudirman
80000161 | 2019-08-29 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASHD0034 | Maintenance | Leak dan breathing system high | Ready to use |  | Amna
80000161 | 2019-08-29 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0108 | Maintenance | General Check | Ready to use |  | Amna
80000161 | 2019-08-29 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0110 | Maintenance | General Check | Ready to use |  | Amna
80000161 | 2019-08-29 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0117 | Maintenance | General Check | Ready to use |  | Amna
4000122 | 2019-08-27 | RSUD Dr. Adjidarmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Survey | Pelaksaaan uji kesesuaian | Ready to use |  | Bayu
80000160 | 2019-08-27 | RSUD Dr. R. Soedjati Soemodiarjo | Medical Compressor | Draeger/Medical Compressor | ASME0001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-08-27
80000160 | 2019-08-27 | RSUD Dr. R. Soedjati Soemodiarjo | Medical Compressor | Draeger/Medical Compressor | ASME0002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-08-27
70000120 | 2019-08-26 | RST Wijayakusuma | Mesin Anasthesia | Draeger/Fabius Plus | ASEC0176 | Service | Leak system error | Ready to use |  | Ariawan
70000120 | 2019-08-26 | RST Wijayakusuma | Mesin Anasthesia | Draeger/Fabius Plus | ASEC0181 | Service | Leak system error | Ready to use |  | Ariawan
9000257 | 2019-08-26 | RSUD Prof. Dr. H. Aloei Saboe | Ventilator | Draeger/Evita V300 | ASME0153 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-26
9000258 | 2019-08-26 | RSUD Prof. Dr. H. Aloei Saboe | Baby Incubator | Draeger/Isolette C2000 | MC37369 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-26
9000258 | 2019-08-26 | RSUD Prof. Dr. H. Aloei Saboe | Baby Incubator | Draeger/Isolette C2000 | MC37370 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-26
2000110 | 2019-08-24 | RSUD Kab. Temanggung | X-Ray Analog | Apelem/Camargue Analog | CPD30673C18 | Service | Penyimpangan kolimasi | Ready to use |  | Eddy
2000109 | 2019-08-23 | RSUD Kab. Temanggung | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/011/C0 | Maintenance | Tidak bisa expose | Ready to use |  | Eddy
4000121 | 2019-08-23 | RSUD Sultan Suriansyah | X-ray Digital | Apelem/Camargue DR | CPD33483019 | Maintenance | Pelaksaaan uji kesesuaian | Ready to use |  | Bayu
3000121 | 2019-08-22 | RSUD Ulin | Bone Mineral Densitometer | Medilink/Medix 90 | J130105M269 | Maintenance | Alat tidak running | Ready to use |  | Imam
14000106 | 2019-08-22 | Medicelle Clinic | Breast Biopsy System | BARD/ENCOR ULTRA | DYCZT514 | Survey | Pemindahan alat dari RS Manyar Medical Center ke Medicelle Clinic | Ready To Use |  | Satria
2000108 | 2019-08-21 | RSUD RAA Soewondo Pati | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/054/C1 | Service | Instalasi user | Ready to use |  | Eddy
16000103 | 2019-08-20 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 161046179 | Service | Mengganti valve | Ready to use |  | Ipan
70000119 | 2019-08-20 | RSUP Dr. Kariadi | Mobile X-ray | Apelem/Rafale EV30 | 009/18/01271 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-08-20
80000181 | 2019-08-20 | RSUD RAA Soewondo Pati | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/054/C1 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-08-20
5000115 | 2019-08-19 | RS Sentra Medika Cibinong | EECP | Lumenair/Vasomedical | 5TC1064 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal | 2020-08-19
3000120 | 2019-08-16 | RSUD Kab Tangerang | C-arm | Apelem/DSA 500 | 5511161245 | Service | Komputer C-Arm tidak berfungsi | Ready to use |  | Imam
16000102 | 2019-08-16 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058TM | 10007769 | Service | Air bowl tidak mengalir | Ready To Use |  | Ivan
16000102 | 2019-08-16 | Klinik drg. Robert H. Dharma | Dental Unit | Kavo/1058TM | 10007769 | Survey | Water bowl tidak mengalir dan suction tidak bekerja | Ready to use |  | Ipan
70000121 | 2019-08-15 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0042 | Service | Ventilator valve error | Ready to use |  | Ariawan
80000158 | 2019-08-15 | RSUD RAA Soewondo Pati | Baby Incubator | Draeger/Isolette 8000 | MC09036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-08-15
9000254 | 2019-08-13 | RSUD Dr. M. M. Dunda | Patient Monitor | Draeger/Vista 120 S | S2SMA0058 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-13
9000254 | 2019-08-13 | RSUD Dr. M. M. Dunda | Patient Monitor | Draeger/Vista 120 S | S2SMA0062 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-13
9000254 | 2019-08-13 | RSUD Dr. M. M. Dunda | Patient Monitor | Draeger/Vista 120 S | S2SMA0066 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-13
9000254 | 2019-08-13 | RSUD Dr. M. M. Dunda | Patient Monitor | Draeger/Vista 120 S | S2SMA0070 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-13
9000254 | 2019-08-13 | RSUD Dr. M. M. Dunda | Patient Monitor | Draeger/Vista 120 S | S2SMA0071 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-13
9000254 | 2019-08-13 | RSUD Dr. M. M. Dunda | Patient Monitor | Draeger/Vista 120 S | S2SMA0073 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-13
9000254 | 2019-08-13 | RSUD Dr. M. M. Dunda | Patient Monitor | Draeger/Vista 120 S | S2SMA0080 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-13
9000254 | 2019-08-13 | RSUD Dr. M. M. Dunda | Patient Monitor | Draeger/Vista 120 S | S2SMA0107 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-13
80000157 | 2019-08-13 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Survey | Gas sensor inaccurate | Ready to use |  | Amna
5000114 | 2019-08-12 | RS Abdi Waluyo | ESWT | Medispec/ED1000 | 909876 | Survey | Aplicator harus di ganti | Ready to use |  | Bilal
2000107 | 2019-08-10 | RSUD Dr. Adjidarmo | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0340 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2020-08-10
9000201 | 2019-08-09 | RSUD Datoe Binangkang | Patient Monitor | Draeger/Vista 120 | S2SLN0004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-09
9000201 | 2019-08-09 | RSUD Datoe Binangkang | Patient Monitor | Draeger/Vista 120 | S2SLN0009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-09
9000201 | 2019-08-09 | RSUD Datoe Binangkang | Patient Monitor | Draeger/Vista 120 | S2SLN0015 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-09
9000201 | 2019-08-09 | RSUD Datoe Binangkang | Patient Monitor | Draeger/Vista 120 | S2SLN0022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-09
9000201 | 2019-08-09 | RSUD Datoe Binangkang | Patient Monitor | Draeger/Vista 120 | S2SLN0029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-09
9000201 | 2019-08-09 | RSUD Datoe Binangkang | Patient Monitor | Draeger/Vista 120 | S2SLN0030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-09
9000201 | 2019-08-09 | RSUD Datoe Binangkang | Patient Monitor | Draeger/Vista 120 | S2SLN0036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-09
9000201 | 2019-08-09 | RSUD Datoe Binangkang | Patient Monitor | Draeger/Vista 120 | S2SLN0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-09
9000201 | 2019-08-09 | RSUD Datoe Binangkang | Patient Monitor | Draeger/Vista 120 | S2SLN0070 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-08-09
4000119 | 2019-08-09 | RSUD Kota Salatiga | ESWL | Medispec/E3000 | 09131174 | Service | Kapasitor HVG meledak | Ready to use |  | Bayu
6000387 | 2019-08-08 | RSKBD Panti Nugroho | Baby Incubator | Draeger/Isolette C2000 | EH37103 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-08-08
6000387 | 2019-08-08 | RSKBD Panti Nugroho | Baby Incubator | Draeger/Isolette C2000 | EH37104 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-08-08
6000387 | 2019-08-08 | RSKBD Panti Nugroho | Baby Incubator | Draeger/Isolette C2000 | KH37079 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-08-08
12000042 | 2019-08-08 | RSUD Dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0109 | Service | mesin kekurangan part O2 sensor | Ready to use |  | Amna
80000156 | 2019-08-07 | RSUD RA Kartini | Lampu Operasi | Draeger/Polaris 600 | ASLL0226 | Service | Lampu mati | Ready to use | Sudah diperbaiki. Claim warranty 18 September 2019 | Amna
11000116 | 2019-08-06 | Puskesmas Polocamba | Bed 2 crank | Paramount Bed/2 crank | 900503200018 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000116 | 2019-08-06 | Puskesmas Polocamba | Bed 2 crank | Paramount Bed/2 crank | 900503200019 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000116 | 2019-08-06 | Puskesmas Polocamba | Bed 2 crank | Paramount Bed/2 crank | 900503200020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
9000203 | 2019-08-06 | RSUD Ainun Habibie | Resuscitaire | Draeger/Resuscitaire | DY03434 | Survey | Training User | Ready to use |  | M. Rizky
9000203 | 2019-08-06 | RSUD Ainun Habibie | Resuscitaire | Draeger/Resuscitaire | DY03437 | Service | Training User | Ready to use |  | M. Rizky
6000386 | 2019-08-06 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASKJ0043 | Maintenance | gas analyzer dan O2 sensor blinking yellow | Ready to use |  | Afrizal
80000155 | 2019-08-05 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | Maintenance | General Check | Ready to use |  | Amna
80000155 | 2019-08-05 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Maintenance | General Check | Ready to use |  | Amna
9000204 | 2019-08-05 | RSUD Prof. Dr. H. Aloei Saboe | Bed Patient | Paramount/PA6445CBBABA | 9000311300071 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky
11000115 | 2019-08-01 | RSUD Kab. Bombana | Mesin Anasthesia | Draeger/Fabius Plus | ASHP0010 | Service | Tidal volume tidak sesuai setting | Ready to use | Sudah dilakukan perbaikan | Nasrul
11000113 | 2019-07-30 | RSUD Bahteramas | Phototherapy | Draeger/PT 4000 | ASKE0254 | Service | Device check | Ready to use |  | Nasrul
11000113 | 2019-07-30 | RSUD Bahteramas | Baby Incubator | Draeger/Caleo | ASKL0010 | Service | Device check | Ready to use |  | Nasrul
11000113 | 2019-07-30 | RSUD Bahteramas | Baby Incubator | Draeger/Caleo | ASKL0021 | Service | Device check | Ready to use |  | Nasrul
11000113 | 2019-07-30 | RSUD Bahteramas | Baby Incubator | Draeger/Caleo | ASLE0001 | Service | Device check | Ready to use |  | Nasrul
70000118 | 2019-07-30 | RSUD Banyumas | Mesin Anasthesia | Draeger/Primus A500 | ASKF0075 | Service | Gas sensor failure | Ready to use |  | Ariawan
11000112 | 2019-07-29 | RSUD Kota Makassar | Ventilator | Draeger/Evita V300 | ASHM0194 | Service | Alarm O2 kalibrasi | Ready to use |  | Nasrul
70000117 | 2019-07-29 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002705 | Service | Penggantian Main pcb | Ready to use |  | Ariawan
70000117 | 2019-07-29 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002722 | Service | Penggantian Main pcb | Ready to use |  | Ariawan
11000111 | 2019-07-26 | RSUD Undata Palu | Infant Warmer | Draeger/Babytherm 8010 | ASME0068 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-07-26
80000154 | 2019-07-23 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Maintenance | Instalasi of Service set 2 years | Ready to use |  | Amna
80000154 | 2019-07-23 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Service | Instalasi of Service set 2 years | Ready to use |  | Amna
3000119 | 2019-07-22 | RSUD Sultan Suriansyah | X-ray Digital | Apelem/Camargue DR | CPD33483019 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2020-07-22
9000093 | 2019-07-22 | RSUD Liunkendage Tahuna | Mesin Anasthesia | Draeger/Fabius Plus | ASBH0201 | Survey | Pengantaran O2 cell dan Selang | Ready to use |  | M. Rizky
6000384 | 2019-07-21 | RSUP dr. Soeradji Tirtonegoro | Jaundice Meter | Draeger/JM 105 | B3502042 | Service | Unit dikembalikan ke PT Draeger Medical Indonesia | Ready to use |  | Afrizal
6000383 | 2019-07-19 | RSUP dr. Soeradji Tirtonegoro | Jaundice Meter | Draeger/JM 105 | B3502776 | Service | Serah terima unit | Ready to use |  | Afrizal
11000110 | 2019-07-18 | RSUP Tadjudin Chalid | Patient Monitor | Draeger/Vista 120 | VTSJK0109 | Service | NIBP tidak berfungsi | Ready to use |  | Nasrul
11000110 | 2019-07-18 | RSUP Tadjudin Chalid | Patient Monitor | Draeger/Vista 120 | VTSJM0080 | Service | NIBP tidak berfungsi | Ready to use |  | Nasrul
70000116 | 2019-07-18 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0072 | Service | Penggantian O2 sensor | Ready to use |  | Ariawan
70000116 | 2019-07-18 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0048 | Service | Penggantian O2 sensor | Ready to use |  | Ariawan
11000109 | 2019-07-17 | RSUD Kota Makassar | Bed Electric | Paramount Bed/PA-6415 | 90031130065 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000109 | 2019-07-17 | RSUD Kota Makassar | Bed Electric | Paramount Bed/PA-6415 | 90031130079 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000109 | 2019-07-17 | RSUD Kota Makassar | Bed Electric | Paramount Bed/PA-6415 | 90031130083 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000109 | 2019-07-17 | RSUD Kota Makassar | Bed Electric | Paramount Bed/PA-6415 | 900031130064 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000109 | 2019-07-17 | RSUD Kota Makassar | Bed Electric | Paramount Bed/PA-6415 | 900031130078 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
13000052 | 2019-07-17 | RSU Haji Medan | Bone Mineral Densitometer | Medilink/Medix DR | D197110560 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Sofian | 2020-07-17
70000113 | 2019-07-17 | RSKBD Panti Nugroho | Ventilator | Draeger/Evita V300 | ASLD0168 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-07-17
70000115 | 2019-07-17 | RSKBD Panti Nugroho | Incubator Transport | Draeger/Isolette TI500 | NC128691 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-07-17
70000115 | 2019-07-17 | RSKBD Panti Nugroho | Incubator Transport | Draeger/Isolette TI500 | NC12870 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-07-17
70000114 | 2019-07-17 | RSKBD Panti Nugroho | Patient Monitor | Draeger/Vista 120 | VTSLK0037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-07-17
80000152 | 2019-07-17 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 Edition | ASEL 0142 | Service | o2 sensor fail | Ready to use |  | Amna
80000152 | 2019-07-17 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 Edition | ASEL 0143 | Maintenance | Instalasi of Service set 2 years | Ready to use |  | Amna
80000153 | 2019-07-17 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASJF0033 | Service | Instalasi of Service set 2 years | Ready to use |  | Amna
70000114 | 2019-07-17 | RSKBD Panti Nugroho | Patient Monitor | Draeger/Vista 120 | VTSLK0038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-07-17
70000114 | 2019-07-17 | RSKBD Panti Nugroho | Patient Monitor | Draeger/Vista 120 | VTSLK0043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-07-17
70000114 | 2019-07-17 | RSKBD Panti Nugroho | Patient Monitor | Draeger/Vista 120 | VTSLK0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-07-17
70000114 | 2019-07-17 | RSKBD Panti Nugroho | Patient Monitor | Draeger/Vista 120 | VTSLK0050 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-07-17
70000114 | 2019-07-17 | RSKBD Panti Nugroho | Patient Monitor | Draeger/Vista 120 | VTSLK0053 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-07-17
6000381 | 2019-07-16 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | FY06852 | Service | Pemasangan spare part evaporator | Ready to use |  | Afrizal
6000381 | 2019-07-16 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | FY06854 | Service | Pemasangan spare part evaporator | Ready to use |  | Afrizal
6000380 | 2019-07-13 | RS Panti Rapih | Mesin Anasthesia | Draeger/Fabius Plus | ARZM0048 | Survey | Pemasangan Service set 3years | Ready to use |  | Afrizal
6000382 | 2019-07-13 | RSUP Dr. Sardjito | Patient Monitor | Draeger/Delta XL | 6008123373 | Service | Monitor tidak bisa inisialisasi | Ready to use | Sudah diperbaiki (6000313) | Afrizal
3000010 | 2019-07-12 | RSUD Malingping | X-ray Digital | Apelem/Camargue DR | IMA703916 | Survey | Training User | Ready to use |  | Imam
2000106 | 2019-07-12 | RSUD DR. AGOESDJAM | Digital Radiografi Fluoroskopi | Apelem/Platinum | E196F0342 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2020-07-12
80000151 | 2019-07-09 | RSUD Kab Batang | Incubator Transport | Draeger/Isolette TI500 | CF10703 | Service | Battery fail alarm | Ready to use |  | Amna
9000089 | 2019-07-08 | RSUD Maria Walanda Maramis | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0062 | Service | O2 sensor failed | Ready to use |  | M. Rizky
6000379 | 2019-07-08 | RSUD Panembahan Senopati Bantul | Compressor | Draeger/GS500 | 3151989 | Survey | Internal failure | Ready to use |  | Afrizal
6000379 | 2019-07-08 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASJH0121 | Service | General check | Ready to use |  | Afrizal
70000112 | 2019-07-06 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0071 | Service | Penggantian O2 sensor | Ready to use |  | Ariawan
70000112 | 2019-07-06 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0073 | Service | Penggantian O2 sensor | Ready to use |  | Ariawan
80000150 | 2019-07-06 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASJE0014 | Service | Sensor check fail | Ready to use |  | Amna
7000111 | 2019-07-06 | RSUI Harapan Anda | Ventilator | Draeger/Babylog VN500 | ASJC0050 | Service | Alarm Humidifier | Ready to use |  | Ariawan
80000149 | 2019-07-05 | RSUD Kajen | Lampu Operasi | Draeger/Polaris 600 | ASMA0360 | Service | Installasi | Ready to use |  | Amna
6000378 | 2019-07-04 | RSA UGM | Mesin Anasthesia | Draeger/Primus | ASKE0210 | Service | VT tidak sesuai setting | Ready to use |  | Afrizal
1000118 | 2019-07-04 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Service | Hasil kolimasi tidak sesuai | Ready to use |  | Bayu
11000108 | 2019-07-03 | RSUD Tenriawaru Bone | Patient Monitor | Draeger/Vista 120 | VTSLL0054 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-07-03
11000108 | 2019-07-03 | RSUD Tenriawaru Bone | Patient Monitor | Draeger/Vista 120 | VTSLL0055 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-07-03
11000108 | 2019-07-03 | RSUD Tenriawaru Bone | Patient Monitor | Draeger/Vista 120 | VTSLL0062 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-07-03
11000108 | 2019-07-03 | RSUD Tenriawaru Bone | Patient Monitor | Draeger/Vista 120 | VTSLL0063 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-07-03
9000094 | 2019-07-03 | RSUD Tani dan Nelayan | Mesin Anasthesia | Draeger/Perseus A500 | ASKJ0080 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-07-03
4000117 | 2019-07-02 | RSUD Dr. Drajat Prawiranegara | ESWL | Medispec/EM1000 | 0610748 | Survey | Uji fungsi dengan tim | Ready to use |  | Bayu | 2021-07-02
9000082 | 2019-07-02 | RSUD Dr. M. M. Dunda | Bed Patient | Paramount/4Motor | 900911400022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky
80000148 | 2019-07-02 | RSUD PANDAN ARANG | SystemB | Medispec/ED1000 | 0713907 | Service | Tembakan tidak teratur | Ready to use |  | Amna
16000101 | 2019-07-01 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 161046179 | Service | Foot switch mengeluarkan suara angin | Ready to use |  | Ipan
16000101 | 2019-07-01 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1610046179 | Service | Foot switch mengeluarkan suara angin | Ready To Use |  | Ivan
80000147 | 2019-07-01 | RSUD Bung Karno | Mesin Anasthesia | Draeger/Primus | ASMA0134 | Survey | Melakukan Uji fungsi dengan tim | Ready to use |  | Amna
80000147 | 2019-07-01 | RSUD Bung Karno | Mesin Anasthesia | Draeger/Primus | ASMA0192 | Survey | Melakukan Uji fungsi dengan tim | Ready to use |  | Amna
80000147 | 2019-07-01 | RSUD Bung Karno | Mesin Anasthesia | Draeger/Primus | ASMA0194 | Survey | Melakukan Uji fungsi dengan tim | Ready to use |  | Amna
80000147 | 2019-07-01 | RSUD Bung Karno | Mesin Anasthesia | Draeger/Primus | ASMA0195 | Survey | Melakukan Uji fungsi dengan tim | Ready to use |  | Amna
9000085 | 2019-07-01 | RSUD Toto Kabila | Baby Incubator | Draeger/Isolette C2000 | BC37248 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-07-01
9000086 | 2019-07-01 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | S2SLN0005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-07-01
9000086 | 2019-07-01 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | S2SLN0013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-07-01
9000086 | 2019-07-01 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | S2SLN0016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-07-01
9000086 | 2019-07-01 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | S2SLN0017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-07-01
9000086 | 2019-07-01 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | S2SLN0023 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-07-01
9000083 | 2019-06-28 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Fabius Plus | ASKK0057 | Service | Selang flow meter bocor | Ready to use |  | M. Rizky
9000077 | 2019-06-28 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | S2SLN0002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-06-28
9000077 | 2019-06-28 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | S2SLN0005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-06-28
9000077 | 2019-06-28 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | S2SLN0013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-06-28
9000077 | 2019-06-28 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | S2SLN0016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-06-28
9000077 | 2019-06-28 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | S2SLN0017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-06-28
9000077 | 2019-06-28 | RSUD Toto Kabila | Patient Monitor | Draeger/Vista 120 | S2SLN0023 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-06-28
11000107 | 2019-06-28 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius | ASAM0044 | Survey | O2 failure | Ready to use |  | Nasrul
6000376 | 2019-06-26 | RS Pratama Dinkes | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0048 | Service | O2 cell kadaluarsa | Ready to use |  | Afrizal
4000116 | 2019-06-26 | RSUD Dr. Drajat Prawiranegara | ESWL | Medispec/EM1000 | 0610748 | Survey | Uji fungsi penembakan batu | Ready to use |  | Bayu | 2021-06-26
11000106 | 2019-06-26 | RSUD Undata Palu | Baby Incubator | Draeger/Isolette C2000 | BC37249 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-06-26
11000105 | 2019-06-26 | RSUD Undata Palu | Patient Monitor | Draeger/Vista 120 | VTSLL0050 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-06-26
11000105 | 2019-06-26 | RSUD Undata Palu | Patient Monitor | Draeger/Vista 120 | VTSLL0056 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-06-26
11000105 | 2019-06-26 | RSUD Undata Palu | Patient Monitor | Draeger/Vista 120 | VTSLL0058 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-06-26
7000110 | 2019-06-26 | RSUD Prof. Dr. Margono Soekarjo | Jaundice Meter | Draeger/JM 105 | B3503802 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan
6000375 | 2019-06-25 | RSUP Dr. Sardjito | Incubator Transport | Draeger/Isolette TI500 | CY02409 | Service | flow oksigen tidak naik | Ready to use |  | Afrizal
7000108 | 2019-06-24 | RSUD Dr. M. Ashari | Mesin Anasthesia | Draeger/Primus | ASMA0131 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-06-24
9000078 | 2019-06-24 | RSUD Liunkendage Tahuna | Mesin Anasthesia | Draeger/Fabius Plus | ASBH0201 | Maintenance | Alarm O2 failed | Ready to use |  | M. Rizky
7000109 | 2019-06-24 | RSUD Suradadi Kab. Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0182 | Maintenance | Kalibrasi O2 sensor | Ready to use |  | Ariawan
3000118 | 2019-06-22 | RSUD Tani dan Nelayan | Digital Radiografi Fluoroskopi | Apelem/Platinum | D196F0339 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2020-06-22
8000145 | 2019-06-22 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Maintenance | General Check | Ready to use |  | Amna
8000145 | 2019-06-22 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Maintenance | General Check | Ready to use |  | Amna
3000116 | 2019-06-20 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | J166F0184 | Service | Keypad concole tidak berfungsi, cek koneksi kabel pada console | Ready to use |  | Imam
1000005 | 2019-06-20 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Service | Alat tidak bisa expose | Ready To Use |  | Sudirman
7000107 | 2019-06-19 | RS Prima Medika | Ventilator | Draeger/Savina 300 | ASHN-0021 | Maintenance | Penggantian O2 sensor | Ready to use |  | Ariawan
9000076 | 2019-06-19 | RSUD Noongan | Mesin Anasthesia | Draeger/Perseus A500 | ASJF0093 | Maintenance | Melakukan system test | Ready to use |  | M. Rizky
9000076 | 2019-06-19 | RSUD Noongan | Mesin Anasthesia | Draeger/Perseus A500 | ASJL0006 | Maintenance | Melakukan system test | Ready to use |  | M. Rizky
9000076 | 2019-06-19 | RSUD Noongan | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0066 | Maintenance | Melakukan system test | Ready to use |  | M. Rizky
8000144 | 2019-06-18 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASHL0063 | Maintenance | Replace Service set 3 years | Ready to use |  | Amna
8000144 | 2019-06-18 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASJE0013 | Service | Replace Service set 3 years | Ready to use |  | Amna
8000143 | 2019-06-17 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASJE0014 | Service | o2 sensor fail | Ready to use | Sudah diperbaiki pada pengecekan selanjutnya (80000150) | Amna
5000113 | 2019-06-14 | RSUD Dr. Drajat Prawiranegara | ESWL | Medispec/EM1000 | 0617048 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal | 2020-06-14
1000004 | 2019-06-14 | RSUD Ulin | ESWL | Medispec/E3000 | 09131173 | Service | Alat tiba tiba mati | Ready To Use |  | Sudirman
9000087 | 2019-06-13 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Evita V300 | ASLD0169 | Service | Touchscreen tidak berfungsi | Ready to use |  | M. Rizky
80000142 | 2019-06-12 | RSUD RA Kartini | Patient Monitor | Draeger/Vista 120 | VTSLK0029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-12
80000142 | 2019-06-12 | RSUD RA Kartini | Patient Monitor | Draeger/Vista 120 | VTSLK0036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-12
80000142 | 2019-06-12 | RSUD RA Kartini | Patient Monitor | Draeger/Vista 120 | VTSLK0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-12
80000142 | 2019-06-12 | RSUD RA Kartini | Patient Monitor | Draeger/Vista 120 | VTSLK0051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-12
80000142 | 2019-06-12 | RSUD RA Kartini | Patient Monitor | Draeger/Vista 120 | VTSLK0054 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-12
4000115 | 2019-06-12 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0917012 | Survey | EMH Expired | Ready to use |  | Bayu
80000127 | 2019-06-11 | RSUD Kab Batang | Patient Monitor | Draeger/Vista 120 S | S2SLN0001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-11
80000127 | 2019-06-11 | RSUD Kab Batang | Patient Monitor | Draeger/Vista 120 S | S2SLN0010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-11
80000127 | 2019-06-11 | RSUD Kab Batang | Patient Monitor | Draeger/Vista 120 S | S2SLN0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-11
80000127 | 2019-06-11 | RSUD Kab Batang | Patient Monitor | Draeger/Vista 120 S | S2SLN0012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-11
80000127 | 2019-06-11 | RSUD Kab Batang | Patient Monitor | Draeger/Vista 120 S | S2SLN0012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-11
80000127 | 2019-06-11 | RSUD Kab Batang | Patient Monitor | Draeger/Vista 120 S | S2SLN0018 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-11
80000127 | 2019-06-11 | RSUD Kab Batang | Patient Monitor | Draeger/Vista 120 S | S2SLN0019 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-11
80000127 | 2019-06-11 | RSUD Kab Batang | Patient Monitor | Draeger/Vista 120 S | S2SLN0025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-11
80000127 | 2019-06-11 | RSUD Kab Batang | Patient Monitor | Draeger/Vista 120 S | S2SLN0026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-11
80000127 | 2019-06-11 | RSUD Kab Batang | Patient Monitor | Draeger/Vista 120 S | S2SLN0035 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-11
80000127 | 2019-06-11 | RSUD Kab Batang | Patient Monitor | Draeger/Vista 120 S | S2SLN0041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-06-11
3000117 | 2019-06-10 | RSD Mayjen HM Ryacudu | X-ray Digital | Apelem/Camargue DR | IMA703939 | Service | System OS tidak running; Penggantian HDD 1TB, reinstall software dan kalibrasi detektor | Ready to use |  | Imam
13000050 | 2019-06-10 | RSUD Dr. Fauziah | ESWL | Medispec/EM1000 | 716249 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Sofian | 2021-06-10
2000105 | 2019-06-02 | RSUP Prof. Dr. R. D. Kandou Manado (KSO) | CT Scan | Philips MCT-480 | 129286 | Service | eeror gantry cannot initialize | Ready to use |  | Eddy
7000106 | 2019-05-31 | RSUD DR. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASEM0003 | Maintenance | No air supply | Pending | Sudah diberikan penawaran | Ariawan
4000114 | 2019-05-29 | RS Advent Bandar Lampung | Mobile X-ray | Apelem/Rafale EV30 | 0051801391 | Survey | Pelaksanaan uji kesesuaian | Ready to use |  | Bayu
6000373 | 2019-05-29 | RSUD Bagas Waras Klaten | Ventilator | Draeger/Babylog VN500 | ASJD0018 | Service | Internal power supply failure | Ready to use |  | Afrizal
5000112 | 2019-05-28 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 108847 | Service | Tekanan handpiece error | Ready to use |  | Bilal
80000141 | 2019-05-28 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDK0300 | Service | Exp Leakage port alarm | Ready to use |  | Amna
6000372 | 2019-05-28 | RSUD Sleman | Incubator Transport | Draeger/Isolette TI500 | JC1027 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-05-28
6000372 | 2019-05-28 | RSUD Sleman | Resuscitaire | Draeger/Resuscitaire | JC13210 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-05-28
6000372 | 2019-05-28 | RSUD Sleman | Resuscitaire | Draeger/Resuscitaire | JC13211 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-05-28
11000114 | 2019-05-28 | RSUD Konawe | Infant Warmer | Draeger/Babytherm 8004 | ASMF0047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-05-28
11000114 | 2019-05-28 | RSUD Konawe | Infant Warmer | Draeger/Babytherm 8004 | ASMF0053 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-05-28
7000105 | 2019-05-28 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASEF0108 | Maintenance | Device check | Ready to use |  | Ariawan
7000105 | 2019-05-28 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASHA0100 | Maintenance | Device check | Ready to use |  | Ariawan
7000105 | 2019-05-28 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHC0171 | Maintenance | General Check | Ready to use |  | Ariawan
11000101 | 2019-05-28 | RSUD Kabupaten Konawa | Baby Incubator | Draeger/Isolette 8000 | HH08604 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-05-28
11000101 | 2019-05-28 | RSUD Kabupaten Konawa | Baby Incubator | Draeger/Isolette 8000 | HH08606 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-05-28
11000101 | 2019-05-28 | RSUD Kabupaten Konawa | Baby Incubator | Draeger/Isolette 8000 | HH08607 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-05-28
11000102 | 2019-05-28 | RSUD Kabupaten Konawa | Incubator Transport | Draeger/Isolette TI500 | MC12868 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-05-28
11000101 | 2019-05-28 | RSUD Kabupaten Konawa | Baby Incubator | Draeger/Isolette 8000 | NC08791 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-05-28
7000104 | 2019-05-28 | RSUI Harapan Anda | Ventilator | Draeger/Savina 300 | ASHA0095 | Service | Unit tidak bisa dimatikan | Ready to use |  | Ariawan
80000140 | 2019-05-27 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLN0009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-27
80000140 | 2019-05-27 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLN0010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-27
80000140 | 2019-05-27 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASLN0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-27
80000140 | 2019-05-27 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASMA0047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-27
80000140 | 2019-05-27 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASMD0040 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-27
80000139 | 2019-05-24 | RSUD RA Kartini | Lampu Operasi | Draeger/Polaris 600 | ASLL0226 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-24
80000139 | 2019-05-24 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Primus | ASMA0193 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-24
9000256 | 2019-05-24 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Evita V300 | ASLD0167 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-05-24
9000256 | 2019-05-24 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Evita V300 | ASLD0170 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-05-24
6000371 | 2019-05-23 | RSUD Kab. Temanggung | Air compressor | Draeger/GS500 | 25141848 | Service | Compressor tidak aktif, diganti dengan unit backup | Ready to use |  | Afrizal
7000103 | 2019-05-22 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASLD0042 | Survey | Claim warranty | Ready to use |  | Ariawan
5000106 | 2019-05-20 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1086410 | Service | Hexagonal pada footswitch | Ready to use |  | Bilal
9000080 | 2019-05-20 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | ASJD0019 | Service | O2 PATO bermasalah | Ready to use |  | M. Rizky
7000102 | 2019-05-16 | RSUD Bumiayu | Baby Incubator | Draeger/BabyLeo TN500 | ASLJ0005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-05-16
7000102 | 2019-05-16 | RSUD Bumiayu | Jaundice Meter | Draeger/JM 105 | B4503805 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-05-16
80000138 | 2019-05-16 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASLD0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-16
80000137 | 2019-05-16 | RSI Sultan Agung | Lampu Operasi | Draeger/Polaris 600 | ASM0357 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-16
80000137 | 2019-05-16 | RSI Sultan Agung | Lampu Operasi | Draeger/Polaris 600 | ASMA0359 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-16
9000073 | 2019-05-15 | RSUD Ainun Habibie | Ventilator | Draeger/Evita V300 | ASLD0172 | Service | Expiratory valve error | Ready to use |  | M. Rizky
11000094 | 2019-05-15 | RSUD Undata Palu | Mesin Anasthesia | Draeger/Primus | ASLD0052 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-05-15
11000092 | 2019-05-14 | RSU MADANI - PALU | Emergency Stretcher | Paramount/PK 2100 | 5719D-900401200046 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000092 | 2019-05-14 | RSU MADANI - PALU | Emergency Stretcher | Paramount/PK 2100 | 5719D-900401200047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000093 | 2019-05-14 | RSU MADANI - PALU | Incubator Transport | Draeger/Isolette TI500 | CH12772 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2020-05-14
9000070 | 2019-05-14 | RSUD Prof. Dr. H. Aloei Saboe | Mesin Anasthesia | Draeger/Fabius Plus | ASJH0080 | Service | O2 sensor failed | Ready to use |  | M. Rizky
9000072 | 2019-05-14 | RSUD Prof. Dr. H. Aloei Saboe | Mesin Anasthesia | Draeger/Fabius Plus | ASJL0005 | Service | O2 sensor failed | Ready to use |  | M. Rizky
9000081 | 2019-05-14 | RSUD Prof. Dr. H. Aloei Saboe | Mesin Anasthesia | Draeger/Fabius Plus | ASJL0008 | Service | O2 sensor failed | Ready to use |  | M. Rizky
9000068 | 2019-05-13 | RSUD Dr Mamdinda | Ventilator | Draeger/Evita V300 | ASHL0175 | Service | Touchscreen tidak berfungsi | Ready to use |  | M. Rizky
9000069 | 2019-05-13 | RSUD Prof. Dr. H. Aloei Saboe | Mesin Anasthesia | Draeger/Fabius Plus | ASJK0228 | Service | O2 sensor failed | Ready to use |  | M. Rizky
4000113 | 2019-05-13 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0917012 | Survey | Uji coba tembakan | Ready to use |  | Bayu
80000135 | 2019-05-10 | RSUD dr. R. Soetijono Blora | Ventilator | Draeger/Babylog VN500 | ASLF0054 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-10
80000135 | 2019-05-10 | RSUD dr. R. Soetijono Blora | Ventilator | Draeger/Babylog VN500 | ASLF0054 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-10
8000136 | 2019-05-10 | RSUD dr. R. Soetijono Blora | Incubator Transport | Draeger/Isolette TI500 | HH12696 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-10
9000074 | 2019-05-09 | RSUD Dr. M. M. Dunda | Mesin Anasthesia | Draeger/Primus | ASLD0027 | Survey | Training User | Ready to use |  | M. Rizky
2000104 | 2019-05-09 | RSUD Dr. M. M. Dunda | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/048/C0 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2020-05-09
7000101 | 2019-05-09 | RSUD Hj. Anna Lasmanah | Resuscitaire | Draeger/Resuscitaire | CY11218 | Service | Error 05 | Ready to use |  | Ariawan
7000101 | 2019-05-09 | RSUD Hj. Anna Lasmanah | Resuscitaire | Draeger/Resuscitaire | CY11222 | Service | Kit corner block | Ready to use |  | Ariawan
5000110 | 2019-05-09 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0917012 | Service | Karet pada EMH bocor | Ready to use |  | Bilal
80000133 | 2019-05-08 | RSUD RA Kartini | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/006/C0 | Maintenance | Detector tidak merespon | Ready to use | Sudah diperbaiki pada pengecekan selanjutnya (8000245) | Amna
3000114 | 2019-05-07 | Laboratorium Klinik Pramita | Bone Mineral Densitometer | Stratos | HII013D245 | Service | Upgrade software | Ready to use |  | Imam
9000075 | 2019-05-07 | RSUD Dr. M. M. Dunda | Mesin Anasthesia | Draeger/Primus | ASLD0027 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-05-07
3000113 | 2019-05-06 | RSUP Dr. Hasan Sadikin | Mamografi | Metaltronica/Helinthus BYM | HELI/0023/CB0 | Service | Setting printer | Ready to use |  | Imam
5000109 | 2019-05-06 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1086410 | Service | Air pada unit dental tidak keluar | Ready to use |  | Bilal
4000112 | 2019-05-06 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0917012 | Service | Tidak fokus pada titik +15 | Ready to use |  | Bayu
4000112 | 2019-05-06 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0917012 | Maintenance | Tidak fokus pada titik +15 | Ready to use |  | Bayu
3000115 | 2019-05-04 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | Ji66F0184 | Service | Monitor consule blank | Ready to use |  | Imam
3000112 | 2019-05-02 | Laboratorium Klinik Pramita | Bone Mineral Densitometer | Stratos | J08015D134 | Survey | Software stratos dan windows crash | Ready to use |  | Imam
6000362 | 2019-05-02 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHA0113 | Service | Blower pada kompressor tidak berfungsi | Ready to use |  | Afrizal
7000101A | 2019-05-02 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08776 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-05-02
7000101A | 2019-05-02 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08777 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-05-02
7000101A | 2019-05-02 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08778 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-05-02
7000101A | 2019-05-02 | RSUD Kardinah Tegal | Baby Incubator | Draeger/Isolette 8000 | EH08779 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-05-02
80000132 | 2019-05-02 | RSUD Kajen | Mesin Anasthesia | Draeger/Primus | ASHA0233 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-02
80000132 | 2019-05-02 | RSUD Kajen | Mesin Anasthesia | Draeger/Primus | ASLD0023 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-05-02
3000111 | 2019-04-30 | RSUD Jend. Ahmad Yani | C-Arm | Apelem/Evo R+ | 55116054524 | Survey | Gambar terbalik | Ready to use |  | Imam
14000104 | 2019-04-30 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1012854 | Service | Pembersihan selang headpiece | Ready To Use |  | Satria
6000363 | 2019-04-30 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette 8000 | FY06853 | Service | Sensor disconnect alarm | Ready to use |  | Afrizal
8000131 | 2019-04-30 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHA0114 | Service | TV not constant | Ready to use |  | Amna
14000103 | 2019-04-29 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1012854 | Service | Headpiece tidak bisa mengeluarkan angin | Ready To Use |  | Satria
5000107 | 2019-04-29 | RS Siloam Asri | Baby Incubator | Medispec/ED1000 | 6111274 | Service | Monitor pada alat mati | Ready to use |  | Bilal
80000130 | 2019-04-29 | RSUD Dr. Soediran Mangun Sumarso | Baby Incubator | Draeger/BabyLeo TN500 | ASLJ0004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-04-29
80000128 | 2019-04-29 | RSUD Dr. Soediran Mangun Sumarso | Resuscitaire | Draeger/Resuscitaire | BC13128 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-04-29
80000128 | 2019-04-29 | RSUD Dr. Soediran Mangun Sumarso | Resuscitaire | Draeger/Resuscitaire | BC13129 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-04-29
80000128 | 2019-04-29 | RSUD Dr. Soediran Mangun Sumarso | Resuscitaire | Draeger/Resuscitaire | BC13130 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-04-29
80000128 | 2019-04-29 | RSUD Dr. Soediran Mangun Sumarso | Resuscitaire | Draeger/Resuscitaire | BC13131 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-04-29
80000128 | 2019-04-29 | RSUD Dr. Soediran Mangun Sumarso | Resuscitaire | Draeger/Resuscitaire | BC13132 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-04-29
80000129 | 2019-04-29 | RSUD Dr. Soediran Mangun Sumarso | Baby Incubator | Draeger/Isolette C2000 | BC37247 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-04-29
80000128 | 2019-04-29 | RSUD Dr. Soediran Mangun Sumarso | Resuscitaire | Draeger/Resuscitaire | NC12996 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-04-29
80000126 | 2019-04-27 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASHD0034 | Service | Replace Service set 2 years | Ready to use |  | Amna
14000102 | 2019-04-26 | RSUD Provinsi NTB | Digital Radiografi Fluoroskopi | Apelem/Platinum | B196F0329 | Instalasi | Alat berfungsi dengan baik | Ready To Use |  | Satria | 2020-04-26
6000361 | 2019-04-26 | RSUP dr. Soeradji Tirtonegoro | Jaundice Meter | Draeger/JM 105 | B3502042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-04-26
80000125 | 2019-04-25 | RSI Sultan Agung | Lampu Operasi | Draeger/Polaris 600 | ASM0357 | Maintenance | Pra Installasi | Ready to use |  | Amna
4000109 | 2019-04-25 | RSUD Provinsi NTB | Mobile X-ray | Apelem/Rafale EV30 | 009/18/01273 | Maintenance | Pelaksanaan uji kesesuaian | Ready to use |  | Bayu
4000109 | 2019-04-25 | RSUD Provinsi NTB | Mobile X-ray | Apelem/Rafale EV30 | 009/18/01274 | Maintenance | Pelaksanaan uji kesesuaian | Ready to use |  | Bayu
4000110 | 2019-04-25 | RSUD Provinsi NTB | Digital Radiografi Fluoroskopi | Apelem/Platinum | B196F0329 | Survey | Pelaksanaan uji kesesuaian | Ready to use |  | Bayu
3000110 | 2019-04-24 | Laboratorium Klinik Pramita | Bone Mineral Densitometer | Stratos | HII015D245 | Service | Cek alat | Ready to use |  | Imam
6000359A | 2019-04-24 | RS Panti Rini | Mesin Anasthesia | Draeger/Fabius Plus | ASFL0007 | Maintenance | Kalibrasi o2 sensor | Ready to use |  | Afrizal
7000099 | 2019-04-24 | RSUD DR. Soedirman - Kebumen | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASLE0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2020-04-24
80000124 | 2019-04-24 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Fabius Plus | ASLK0116 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-04-24
80000124 | 2019-04-24 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Primus | ASLD0040 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-04-24
7000098 | 2019-04-22 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08430 | Service | LH patah | Ready to use |  | Ariawan
7000098 | 2019-04-22 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08433 | Survey | LH patah | Ready to use |  | Ariawan
6000360 | 2019-04-21 | RSUD Tidar Magelang | Resuscitaire | Draeger/Resuscitaire | CY11161 | Maintenance | General check | Ready to use |  | Afrizal
80000123 | 2019-04-18 | RSJD Dr. Amino Gondohutomo | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0145 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-04-18
11000089 | 2019-04-18 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 9111371 | Service | Penggantian unit EMH | Ready to use |  | Nasrul
7000097 | 2019-04-16 | RSUD Dr. Soeselo Slawi | Baby Incubator | Draeger/Isolette C2000 | CF32766 | Service | Access panel patah | Ready to use |  | Amna
80000122 | 2019-04-16 | RSUD Kraton Pekalongan | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASHM0009 | Service | Alarm paw measurment inop | Ready to use |  | Amna
3000109 | 2019-04-15 | RSUD Toto Kabila | X-ray Digital | Apelem/Camargue DR | CPD327474K18 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2021-04-15
80000121 | 2019-04-15 | RS Indriati Solo | Mesin Anasthesia | Draeger/Primus | ASJF0036 | Service | PFM Failure | Ready to use |  | Amna
4000108 | 2019-04-12 | RSUD Sultan Syarif Al-Idrus | X-Ray Analog | Apelem/Camargue Analog | CPD30841D18 | Pra Instalasi | Setting bucky,kalibrasi tube,pelaksaan uji paparan | Ready to use |  | Bayu | 2021-04-12
9000253 | 2019-04-12 | RSUD Datoe Binangkang | Lampu Periksa | Draeger/Polaris 50 | 109580003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-12
9000253 | 2019-04-12 | RSUD Datoe Binangkang | Lampu Periksa | Draeger/Polaris 50 | 1094901004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-12
9000253 | 2019-04-12 | RSUD Datoe Binangkang | Lampu Periksa | Draeger/Polaris 50 | 1095580002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-12
9000253 | 2019-04-12 | RSUD Datoe Binangkang | Lampu Periksa | Draeger/Polaris 50 | 1095580004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-12
9000253 | 2019-04-12 | RSUD Datoe Binangkang | Lampu Periksa | Draeger/Polaris 50 | 1095580005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-12
9000253 | 2019-04-12 | RSUD Datoe Binangkang | Lampu Periksa | Draeger/Polaris 50 | 1095656001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-12
9000253 | 2019-04-12 | RSUD Datoe Binangkang | Lampu Periksa | Draeger/Polaris 50 | 1095656002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-12
9000259 | 2019-04-12 | RSUD Datoe Binangkang | Lampu Operasi | Draeger/Polaris 600 | ASLL0228 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-12
9000259 | 2019-04-12 | RSUD Datoe Binangkang | Lampu Operasi | Draeger/Polaris 600 | ASLL0234 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-12
80000120 | 2019-04-12 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0011 | Maintenance | General Check | Ready to use |  | Amna
9000259 | 2019-04-12 | RSUD Datoe Binangkang | Lampu Operasi | Draeger/Polaris 600 | ASMA2336 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-12
11000085 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Transport Stratchart | Paramount | 1178120-00014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000085 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Transport Stratchart | Paramount | 1178130-00008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000085 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Transport Stratchart | Paramount | 1178130-00009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000085 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Transport Stratchart | Paramount | 1178130-0001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000085 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Transport Stratchart | Paramount | 1178130-00010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000085 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Transport Stratchart | Paramount | 1178130-0002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000085 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Transport Stratchart | Paramount | 1178130-0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000085 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Transport Stratchart | Paramount | 1178130-0004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000085 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Transport Stratchart | Paramount | 1178130-0005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000087 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | ICU Bed | Paramount Bed/DA 66250DXS | 9003061-00006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000087 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | ICU Bed | Paramount Bed/DA 66250DXS | 9003061-00008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000087 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | ICU Bed | Paramount Bed/DA 66250DXS | 9003061-00012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000087 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | ICU Bed | Paramount Bed/DA 66250DXS | 9003061-00013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000087 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | ICU Bed | Paramount Bed/DA 66250DXS | 9003061-00024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000087 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | ICU Bed | Paramount Bed/DA 66250DXS | 9003061-0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000087 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | ICU Bed | Paramount Bed/DA 66250DXS | 9003061-0014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000087 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | ICU Bed | Paramount Bed/DA 66250DXS | 9011013-00082 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000087 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | ICU Bed | Paramount Bed/DA 66250DXS | 9011013-0083 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000086 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Medical Cart | Paramount Bed/KY40QR | OS008257-00003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000086 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Medical Cart | Paramount Bed/KY40QR | OS008257-00005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000086 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Medical Cart | Paramount Bed/KY40QR | OS008257-00014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000086 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Medical Cart | Paramount Bed/KY40QR | OS008257-00055 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000086 | 2019-04-11 | RSUD Prof. ANWAR MAKKATUTU | Medical Cart | Paramount Bed/KY40QR | OS008257-00058 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
9000066 | 2019-04-11 | RSUD Kab. Bolaang Mongondow | Printer Fuji | DRYDIX Smart 2000 | 66916486 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-11
7000096 | 2019-04-10 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASKL0108 | Service | Blower mati total | Ready to use |  | Ariawan
9000065 | 2019-04-10 | RSUD KOTAMOBAGU | Printer Fuji | DRYDIX Smart 6000 | 76435078 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-10
9000066 | 2019-04-10 | RSUD KOTAMOBAGU | Printer Fuji | DRYDIX Smart 6000 | 76435092 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | M. Rizky | 2020-04-10
11000083 | 2019-04-09 | RSUD Bahteramas | Baby Incubator | Draeger/Caleo | ASKL0010 | Service | Device check | Ready to use |  | Nasrul
11000083 | 2019-04-09 | RSUD Bahteramas | Baby Incubator | Draeger/Caleo | ASKL0021 | Service | Device check | Ready to use |  | Nasrul
11000083 | 2019-04-09 | RSUD Bahteramas | Baby Incubator | Draeger/Caleo | ASLG-0001 | Service | Device check | Ready to use |  | Nasrul
11000082 | 2019-04-09 | RSUD Bahteramas | Incubator Transport | Draeger/Isolette TI500 | MH12469 | Service | Battery failure | Ready to use |  | Nasrul
11000082 | 2019-04-09 | RSUD Bahteramas | Incubator Transport | Draeger/Isolette TI500 | MH12470 | Service | Battery failure | Ready to use |  | Nasrul
80000118 | 2019-04-09 | RSUD Dr. Soeratno Gemolong | Mesin Anasthesia | Draeger/Fabius Plus | ASJM0264 | Survey | Layar blank | Ready to use |  | Amna
6000367 | 2019-04-09 | RSUD Panembahan Senopati Bantul | Mesin Anasthesia | Draeger/Primur | ASHJ0223 | Survey | Flow sensor inoperable | Ready to use |  | Afrizal
6000367 | 2019-04-09 | RSUD Panembahan Senopati Bantul | Mesin Anasthesia | Vapor 2000 Sevoflurane | ASHJ1282 | Service | Penggantian aksesoris | Ready to use |  | Afrizal
3000106 | 2019-04-08 | Klinic Aic Medical Imaging & Diagnostic | Ventilator | Medispec | 814008 | Service | Aplicator not detect | Ready to use |  | Imam
6000358 | 2019-04-08 | RSUP Dr. Sardjito | Ventilator | Draeger/Savina 300 | ARTK0043 | Service | Pemasangan Service set | Ready to use |  | Afrizal
7000095 | 2019-04-06 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASJH0125 | Service | Kalibrsi gas sensor failed | Ready to use |  | Ariawan
9000062 | 2019-04-05 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Babylog VN500 | ASJD0019 | Service | Kerusakan pada O2 | Ready to use |  | M. Rizky
3000108 | 2019-04-04 | RSUD Kab Tangerang | Mamografi | Metaltronica/Lilyum | ILILHFI/03911/CO | Survey | Perbaikan colimator | Ready to use |  | Imam
4000107 | 2019-04-04 | RSUD Kab. Bekasi | C-Arm | Apelem/Evo R+ | 55118026649 | Survey | Training User | Ready to use |  | Bayu
6000356 | 2019-04-04 | RSUD Tidar Magelang | Baby Incubator | Draeger/Isolette 8000 | HH08605 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-04-04
6000356 | 2019-04-04 | RSUD Tidar Magelang | Baby Incubator | Draeger/Isolette 8000 | HH08612 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-04-04
6000356 | 2019-04-04 | RSUD Tidar Magelang | Baby Incubator | Draeger/Isolette 8000 | NC08792 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-04-04
9000063 | 2019-04-02 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Evita V300 | ASLD0169 | Service | Pengecekan alat setelah Service | Ready to use |  | M. Rizky
80000117 | 2019-04-01 | RS Indriati Solo | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0036 | Service | Gas elgent keluar | Ready to use |  | Amna
9000064 | 2019-04-01 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Evita V300 | ASHL0175 | Maintenance | Selalu keluar merah saat kalibrasi | Ready to use |  | M. Rizky
5000105 | 2019-03-29 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Maintenance | Aplicator ED 1000 kurang air | Ready to use |  | Bilal
7000093 | 2019-03-29 | RSUD DR. R. Goeteng Taroenadibrata | Mesin Anasthesia | Draeger/Primus | ASJF0034 | Service | General Check | Ready to use |  | Ariawan
4000106 | 2019-03-29 | RSUD Provinsi NTB | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/032/C0 | Maintenance | Data gambar 2018 tidak ada di workstation | Ready to use |  | Bayu
7000091 | 2019-03-28 | RS Pala raya | Ventilator | Draeger/Savina 300 | ASKJ0001 | Service | General Check | Ready to use |  | Ariawan
80000116 | 2019-03-28 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASJF0033 | Service | Device is white | Ready to use |  | Amna
80000116 | 2019-03-28 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Service | Device is white | Ready to use |  | Amna
80000116 | 2019-03-28 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Service | Device is white | Ready to use |  | Amna
7000092 | 2019-03-28 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0027 | Service | Ventilator leak 300 | Ready to use |  | Ariawan
80000115 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
80000115 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
80000113 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
80000113 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0032 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
80000113 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
80000114 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
80000113 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
80000114 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0060 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
80000114 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0061 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
80000115 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0067 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
80000114 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0069 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
80000115 | 2019-03-27 | RSUD Ambarawa | Patient Monitor | Draeger/Vista 120 | VTSLK0070 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-27
11000081 | 2019-03-27 | RSUD Batara Guru | Patient Monitor | Draeger/Vista 120 | V2SHK0015 | Service | Check unit | Ready to use |  | Nasrul
11000081 | 2019-03-27 | RSUD Batara Guru | Patient Monitor | Draeger/Vista 120 | V2SHK0017 | Service | Check unit | Ready to use |  | Nasrul
11000081 | 2019-03-27 | RSUD Batara Guru | Patient Monitor | Draeger/Vista 120 | V2SHK0024 | Service | Check unit | Ready to use |  | Nasrul
11000081 | 2019-03-27 | RSUD Batara Guru | Patient Monitor | Draeger/Vista 120 | V3SHJ0014 | Service | Check unit | Ready to use |  | Nasrul
5000104 | 2019-03-26 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1086417 | Service | Handpiece tidak mengeluarkan air | Ready to use |  | Bilal
6000355 | 2019-03-26 | RSJ Prof Dr Soerojo Magelang | Incubator Transport | Draeger/Isolette TI500 | CH12771 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-03-26
6000354 | 2019-03-26 | RSJ Prof Dr Soerojo Magelang | Patient Monitor | Draeger/Vista 120 | VTSLJ0029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2020-03-26
11000080 | 2019-03-26 | RSUD Andi Djemma Masamba | Ventilator | Draeger/Savina 300 | ASHD0010 | Service | O2 sensor failed | Ready to use |  | Nasrul
11000079 | 2019-03-26 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | AY06913 | Service | Cek bakterial filter | Ready to use |  | Nasrul
11000079 | 2019-03-26 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | AY06914 | Service | Cek bakterial filter | Ready to use |  | Nasrul
11000079 | 2019-03-26 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | AY06915 | Service | Cek bakterial filter | Ready to use |  | Nasrul
11000079 | 2019-03-26 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | FR07841 | Service | Cek bakterial filter | Ready to use |  | Nasrul
11000079 | 2019-03-26 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | FR07842 | Service | Cek bakterial filter | Ready to use |  | Nasrul
6000353 | 2019-03-22 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLL0044 | Service | Rotary knob inoperable | Ready to use |  | Afrizal
7000089 | 2019-03-22 | RSUD Banyumas | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0075 | Service | Gas sensor failure | Ready to use |  | Ariawan
9000058 | 2019-03-21 | RSUD DR. SAM RATULANGI TONDANO | Ventilator | Draeger/Savina 300 | ASJM0392 | Service | Alarm battery internal discharge | Ready to use |  | M. Rizky
4000103 | 2019-03-21 | RSUD Provinsi NTB | Mobile X-ray | Apelem/Rafale EV30 | 009/18/01273 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2020-03-21
4000104 | 2019-03-21 | RSUD Provinsi NTB | Mobile X-ray | Apelem/Rafale EV30 | 009/18/01274 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2020-03-21
11000077 | 2019-03-20 | RSUD Kota Makassar | X-ray Digital | Apelem/Camargue DR | 84R481 | Service | Komputer tidak bisa ON | Ready to use |  | Nasrul
9000057 | 2019-03-20 | RSUD Maria Walanda Maramis | Ventilator | Draeger/Savina 300 | ASHL0291 | Maintenance | Kerusakan pada expiratory valve | Ready to use |  | M. Rizky
7000088 | 2019-03-19 | RS Prima Medika | Ventilator | Draeger/Savina 300 | ASHN0071 | Survey | General Check | Ready to use |  | Ariawan
6000352 | 2019-03-19 | RSUP dr. Soeradji Tirtonegoro | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0109 | Service | Kunjungan | Ready to use |  | Afrizal
7000087 | 2019-03-18 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASEF0030 | Maintenance | Air oksigen o2 failed | Ready to use |  | Ariawan
80000111 | 2019-03-15 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0458 | Maintenance | General Check | Ready to use |  | Amna
80000111 | 2019-03-15 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0478 | Maintenance | General Check | Ready to use |  | Amna
80000112 | 2019-03-15 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHH0122 | Maintenance | General Check | Ready to use |  | Amna
80000112 | 2019-03-15 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0179 | Maintenance | General Check | Ready to use |  | Amna
80000112 | 2019-03-15 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Primus | ASHJ0180 | Maintenance | General Check | Ready to use |  | Amna
80000112 | 2019-03-15 | RSUD Tugurejo | Mesin Anasthesia | Draeger/Presius A 500 | ASHL0016 | Service | General Check | Ready to use |  | Amna
9000059 | 2019-03-13 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Evita V300 | ASLD0169 | Service | Penggantian kokpid dan inspirasi valve | Ready to use |  | M. Rizky
11000075 | 2019-03-13 | RSUD Kota Makassar | Lampu Operasi | Draeger/Polaris 760 DC | 100000000123141 | Service | Mati total | Ready to use |  | Nasrul
9000060 | 2019-03-12 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJE0030 | Service | Test fungsi alat | Ready to use |  | M. Rizky
9000060 | 2019-03-12 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJE0031 | Service | Test fungsi alat | Ready to use |  | M. Rizky
9000060 | 2019-03-12 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJE0053 | Service | Test fungsi alat | Ready to use |  | M. Rizky
9000060 | 2019-03-12 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJE0059 | Service | Test fungsi alat | Ready to use |  | M. Rizky
9000060 | 2019-03-12 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJE0060 | Service | Test fungsi alat | Ready to use |  | M. Rizky
9000060 | 2019-03-12 | RSUD Ainun Habibie | Baby Incubator | Draeger/Caleo | ASJE0061 | Service | Test fungsi alat | Ready to use |  | M. Rizky
9000061 | 2019-03-12 | RSUD Ainun Habibie | Ventilator | Draeger/Evita V300 | ASLD0170 | Service | Alarm error | Ready to use |  | M. Rizky
7000086 | 2019-03-11 | RSUD DR. R. Goeteng Taroenadibrata | Ventilator | Draeger/Babylog VN500 | ASKL0081 | Service | pemasangan power supply | Ready to use |  | Ariawan
80000109 | 2019-03-11 | RSUD Dr. Soediran Mangun Sumarso | Ventilator | Draeger/Savina 300 | ARWL0203 | Service | Alarm int battery inop | Ready to use |  | Amna
80000110 | 2019-03-11 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASLL0147 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-03-11
11000074 | 2019-03-11 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Savina 300 | ASHD0008 | Service | Touchscreen rusak | Ready to use |  | Nasrul
80000108 | 2019-03-09 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Fabius 65 | ARUC 0012 | Service | Error gas supply | Ready to use |  | Amna
80000108 | 2019-03-09 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0066 | Service | Error gas supply | Ready to use |  | Amna
20000107 | 2019-03-08 | RSUP Dr. Kariadi | Resuscitaire | Draeger/RW82VHA | AF10367 | Service | Error 05 | Ready to use |  | Amna
80000106 | 2019-03-06 | RS Indriati Solo | Mesin Anasthesia | Draeger/Primus | ASJF0036 | Service | Vapor 120 tidak sesuai setting | Ready to use |  | Amna
80000106 | 2019-03-06 | RS Indriati Solo | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0063 | Maintenance | Vapor 120 tidak sesuai setting | Ready to use |  | Amna
80000105 | 2019-03-05 | RSUD Dr. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | V3SE0041 | Maintenance | Instalsi Service set | Ready to use |  | Amna
80000105 | 2019-03-05 | RSUD Dr. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | V3SEF0056 | Maintenance | Instalsi Service set | Ready to use |  | Amna
80000105 | 2019-03-05 | RSUD Dr. Soeprapto Cepu | Patient Monitor | Draeger/Vista 120 | V3SEF0064 | Service | Instalsi Service set | Ready to use |  | Amna
11000072 | 2019-03-05 | RSUD Kota Makassar | Incubator Transport | Draeger/Isolette TI500 | MY11117 | Service | Battery failure | Ready to use |  | Nasrul
11000072 | 2019-03-05 | RSUD Kota Makassar | Incubator Transport | Draeger/Isolette TI500 | MY11118 | Service | Battery failure | Ready to use |  | Nasrul
11000073 | 2019-03-05 | RSUP Dr. Wahidin Sudirohusodo | Ventilator | Draeger/Babylog VN500 | ASLA0082 | Service | Error system alarm failure | Ready to use |  | Nasrul
9000053 | 2019-03-02 | RSUP Prof. Dr. R. D. Kandou Manado | Bed Patient | PA64230C0513L | 90011-0510-0008 | Service | Kerusakan pada remote kontrol | Ready to use |  | M. Rizky
9000054 | 2019-03-02 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASHD0005 | Service | Safety valve error | Ready to use |  | M. Rizky
9000056 | 2019-03-02 | RSUP Prof. Dr. R. D. Kandou Manado | Ventilator | Draeger/Savina 300 | ASJD0179 | Service | Indicator batttery tidak keluar | Ready to use |  | M. Rizky
11000071 | 2019-03-01 | RSUP Dr. Wahidin Sudirohusodo | Lampu Operasi | Draeger/Polaris 760 DC | 100000000080970 | Service | Kamera error | Ready to use |  | Nasrul
9000052 | 2019-03-01 | RSUP Prof. Dr. R. D. Kandou Manado | Bed Patient | PA64230C0513L | 90011-0510-0008 | Service | Remote tidak berfungsi | Ready to use |  | M. Rizky
5000103 | 2019-02-28 | RSUD Tarakan | ESWL | Medispec/EM1000 | 0917012 | Service | Ada udara di dalam EMH | Ready to use |  | Bilal
3000102 | 2019-02-27 | RSUP Prof. Dr. R. D. Kandou Manado | Digital Radiografi Fluoroskopi | Apelem/Platinum | JI66F0184 | Service | Celling tidak berfungsi | Ready to use |  | Imam
3000103 | 2019-02-27 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 10131182 | Service | Inverter Fault | Ready to use |  | Imam
3000104 | 2019-02-27 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/E3000 | 7111298 | Survey | Tidak bisa mompa kuiras | Ready to use |  | Imam
3000105 | 2019-02-27 | RSUP Prof. Dr. R. D. Kandou Manado | ESWL | Medispec/EM1000 | 9111365 | Survey | Aplicator sudah lemah | Ready to use | Alat sudah digudangkan | Imam
7000084 | 2019-02-27 | RSUD Kardinah Tegal | Ventilator | Draeger/Evita V300 | ASHM0104 | Service | Error alarm failure | Ready to use |  | Ariawan
80000103 | 2019-02-27 | RSUP Dr. Kariadi | Mesin Anasthesia | Draeger/Fabius 65 | ARUC0012 | Survey | o2, N2O Air manometer error | Ready to use |  | Amna
5000102 | 2019-02-26 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 161046179 | Service | Selang pada handpiece mampet | Ready to use |  | Bilal
80000102 | 2019-02-26 | RSI Sunan Kudus | Ventilator | Draeger/Savina 300 | ARUE0022 | Service | Alarm no internal battery | Ready to use |  | Amna
80000102 | 2019-02-26 | RSI Sunan Kudus | Ventilator | Draeger/Savina 300 | ARUM0308 | Service | Alarm no internal battery | Ready to use |  | Amna
2000103 | 2019-02-26 | RSUD Kab. Temanggung | ESWL | Medispec/EM1000 | 1115056 | Pra Instalasi | CPU C-Arm tidak bisa menyala | Ready to use |  | Eddy
3000101 | 2019-02-25 | RSUD Kab. Bolaang Mongondow | X-ray Digital | Apelem/Camargue DR | CPD30503B18 | Service | Instalasi work station dan detector | Ready to use |  | Imam | 2021-02-25
80000101 | 2019-02-25 | RSUP Dr. Kariadi | Baby Incubator | Draeger/Caleo | ASJM0048 | Maintenance | Devicce error 31 | Ready to use |  | Amna
4000102 | 2019-02-23 | RSUD Dr. M. M. Dunda | C-Arm | Apelem/Evo R+ | 55118013636 | Survey | Pelaksanaan uji kesesuaian dan uji paparan | Ready to use |  | Bayu
9000051 | 2019-02-23 | RSUD Dr. M. M. Dunda | C-Arm | Apelem/Evo R+ | C5128013636 | Survey | Pelaksanaan uji kesesuaian dan uji paparan | Ready to use |  | Nasrul
4000101 | 2019-02-22 | RSUD Kab. Bolaang Mongondow | X-ray Digital | Apelem/Camargue DR | CPD30503B18 | Survey | Pelaksanaan uji kesesuaian dan perapatan kabel | Ready to use |  | Bayu
6000359B | 2019-02-22 | RSUP Dr. Sardjito | Jaundice Meter | Draeger/JM 105 | B3502776 | Service | Out of calibrasi | Ready to use |  | Afrizal
7000083 | 2019-02-21 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0049 | Service | Semua voice error | Ready to use | Disarankan penggrebesan module control electronic mainboard, valve block brake arm, gas connector | Ariawan
7000082 | 2019-02-20 | RSUD DR. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CCI03924 | Service | Sensor modul failure | Ready to use |  | Ariawan
11000070 | 2019-02-20 | RSUD Tenriawaru Bone | X-ray Digital | Apelem/Camargue DR | CPD26084L16 | Service | Printer error | Ready to use |  | Nasrul
11000069 | 2019-02-18 | RSUD SYEKH YUSUP KAB. GOWA | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0024 | Service | O2 sensor failed | Ready to use |  | Nasrul
11000069 | 2019-02-18 | RSUD SYEKH YUSUP KAB. GOWA | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0028 | Service | O2 sensor failed | Ready to use |  | Nasrul
11000069 | 2019-02-18 | RSUD SYEKH YUSUP KAB. GOWA | Mesin Anasthesia | Draeger/Primus | ASHL0057 | Survey | O2 sensor failed | Ready to use |  | Nasrul
2000102 | 2019-02-15 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D156F0271 | Service | Instalasi user | Ready to use |  | Eddy
7000081 | 2019-02-15 | RSUI Harapan Anda | Ventilator | Draeger/Savina 300 | ASHA0095 | Service | Leak sistem ventilator fail | Ready to use |  | Ariawan
2000101 | 2019-02-14 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | I176F0243 | Survey | Perbaikan sensor grid | Ready to use |  | Eddy
9000050 | 2019-02-13 | RSUD Tani dan Nelayan | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0052 | Service | Po2 bermasalah | Ready to use |  | Nasrul
9000048 | 2019-02-12 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/V 300 | ASLD0169 | Service | Touchscreen tidak berfungsi | Ready to use |  | Nasrul
5000101 | 2019-02-12 | RSUD Sultan Syarif Al-Idrus | X-Ray Analog | Apelem/Camargue Analog | CPD30841D18 | Pra Instalasi | Belum ada listrik lanjutan Service report bayu | Ready to use |  | Bilal
11000067 | 2019-02-12 | RSUD Dr. M. M. Dunda | Ventilator | Draeger/Evita V300 | ASLD0169 | Pra Instalasi | Touchscreen tidak berfungsi | Ready to use |  | Nasrul
9000049 | 2019-02-12 | RSUD Toto Kabila | Mesin Anasthesia | Draeger/Fabius Plus | ASKK0057 | Service | O2 sensor failed | Ready to use |  | Nasrul
7000080 | 2019-02-11 | RSUD DR. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CCI03923 | Service | Sensor modul failure | Ready to use |  | Ariawan
12000095 | 2019-02-08 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASFL0121 | Service | System check o2 sensor fail | Ready to use |  | Amna
12000096 | 2019-02-08 | RSUD Banyumas | Incubator Transport | Draeger/Isolette TI500 | KF10743 | Survey | Battery error | Ready to use |  | Amna
11000060B | 2019-02-07 | RSUP Tadjudin Chalid | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASJK0068 | Service | Tabung oxygen tidak ada | Ready to use |  | Nasrul
11000066 | 2019-02-07 | RSUP Tadjudin Chalid | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASJK0068 | Service | Tabung oxygen tidak ada | Ready to use |  | Nasrul
12000094 | 2019-02-06 | RS Dr. Oen Solo Baru | Resuscitaire | Draeger/RW82VHA | CY11241 | Survey | Error 05 | Ready to use |  | Amna
6000046 | 2019-02-02 | RSUP dr. Soeradji Tirtonegoro | Resuscitaire | Draeger/Resuscitaire | KH19232 | Maintenance | Training User | Ready to use |  | Afrizal
12000092 | 2019-01-29 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Primus | ASHD0034 | Service | General Check | Ready to use |  | Amna
6000044 | 2019-01-28 | RSUD Dr. Tjitrowardojo Purworejo | Ventilator | Draeger/Savina 300 | ASBM0243 | Service | Check software dan hardware | Ready to use |  | Afrizal
11000064 | 2019-01-27 | RSUD Djafar Harun | X-ray Digital | Apelem/Camargue DR | 845889 | Service | Meja pasien tidak dapat bekerja dan hasil foto tidak muncul; Perbaikan driver board | Ready to use |  | Nasrul
12000090 | 2019-01-25 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002705 | Service | Padle bocor | Ready to use |  | Amna
12000091 | 2019-01-25 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002719 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2020-01-25
12000086 | 2019-01-24 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002721 | Service | Padle bocor | Ready to use |  | Amna
12000088 | 2019-01-24 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002721 | Service | Padle bocor | Ready to use |  | Amna
12000086 | 2019-01-24 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002722 | Service | General Check | Ready to use |  | Amna
12000089 | 2019-01-24 | RSUP Dr. Kariadi | Dental Unit | Kavo/1058TM | 1002722 | Service | General Check | Ready to use |  | Amna
12000085 | 2019-01-23 | RSUD RAA Soewondo Pati | Baby Incubator | Draeger/Isolette 8000 | HF05849 | Service | General Check | Ready to use |  | Amna
12000085 | 2019-01-23 | RSUD RAA Soewondo Pati | Baby Incubator | Draeger/Isolette 8000 | HF05850 | Service | General Check | Ready to use |  | Amna
12000085 | 2019-01-23 | RSUD RAA Soewondo Pati | Baby Incubator | Draeger/Isolette 8000 | HF05851 | Service | General Check | Ready to use |  | Amna
12000085 | 2019-01-23 | RSUD RAA Soewondo Pati | Baby Incubator | Draeger/Isolette 8000 | HF05855 | Maintenance | General Check | Ready to use |  | Amna
3000038 | 2019-01-22 | RSUP Dr. Hasan Sadikin | X-ray Digital | Apelem EOL | CPD177971214 | Service | Tidak bisa exspose | Ready to use |  | Imam
7000078 | 2019-01-22 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/fabius Tiro | ASBU0248 | Maintenance | General Check | Ready to use |  | Ariawan
7000078 | 2019-01-22 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0025 | Maintenance | General Check | Ready to use |  | Ariawan
7000078 | 2019-01-22 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0027 | Maintenance | General Check | Ready to use |  | Ariawan
7000078 | 2019-01-22 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0039 | Maintenance | General Check | Ready to use |  | Ariawan
7000078 | 2019-01-22 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0047 | Survey | General Check | Ready to use |  | Ariawan
6000043 | 2019-01-21 | RSUD Tidar Magelang | ESWL | Medispec/E3000 | 0317018 | Survey | Filter kotor | Ready to use |  | Afrizal
3000037 | 2019-01-18 | RSUP Dr. Mohammad Hoesin | Bone Mineral Densitometer | Medilink/Medix 90 | L11015D248 | Service | X-Ray Crash | Ready to use |  | Imam
11000065 | 2019-01-17 | RSUD Toto Kabila | Lampu Operasi | Draeger/Polaris 600 | ASKM0658 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000065 | 2019-01-17 | RSUD Toto Kabila | Lampu Operasi | Draeger/Polaris 600 | ASKM0659 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
7000077 | 2019-01-16 | RSUD Hj. Anna Lasmanah | Resuscitaire | Draeger/Resuscitaire | KF10740 | Service | Heater mati total | Ready to use |  | Ariawan
3000036 | 2019-01-14 | RSUD Tualang | X-ray Digital | Apelem/Camargue DR | CPD30502813 | Service | Pemindahan alat DR | Ready to use |  | Imam | 2021-01-14
12000083 | 2019-01-11 | RS Bhina Bhakti Husada | Phototherapy | Draeger/PT 4000 | ASKL0006 | Service | While light off | Ready to use |  | Amna
2000029 | 2019-01-11 | RSUD KRT. Setjonegoro Wonosobo | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/044/C0 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2021-01-11
2000030 | 2019-01-10 | RSUD KRT. Setjonegoro Wonosobo | Digital Radiografi Fluoroskopi | Apelem/Platinum | K126F0302 | Survey | Training User | Ready to use |  | Eddy | 2020-01-10
2000028 | 2019-01-09 | RSUD Kab. Temanggung | C-Arm | Apelem/Evo R+ | 55115011469 | Service | C-Arm pada unit tidak bisa maju mundur | Ready to use |  | Eddy
7000074 | 2019-01-08 | RSUD Dr. Moewardi | Mesin Anasthesia | Draeger/Fabius Plus | ASEE0707 | Survey | General Check | Ready to use |  | Ariawan
2000027 | 2019-01-08 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Survey | Training User | Ready to use |  | Eddy | 2021-01-08
12000082 | 2019-01-07 | Siloam Hospital Semarang | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASKL0038 | Maintenance | Device Check | Ready to use |  | Amna
12000082 | 2019-01-07 | Siloam Hospital Semarang | Ventilator | Draeger/Savina 300 | ASKN0288 | Maintenance | Device Check | Ready to use |  | Amna
12000082 | 2019-01-07 | Siloam Hospital Semarang | Incubator Transport | Draeger/Isolette TI500 | CR12070 | Maintenance | Device Check | Ready to use |  | Amna
12000082 | 2019-01-07 | Siloam Hospital Semarang | Baby Incubator | Draeger/Isolette C2000 | CR33565 | Service | Device Check | Ready to use |  | Amna
12000081 | 2019-01-05 | RSUD RAA Soewondo Pati | Baby Incubator | Draeger/Isolette 8000 | HF05855 | Service | Controller Failure 4 | Ready to use |  | Amna
6000349 | 2019-01-04 | RSA UGM | Mesin Anasthesia | Draeger/Fabius Plus | ASEM0135 | Service | APL tidak sesuai setting | Ready to use |  | Afrizal
7000079 | 2019-01-04 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0056 | Service | Leak sistem lebih dari 300L/menit | Ready to use |  | Ariawan
7000079 | 2019-01-04 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Fabius Plus | ASHK0058 | Maintenance | Cleaning cpsy dan piston | Ready to use |  | Ariawan
7000073 | 2019-01-03 | RSUD Banyumas | Ventilator | Draeger/Savina 300 | ARXL0174 | Maintenance | General Check | Ready to use |  | Ariawan
7000073 | 2019-01-03 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASKK0230 | Survey | General Check | Ready to use |  | Ariawan
12000080 | 2019-01-02 | RSUD Kab Batang | Baby Incubator | Draeger/Isolette C2000 | HH36864 | Service | Alarm skin temperature low | Ready to use |  | Amna
11000063 | 2018-12-28 | RSUD Djafar Harun | X-ray Digital | Apelem/Camargue DR | 45889 | Survey | Meja pasien tidak dapat digunakan | Ready to use |  | Nasrul
7000072 | 2018-12-27 | RS Bhakti Asih Brebes | Ventilator | Draeger/Evita V300 | ASKD0011 | Maintenance | General Check | Ready to use |  | Ariawan
7000071 | 2018-12-26 | RS. MITRA SIAGA - TEGAL | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0183 | Maintenance | General Check | Ready to use |  | Ariawan
3000035 | 2018-12-22 | RSUD KOTAMOBAGU | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/043/C0 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2019-12-22
2000024 | 2018-12-22 | RSUD KOTAMOBAGU | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0270 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2019-12-22
11000062 | 2018-12-22 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 9111371 | Service | Cek EMH | Ready to use |  | Nasrul
6000042 | 2018-12-21 | RS PKU Muhammadiyah wonosobo | Ventilator | Draeger/Savina 300 | ARVA0189 | Service | Batterai drop, Vte tidak stabil, 02 sensor kadaluarsa | Ready to use |  | Afrizal
6000041 | 2018-12-20 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Primus | ASJE0011 | Survey | Gas sensor fail | Ready to use |  | Afrizal
6000041 | 2018-12-20 | RSUD Kab. Temanggung | Mesin Anasthesia | Draeger/Primus | ASJE0011 | Service | cek software dan hardware | Ready to use |  | Afrizal
11000061 | 2018-12-20 | RSUD Trikora Jalakan Kab Banggi | Mesin Anasthesia | Draeger/Primus | ASDH0036 | Service | ventilator failed | Ready to use |  | Nasrul
7000070 | 2018-12-19 | RSUD PREMBUN KAB. KEBUMEN | Incubator Transport | Draeger/Isolette TI500 | HH12657 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-12-19
3000034 | 2018-12-18 | RSUD Kab Tangerang | Mamografi | Metaltronica/Lilyum | ILILHFI/03911/CO | Service | Lampu colimator mati | Ready to use |  | Imam
6000039 | 2018-12-17 | RS Peratama Yogyakarta | Patient Monitor | Draeger/Vista 120 | VTSKL0027 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-12-17
6000040 | 2018-12-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0007 | Service | Paw sering terlalu tinggi | Ready to use |  | Afrizal
3000033 | 2018-12-14 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Software komputer not running | Ready to use |  | Imam
4000023 | 2018-12-14 | RSUD KRT. Setjonegoro Wonosobo | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0302 | Survey | Pelaksanaan uji kesesuaian dan uji paparan | Ready to use |  | Bayu | 2019-12-14
2000025 | 2018-12-13 | RSUD KRT. Setjonegoro Wonosobo | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/044/C0 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2019-12-13
12000079 | 2018-12-11 | RSUD Dr. R. Soedjati Soemodiarjo | Lampu Operasi | Draeger/Polaris 600 | ASKM0657 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
6000034 | 2018-12-10 | RSUD KRT. Setjonegoro Wonosobo | Pendant | Draeger/Movita | ASKM0003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-12-10
6000034 | 2018-12-10 | RSUD KRT. Setjonegoro Wonosobo | Pendant | Draeger/Movita | ASKM0004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-12-10
6000034 | 2018-12-10 | RSUD KRT. Setjonegoro Wonosobo | Pendant | Draeger/Movita | ASLL0157 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-12-10
6000034 | 2018-12-10 | RSUD KRT. Setjonegoro Wonosobo | Pendant | Draeger/Movita | ASLL0158 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-12-10
6000037 | 2018-12-10 | RSUD KRT. Setjonegoro Wonosobo | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/044/C0 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-12-10
6000038 | 2018-12-10 | RSUD KRT. Setjonegoro Wonosobo | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0302 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-12-10
12000077 | 2018-12-07 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Service | LCD blanck | Ready to use |  | Amna
12000078 | 2018-12-05 | RSUD Dr. R. Soedjati Soemodiarjo | Lampu Operasi | Draeger/Polaris 600 | ASKM0657 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2019-12-05
4000022 | 2018-12-05 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Survey | Pelaksanaan uji kesesuaian dan uji paparan | Ready to use |  | Bayu | 2019-12-05
5000014 | 2018-12-03 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 1086417 | Service | Meja mengeluarkan bunyi bising | Ready to use |  | Bilal
5000013 | 2018-12-03 | Klinik Smile Works | Dental Unit | Kavo/Systematica 1060 T | 161046179 | Service | Pedal kaki error | Ready to use |  | Bilal
2000022 | 2018-12-03 | RSUD Kab. Temanggung | Digital Radiografi Fluoroskopi | Apelem/Platinum | K186F0300 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2019-12-03
11000057 | 2018-12-01 | RSU Undata Palu | Emergency Stretcher | Paramount/PK 2100 | 9009095-00064 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000057 | 2018-12-01 | RSU Undata Palu | Emergency Stretcher | Paramount/PK 2100 | 9009095-00065 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000059 | 2018-12-01 | RSU Undata Palu | Box Baby | Paramount/PB 1100B | 9011028-00007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000059 | 2018-12-01 | RSU Undata Palu | Box Baby | Paramount/PB 1100B | 9011028-00008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000059 | 2018-12-01 | RSU Undata Palu | Box Baby | Paramount/PB 1100B | 9011028-00009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000059 | 2018-12-01 | RSU Undata Palu | Box Baby | Paramount/PB 1100B | 9011028-00011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000059 | 2018-12-01 | RSU Undata Palu | Box Baby | Paramount/PB 1100B | 9011028-00019 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000059 | 2018-12-01 | RSU Undata Palu | Box Baby | Paramount/PB 1100B | 901128-00012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000059 | 2018-12-01 | RSU Undata Palu | Box Baby | Paramount/PB 1100B | 901128-00017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000059 | 2018-12-01 | RSU Undata Palu | Box Baby | Paramount/PB 1100B | 901128-0018 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
5000012 | 2018-11-30 | RS Marinir Cilandak | X-Ray Analog | Apelem/Camargue Analog | CPD26753A17 | Service | Meja tidak bisa ngebreak karna magnet pada meja kendor; Magnet sudah dikencangkan | Ready to use |  | Bilal
5000011 | 2018-11-30 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Service | Aplikator tidak menembak normal | Ready to use |  | Bilal
5000010 | 2018-11-30 | RS Siloam Asri | ESWT | Medispec/ED1000 | 6111274 | Service | Aplikator kempes | Ready to use |  | Bilal
11000058 | 2018-11-30 | RSU Undata Palu | Ventilator | Draeger/Isolette 8000 | HH08597 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-11-30
5000009 | 2018-11-29 | Apartemen setiabudi | ESWT | Medispec/ED1000 | 1116125 | Service | Aplikator tidak mau menembak | Ready to use |  | Bilal
7000069 | 2018-11-27 | RSUD DR. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | AH36595 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-11-27
6000035 | 2018-11-27 | RSUD KRT. Setjonegoro Wonosobo | Patient Monitor | Draeger/Infinity Vista XL | 6007858268 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-11-27
6000035 | 2018-11-27 | RSUD KRT. Setjonegoro Wonosobo | Patient Monitor | Draeger/Infinity Vista XL | 6008274673 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-11-27
6000035 | 2018-11-27 | RSUD KRT. Setjonegoro Wonosobo | Patient Monitor | Draeger/Infinity Vista XL | 6008274771 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-11-27
6000035 | 2018-11-27 | RSUD KRT. Setjonegoro Wonosobo | Patient Monitor | Draeger/Infinity Vista XL | 6008278072 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-11-27
11000056 | 2018-11-24 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | AY06913 | Service | Humidity Low | Ready to use |  | Nasrul
11000056 | 2018-11-24 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | AY06914 | Service | Humidity Low | Ready to use |  | Nasrul
11000056 | 2018-11-24 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | AY06915 | Service | Humidity Low | Ready to use |  | Nasrul
11000056 | 2018-11-24 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | FR07841 | Service | Humidity Low | Ready to use |  | Nasrul
11000056 | 2018-11-24 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | FR07842 | Service | Humidity Low | Ready to use |  | Nasrul
3000032 | 2018-11-22 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G176F0278 | Service | Monitor flouroscopy mati | Ready to use |  | Imam
4000020 | 2018-11-22 | RSUD Balaraja | C-Arm | Apelem/Evo R+ | 55118034657 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2019-11-22
5000006 | 2018-11-22 | RSUD KOTAMOBAGU | C-Arm | Apelem/Evo R+ | 55118031654 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal | 2019-11-22
7000068 | 2018-11-21 | RSUD DR. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CU03924 | Service | Unit mati total dan arm pintu patah | Ready to use |  | Ariawan
12000074 | 2018-11-19 | RSUD RA Kartini | Ventilator | Draeger/Evita V300 | ASHD0187 | Service | System check nebulixer fail | Ready to use |  | Amna
7000066 | 2018-11-19 | RSUD Banyumas | Lampu Operasi | Draeger/Polaris 760 DC | 1000-0000-0122-490 | Service | kabel RJ.45 | Ready to use |  | Ariawan
7000066 | 2018-11-19 | RSUD Banyumas | Lampu Operasi | Draeger/Polaris 760 DC | 1000-0000-0123-586 | Service | kabel Power | Ready to use |  | Ariawan
3000031 | 2018-11-16 | RSD Mayjen HM Ryacudu | X-ray Digital | Apelem/Camargue DR | IMA 703939 | Survey | Pelaksanaan Uji Kesesuaian | Ready to use |  | Imam | 2020-11-16
4000019 | 2018-11-16 | RSUD KOTA TARAKAN | Mamografi | Metaltronica/Lilyum | 1CILHF/0709/CB2 | Survey | Pelaksanaan Uji Paparan & Uji Kesesuaian | Ready to use |  | Bayu
6000033 | 2018-11-16 | RSUD Sleman | Jaundice Meter | Draeger/JM 105 | B3503649 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-11-16
2000021 | 2018-11-12 | RSUD Ratu Zalecha | Digital Radiografi Fluoroskopi | Apelem/Platinum | D186F0271 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2019-11-12
12000073 | 2018-11-09 | RSUD Kab Batang | Baby Incubator | Draeger/Isolette C2000 | HH 36862 | Survey | Training User | Ready to use |  | Amna
12000073 | 2018-11-09 | RSUD Kab Batang | Baby Incubator | Draeger/Isolette C2000 | HH 36863 | Survey | Training User | Ready to use |  | Amna
12000073 | 2018-11-09 | RSUD Kab Batang | Baby Incubator | Draeger/Isolette C2000 | HH 36864 | Service | Training User | Ready to use |  | Amna
12000072 | 2018-11-08 | RSUP Dr. Sardjito | Central Monitor | Draeger/Vista 120 CMS | 107XL00283 | Service | Bed 8 dan 10 pada central OFF | Ready to use |  | Amna
12000071 | 2018-11-07 | RS Bhina Bhakti Husada | Central Monitor | Draeger/Vista 120 CMS | VLSKM 0058 | Service | Number of pending patient | Ready to use |  | Amna
12000070 | 2018-11-06 | RSUD Kab Batang | Baby Incubator | Draeger/Isolette C2000 | HH 36862 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2019-11-06
12000070 | 2018-11-06 | RSUD Kab Batang | Baby Incubator | Draeger/Isolette C2000 | HH 36863 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2019-11-06
12000070 | 2018-11-06 | RSUD Kab Batang | Baby Incubator | Draeger/Isolette C2000 | HH 36864 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2019-11-06
2000020 | 2018-11-06 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Service | Moving X-Ray berat | Ready to use |  | Eddy
5000007 | 2018-11-05 | RSUD Tani dan Nelayan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/018/C0 | Service | Pedal press tidak bisa dijalankan | Ready to use |  | Bilal
11000060 | 2018-11-04 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 9111375 | Service | Touchscreen tidak berfungsi | Ready to use |  | Nasrul
12000069 | 2018-11-02 | RSUD Kajen | Resuscitaire | Draeger/Resuscitaire | MU22319 | Service | Kit corner blocks patah | Ready to use |  | Amna
3000030 | 2018-10-31 | Laboratorium Klinik Pramita | Bone Mineral Densitometer | DMS Apelem | H11015D245 | Maintenance | software crush | Ready to use |  | Imam
12000067 | 2018-10-30 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 Edition | ASEL0142 | Maintenance | General Check | Ready to use |  | Amna
12000067 | 2018-10-30 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 Edition | ASEL0143 | Maintenance | General Check | Ready to use |  | Amna
12000068 | 2018-10-30 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASJF0033 | Maintenance | Device Check | Ready to use |  | Amna
12000068 | 2018-10-30 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Maintenance | Device Check | Ready to use |  | Amna
12000068 | 2018-10-30 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Maintenance | General Check | Ready to use |  | Amna
2000019 | 2018-10-29 | RSD Mayjen HM Ryacudu | X-ray Digital | Apelem/Camargue DR | IMA703939 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2019-10-29
7000063 | 2018-10-27 | RSUD Provinsi NTB | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/032/C0 | Service | Pedal pada kompresor tidak bisa naik | Ready to use |  | Ariawan
3000029 | 2018-10-26 | RSUD Malingping | X-ray Digital | Apelem/Camargue DR | IMA 703916 | Service | Workstation tidak bisa print | Ready to use |  | Imam
3000027 | 2018-10-25 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | penggantian hardisk Workstation | Ready to use |  | Imam
3000028 | 2018-10-25 | RSUP Dr. Mohammad Hoesin (KSO) | MRI | Nonin/7500 RO | 50239008 | Maintenance | SPO2 sensor not connect | Ready to use |  | Imam
7000062 | 2018-10-25 | RSKBD Panti Nugroho | Ventilator | Draeger/Savina 300 | ASLC0399 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-25
6000031 | 2018-10-25 | RSUD Bagas Waras Klaten | Mesin Anasthesia | Draeger/Perseus A500 | ASJD006 | Service | tukar flow sensor inspirasi dengan flow sensor ekspirasi | Ready to use |  | Afrizal
6000031 | 2018-10-25 | RSUD Bagas Waras Klaten | Mesin Anasthesia | Draeger/Perseus A500 | ASJD0067 | Service | VTE kadang tidak tercapai | Ready to use |  | Afrizal
12000066 | 2018-10-25 | RSUD Kab Batang | Incubator Transport | Draeger/Isolette TI500 | CF10703 | Survey | Roda berat digunakan | Ready to use |  | Amna
12000065 | 2018-10-24 | RS Bhina Bhakti Husada | Central Monitor | Draeger/Vista 120 CMS | VLSKM 0058 | Service | Setting tanggal tidak sesuai | Ready to use |  | Amna
7000061 | 2018-10-24 | RSUD Hj. Anna Lasmanah | Baby Incubator | Draeger/Isolette 8000 | CF06019 | Service | Humidity error check water supply | Ready to use |  | Ariawan
7000061 | 2018-10-24 | RSUD Hj. Anna Lasmanah | Baby Incubator | Draeger/Isolette 8000 | CF06020 | Service | cleaning pada bagian chamber | Ready to use |  | Ariawan
6000029 | 2018-10-24 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASAM0007 | Service | leak system, manual baging terasa berat | Ready to use |  | Afrizal
6000030 | 2018-10-24 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0076 | Maintenance | pemasangan claim warranty monitor touch screeen | Ready to use |  | Afrizal
12000063 | 2018-10-23 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASEH 0118 | Service | Penggantian Service set | Ready to use |  | Amna
7000060 | 2018-10-23 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0049 | Service | Pemindahan stepper head arm pendant | Ready to use | Disarankan penggrebesan module control electronic mainboard, valve block brake arm, gas connector | Ariawan
12000062 | 2018-10-22 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL 0012 | Maintenance | General Check | Ready to use |  | Amna
12000062 | 2018-10-22 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL 0013 | Service | Device Check | Ready to use |  | Amna
7000059 | 2018-10-22 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0001 | Survey | Air supply tidak berfungsi | Ready to use | Disarankan penggrebesan module control electronic mainboard, valve block brake arm, gas connector | Ariawan
6000028 | 2018-10-19 | RS Tidar Budi Rahayu | Baby Incubator | Draeger/Caleo | ASJM 0058 | Survey | cek unit OK | Ready to use |  | Afrizal
6000028 | 2018-10-19 | RS Tidar Budi Rahayu | Baby Incubator | Draeger/Caleo | ASJM0061 | Service | pemasangan claim warranty intermediate element | Ready to use |  | Afrizal
3000024 | 2018-10-18 | RSUD Kardinah Tegal | ESWL | Medispec/EM1000 | 3111118 | Service | connector aplicator rusak | Ready to use |  | Imam
3000025 | 2018-10-18 | RSUD Kardinah Tegal | ESWL | Medispec/EM1000 | 109117 | Service | Tembakan tidak maksimal | Ready to use |  | Imam
6000026 | 2018-10-18 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASHM0007 | Service | jacksoon reesee ada bocor pada katub | Ready to use |  | Afrizal
4000018 | 2018-10-17 | RSUD Bahteramas | C-Arm | Apelem/Evo R+ | 55118027650 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2019-10-17
7000058 | 2018-10-17 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0027 | Service | Leak pada ventilator sehingga tabung volume tidak tercapai | Ready to use |  | Ariawan
12000061 | 2018-10-16 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL 0012 | Maintenance | General Check | Ready to use |  | Amna
12000061 | 2018-10-16 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL 0013 | Service | Device Check | Ready to use |  | Amna
6000027 | 2018-10-16 | RS Panti Rapih | Mesin Anasthesia | Draeger/Fabius Plus | ASDA0088 | Service | cek software dan hardware | Ready to use |  | Afrizal
6000089 | 2018-10-16 | RS Panti Rapih | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0193 | Survey | pemasangan maintence co27 2 tahun | Ready to use |  | Afrizal
6000089 | 2018-10-16 | RS Panti Rapih | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0193 | Survey | cek software dan hardware | Ready to use |  | Afrizal
2000018 | 2018-10-16 | RSUD Tani dan Nelayan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/018/C0 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2020-10-16
7000055 | 2018-10-16 | RSUD Brebes | Pendant | Draeger/Movita | ASKL0018 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-16
7000055 | 2018-10-16 | RSUD Brebes | Pendant | Draeger/Movita | ASKL0172 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-16
7000055 | 2018-10-16 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-16
7000055 | 2018-10-16 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-16
7000055 | 2018-10-16 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0019 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-16
7000055 | 2018-10-16 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-16
7000055 | 2018-10-16 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0049 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-16
7000055 | 2018-10-16 | RSUD Brebes | Pendant | Draeger/Movita | ASKM0050 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-16
7000055 | 2018-10-16 | RSUD Brebes | Pendant | Draeger/Movita | ASLJ0545 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-16
6000023 | 2018-10-15 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHA0113 | Service | Setting vte tidak sesuai | Ready to use |  | Afrizal
6000023 | 2018-10-15 | RSUD Kab. Temanggung | Ventilator | Draeger/Evita V300 | ASHN0038 | Service | backup in inspiratory valve | Ready to use |  | Afrizal
12000060 | 2018-10-15 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0093 | Maintenance | Device Check | Ready to use |  | Amna
12000060 | 2018-10-15 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0094 | Maintenance | General Check | Ready to use |  | Amna
6000025 | 2018-10-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-10-15
6000025 | 2018-10-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASLD0043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-10-15
7000057 | 2018-10-15 | RSUD PREMBUN KAB. KEBUMEN | Resuscitaire | Draeger/Resuscitaire | HR12031 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-15
11000052 | 2018-10-12 | RSU Fatimah Pare-Pare | Bed Patient | Paramount BED/PA-99295 | 9005121-00467 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000052 | 2018-10-12 | RSU Fatimah Pare-Pare | Bed Patient | Paramount BED/PA-99295 | 9005121-00468 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000052 | 2018-10-12 | RSU Fatimah Pare-Pare | Bed Patient | Paramount BED/PA-99295 | 9005121-00497 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000052 | 2018-10-12 | RSU Fatimah Pare-Pare | Bed Patient | Paramount BED/PA-99295 | 9005121-00542 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000052 | 2018-10-12 | RSU Fatimah Pare-Pare | Bed Patient | Paramount BED/PA-99295 | 9005121-00585 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000052 | 2018-10-12 | RSU Fatimah Pare-Pare | Bed Patient | Paramount BED/PA-99295 | 9005121-00622 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000052 | 2018-10-12 | RSU Fatimah Pare-Pare | Bed Patient | Paramount BED/PA-99295 | 9005121-00623 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000052 | 2018-10-12 | RSU Fatimah Pare-Pare | Bed Patient | Paramount BED/PA-99295 | 9005121-00648 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000052 | 2018-10-12 | RSU Fatimah Pare-Pare | Bed Patient | Paramount BED/PA-99295 | 9005121-00655 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000052 | 2018-10-12 | RSU Fatimah Pare-Pare | Bed Patient | Paramount BED/PA-99295 | 9005121-00998 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
2000017 | 2018-10-11 | RSUD Ratu Zalecha | Panoramic | Panoramic/CS9000C/Carestream | DFB1004 | Service | pemindahan unit keruang baru | Ready to use |  | Eddy
12000059 | 2018-10-10 | RSUD RA Kartini | Medical Compressor | Draeger/Medical Compressor | ASKL0084 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2019-10-10
11000051 | 2018-10-10 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASAM 0044 | Service | VR tidak sesuai setting | Ready to use |  | Nasrul
11000051 | 2018-10-10 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASAM 0045 | Service | VR tidak sesuai setting | Ready to use |  | Nasrul
7000056 | 2018-10-08 | RST Wijayakusuma | Ventilator | Draeger/Savina 300 | ASEL0196 | Service | Battery kurang dari 30% | Ready to use |  | Ariawan
12000058 | 2018-10-08 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0108 | Service | PS Heatre Off | Ready to use |  | Amna
11000050 | 2018-10-08 | RSUP Dr. Wahidin Sudirohusodo | Incubator Transport | Draeger/Isolette TI500 | KG 0144 | Service | Incubator battery fail | Ready to use |  | Nasrul
11000050 | 2018-10-08 | RSUP Dr. Wahidin Sudirohusodo | Incubator Transport | Draeger/Isolette TI500 | KG 08143 | Service | Incubator battery fail | Ready to use |  | Nasrul
3000023 | 2018-10-05 | RSUD Kab. Temanggung | C-Arm | Apelem/Evo R+ | 55115011469 | Service | unit blank | Ready to use |  | Imam
3000022 | 2018-10-04 | RSUD Kab. Temanggung | ESWL | Medispec/EM1000 | 1115056 | Service | tanggal tidak sesuai | Ready to use |  | Imam
6000022 | 2018-10-04 | RSA UGM | Incubator Transport | Draeger/Isolette TI500 | AH12604 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-10-04
5000005 | 2018-10-03 | Klinic Aic Medical Imaging & Diagnostic | EECP | Lumenair/Vasomedical | 5TC1127 | Survey | Alat tidak berfungsi dengan baik | Ready to use |  | Bilal
11000049B | 2018-10-02 | RSUD Andi Makassau | Jaundice Meter | Draeger/JM 105 | B3503453 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-10-02
12000056 | 2018-10-02 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0108 | Service | PS Heatre Off | Ready to use |  | Amna
7000053 | 2018-10-01 | RSUD PREMBUN KAB. KEBUMEN | Ventilator | Draeger/Babylog VN500 | ASCF0056 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-01
7000053 | 2018-10-01 | RSUD PREMBUN KAB. KEBUMEN | Ventilator | Draeger/Babylog VN500 | ASCF0057 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-01
7000053 | 2018-10-01 | RSUD PREMBUN KAB. KEBUMEN | Ventilator | Draeger/Savina 300 | ASCF0163 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-10-01
3000021 | 2018-09-27 | RSUD Kab. Bolaang Mongondow | X-ray Digital | Apelem/Camargue DR | CPD30503B18 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2019-09-27
7000052 | 2018-09-27 | RSUD Cilacap | Baby Incubator | Draeger/Isolette C2000 | CF32804 | Maintenance | cleaning pan cover | Ready to use |  | Ariawan
7000052 | 2018-09-27 | RSUD Cilacap | Baby Incubator | Draeger/Isolette C2000 | CF32807 | Service | probe skin | Ready to use |  | Ariawan
3000020 | 2018-09-21 | RSUP Dr. Hasan Sadikin | ESWL | Medispec/E3000 | 05IIII47 | Survey | Inverter fault | Ready to use |  | Imam
12000054 | 2018-09-21 | RSUD Hj. Anna Lasmanah | Mesin Anasthesia | Draeger/Primus | ASJF0042 | Survey | System check breating fail | Ready to use |  | Amna
11000049A | 2018-09-20 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | AY 06914 | Survey | Cek heater chamber | Ready to use |  | Nasrul
11000049A | 2018-09-20 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | AY 06915 | Survey | Cek heater chamber | Ready to use |  | Nasrul
11000049A | 2018-09-20 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | FR 07841 | Service | Cek heater chamber | Ready to use |  | Nasrul
4000015 | 2018-09-20 | RSUD TARAKAN | C-Arm | Apelem/Evo R+ | 55118000653 | Service | Printer tidak konek | Ready to use |  | Bayu
2000016 | 2018-09-20 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Table error | Ready to use |  | Eddy
12000055 | 2018-09-19 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius CE | ARYJ 0111 | Service | o2 sensor fail | Ready to use |  | Amna
12000055 | 2018-09-19 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASEL 0138 | Service | o2 sensor fail | Ready to use |  | Amna
7000050 | 2018-09-17 | RSUD Brebes | Jaundice Meter | Draeger/JM 105 | B3503457 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-09-17
7000050 | 2018-09-17 | RSUD Brebes | Baby Incubator | Draeger/Isolette C2000 | PH36453 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-09-17
7000050 | 2018-09-17 | RSUD Brebes | Baby Incubator | Draeger/Isolette C2000 | PH36454 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-09-17
7000050 | 2018-09-17 | RSUD Brebes | Baby Incubator | Draeger/Isolette C2000 | PH36455 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-09-17
7000050 | 2018-09-17 | RSUD Brebes | Baby Incubator | Draeger/Isolette C2000 | PH36456 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-09-17
3000018 | 2018-09-14 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Upgrade SSD komputer | Ready to use |  | Imam
7000049 | 2018-09-13 | RSUD DR. Soedirman - Kebumen | Incubator Transport | Draeger/Isolette TI500 | BH12348 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-09-13
7000049 | 2018-09-13 | RSUD DR. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | NH35832 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-09-13
12000053 | 2018-09-13 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASJE0005 | Service | System check ventilator fail | Ready to use |  | Amna
7000048 | 2018-09-12 | RSUD DR. R. Goeteng Taroenadibrata | Patient Monitor | Draeger/Vista 120 | VTSKL0013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-09-12
7000047 | 2018-09-10 | RSUD Banyumas | Ventilator | Draeger/Evita V300 | ASEL0120 | Maintenance | Kompresor mati total | Ready to use |  | Ariawan
12000051 | 2018-09-07 | RSUD BENDAN | Mesin Anasthesia | Draeger/Fabius Tiro | ASHK0143 | Maintenance | General Check | Ready to use |  | Amna
11000048 | 2018-09-07 | RSUD Kota Bau-Bau | Incubator Transport | Draeger/Isolette TI500 | MH12473 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-09-07
12000052 | 2018-09-07 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0011 | Maintenance | General Check | Ready to use |  | Amna
12000052 | 2018-09-07 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0014 | Service | General Check | Ready to use |  | Amna
7000046 | 2018-09-06 | RSUD Kardinah Tegal | ESWL | Medispec/EM1000 | 109117 | Maintenance | penggantian EMH | Ready to use |  | Ariawan
4000016 | 2018-09-06 | Setiabudi Sky Garden | ESMR | Medispec/ED1000 | 1116125 | Service | pengisian air | Ready to use |  | Bayu
5000004 | 2018-09-05 | RSUD KOTA TARAKAN | Mamografi | Metaltronica/Lilyum | ILILHF/0709/CB2 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal | 2019-09-05
12000050 | 2018-09-05 | RSUD RAA Soewondo Pati | Phototherapy | Draeger/PT 4000 | ASHE0009 | Service | Hours center lebih dari 1000 jam | Ready to use |  | Amna
12000050 | 2018-09-05 | RSUD RAA Soewondo Pati | Phototherapy | Draeger/PT 4000 | ASHE0012 | Service | Hours center lebih dari 1000 jam | Ready to use |  | Amna
12000050 | 2018-09-05 | RSUD RAA Soewondo Pati | Phototherapy | Draeger/PT 4000 | ASHE0020 | Service | Hours center lebih dari 1000 jam | Ready to use |  | Amna
11000047 | 2018-09-04 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0030 | Service | Internal Battery is missing | Ready to use |  | Nasrul
11000047 | 2018-09-04 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0392 | Service | Internal Battery is missing | Ready to use |  | Nasrul
1000044 | 2018-09-04 | RSUD Kardinah Tegal | ESWL | Medispec/EM1000 | 109117 | Service | EMH patah | Ready to use |  | Ariawan
7000045 | 2018-09-04 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASEF0030 | Service | Presure regulator mendengung | Ready to use |  | Ariawan
2000015 | 2018-09-03 | RSUD Kab. Bekasi | C-Arm | Apelem/Evo R+ | 55118026649 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2019-09-03
7000043 | 2018-09-03 | RSUD Brebes | Ventilator | Draeger/Babylog VN500 | ASLF0032 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-09-03
7000043 | 2018-09-03 | RSUD Brebes | Ventilator | Draeger/Babylog VN500 | ASLF0051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-09-03
7000043 | 2018-09-03 | RSUD Brebes | Ventilator | Draeger/Babylog VN500 | ASLF0053 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-09-03
7000054 | 2018-09-02 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ARWT0079 | Service | Pergantian Service setting 3 tahunan | Ready to use |  | Ariawan
7000054 | 2018-09-02 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0072 | Service | Pergantian Service setting 6 tahunan | Ready to use |  | Ariawan
7000054 | 2018-09-02 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Primus | ASDM0073 | Service | Pergantian Service setting 6 tahunan | Ready to use |  | Ariawan
7000054 | 2018-09-02 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0071 | Service | Pergantian Service setting 3 tahunan | Ready to use |  | Ariawan
7000054 | 2018-09-02 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0072 | Service | Pergantian Service setting 3 tahunan | Ready to use |  | Ariawan
7000054 | 2018-09-02 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEF0073 | Service | Pergantian Service setting 3 tahunan | Ready to use |  | Ariawan
7000054 | 2018-09-02 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0048 | Service | Pergantian Service setting 3 tahunan | Ready to use |  | Ariawan
4000014 | 2018-08-31 | RS Dr. Suyoto | Mobile X-ray | Apelem/Rafale EV30 | 009/17/01212 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2019-08-31
12000049 | 2018-08-31 | RSUD Dr. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0041 | Service | 02 sensor fail | Ready to use |  | Amna
11000046 | 2018-08-31 | RSUD Lakipadada | Bed Patient | Paramount Bed/PA 6415 | 9005061-00007 | Service | Bagian kaki tidak bisa naik turun | Ready to use |  | Nasrul
2000014 | 2018-08-30 | RSUD Dr. M. M. Dunda | C-Arm | Apelem/Evo R+ | 55118013636 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2020-08-30
11000045 | 2018-08-30 | RSUD SYEKH YUSUP KAB. GOWA | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0024 | Service | Alarm O2 sensor Failed | Ready to use |  | Nasrul
11000045 | 2018-08-30 | RSUD SYEKH YUSUP KAB. GOWA | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0028 | Service | Alarm O2 sensor Failed | Ready to use |  | Nasrul
11000045 | 2018-08-30 | RSUD SYEKH YUSUP KAB. GOWA | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0057 | Service | Alarm O2 sensor Failed | Ready to use |  | Nasrul
3000016 | 2018-08-29 | RSUD Tualang | X-ray Digital | Apelem/Camargue DR | CPD30502813 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2020-08-29
2000013 | 2018-08-29 | RSUD Tani dan Nelayan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/018/C0 | Service | Instalasi ulang tomo & biopsy | Ready to use |  | Eddy | 2020-08-29
11000043 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Examination Table | Paramount Bed/PK-1115 | 9007010-000043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000043 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Examination Table | Paramount Bed/PK-1115 | 9007010-00038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000043 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Examination Table | Paramount Bed/PK-1115 | 9007010-00044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000043 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Examination Table | Paramount Bed/PK-1115 | 9007010-00045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000043 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Examination Table | Paramount Bed/PK-1115 | 9007010-0041 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000043 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Examination Table | Paramount Bed/PK-1115 | 9007010-0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000043 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Examination Table | Paramount Bed/PK-1115 | 9007010-0046 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000043 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Examination Table | Paramount Bed/PK-1115 | 9007010-0070 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000043 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Examination Table | Paramount Bed/PK-1115 | 9007010-039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000044 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Bed I Crank | Paramount Bed/DA aa210 | 9007063-00100 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000044 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Bed I Crank | Paramount Bed/DA aa210 | 9007063-00112 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000044 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Bed I Crank | Paramount Bed/DA aa210 | 9007063-00113 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000044 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Bed I Crank | Paramount Bed/DA aa210 | 9007063-00154 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000044 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Bed I Crank | Paramount Bed/DA aa210 | 9007063-00155 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000044 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Bed I Crank | Paramount Bed/DA aa210 | 9007063-00156 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000044 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Bed I Crank | Paramount Bed/DA aa210 | 9007063-00157 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000044 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Bed I Crank | Paramount Bed/DA aa210 | 9007063-00158 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000044 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Bed I Crank | Paramount Bed/DA aa210 | 9007063-00159 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000044 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Bed I Crank | Paramount Bed/DA aa210 | 9007063-00161 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000043 | 2018-08-27 | Dinas Kesehatan Kab. Bulukumba | Examination Table | Paramount Bed/PK-1115 | 9007101-00033 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
12000048 | 2018-08-24 | RSUD Kajen | Resuscitaire | Draeger/Resuscitaire | HF10402 | Survey | Error 05 | Ready to use |  | Amna
7000041 | 2018-08-23 | RSUD Bumiayu | Medical Compressor | Draeger/Medical Compressor | ASKL0034 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-08-23
7000041 | 2018-08-23 | RSUD Bumiayu | Medical Compressor | Draeger/Medical Compressor | ASKL0036 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-08-23
7000041 | 2018-08-23 | RSUD Bumiayu | Medical Compressor | Draeger/Medical Compressor | ASKL0037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-08-23
7000041 | 2018-08-23 | RSUD Bumiayu | Medical Compressor | Draeger/Medical Compressor | ASKL0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-08-23
12000047 | 2018-08-21 | RSUD Dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0514 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2019-08-21
12000046 | 2018-08-20 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2019-08-20
12000046 | 2018-08-20 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2019-08-20
4000013 | 2018-08-16 | RSUD Ulin | Digital Radiografi Fluoroskopi | Apelem/Platinum | J176F0243 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bayu | 2019-08-16
3000015 | 2018-08-15 | RSUD Tualang | X-ray Digital | Apelem/Camargue DR | CPD30502813 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2019-08-15
7000039 | 2018-08-15 | RSUD Dr. Moewardi | Bone Mineral Densitometer | Medilink/Medix DR | A18016M301 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-08-15
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 000138 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00109 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00132 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00134 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00135 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00136 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00137 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000039 | 2018-08-15 | RSU MADANI - PALU | Physiciatric | Paramount Bed | 00138 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000040 | 2018-08-15 | RSU MADANI - PALU | Incubator Transport | Draeger/Isolette TI500 | MH12471 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-08-15
12000045 | 2018-08-15 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASEF0030 | Service | Systeam check nebulizer | Ready to use |  | Amna
12000043 | 2018-08-13 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASHM0015 | Service | Leak Hight | Ready to use |  | Amna
12000043 | 2018-08-13 | RSUD Dr. Soeselo Slawi | Mesin Anasthesia | Draeger/Primus | ASJE0005 | Service | Ventilator piston trouble | Ready to use |  | Amna
11000038 | 2018-08-13 | RSUD Kota Makassar | X-ray Digital | Apelem/Camargue DR | 84R481 | Service | Printer fault | Ready to use |  | Nasrul
2000012 | 2018-08-10 | RSU Haji Medan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/033/C0 | Maintenance | alat mati total | Ready to use |  | Eddy | 2020-08-10
4000012 | 2018-08-10 | RSUD Kab. Temanggung | X-Ray Analog | Apelem/Camargue Analog | CPD30673C18 | Survey | Pelaksanaan Uji Fungsi & Uji Paparan | Ready to use |  | Bayu | 2020-08-10
7000037 | 2018-08-09 | RSUD Dr. M. Ashari | Patient Monitor | Draeger/Vista 120 | VTSLC0051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-08-09
7000037 | 2018-08-09 | RSUD Dr. M. Ashari | Patient Monitor | Draeger/Vista 120 | VTSLC0059 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-08-09
7000037 | 2018-08-09 | RSUD Dr. M. Ashari | Patient Monitor | Draeger/Vista 120 | VTSLC0086 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-08-09
7000037 | 2018-08-09 | RSUD Dr. M. Ashari | Patient Monitor | Draeger/Vista 120 | VTSLC0168 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-08-09
7000037 | 2018-08-09 | RSUD Dr. M. Ashari | Patient Monitor | Draeger/Vista 120 | VTSLC109 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-08-09
2000011 | 2018-08-08 | RSUP Dr. H. Adam Malik Medan | Bone Mineral Densitometer | Medilink/Medix DR | J17016M280 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2020-08-08
12000041 | 2018-08-07 | RSUD K.R.M.T Wongsonegoro | Baby Incubator | Draeger/Isolette 8000 | MH00429 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2019-08-07
11000033 | 2018-08-07 | DINAS KESEHATAN KAB. BUOL | Incubator Transport | Draeger/Isolette TI500 | DN12426 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-08-07
3000013 | 2018-08-03 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Service | Software platinum not running | Ready to use |  | Imam
3000014 | 2018-08-03 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Service | Software not running | Ready to use |  | Imam
12000040 | 2018-08-03 | RSUD Dr. Soeprapto Cepu | Mesin Anasthesia | Draeger/Fabius Plus | ASJF0041 | Maintenance | Device Check | Ready to use |  | Amna
12000040 | 2018-08-03 | RSUD Dr. Soeprapto Cepu | Ventilator | Draeger/Savina 300 | ASJK0211 | Survey | Device Check | Ready to use |  | Amna
4000011 | 2018-08-02 | RS. Dr. ERNALDI BAHAR - PALEMBANG | X-Ray Analog | Apelem/Camargue Analog | CPD30672C18 | Maintenance | Pelaksanaan uji kesesuaian | Ready to use |  | Bayu | 2019-08-02
5000003 | 2018-08-02 | RSUD TARAKAN | C-Arm | Apelem/Evo R+ | 55118030653 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Bilal | 2019-08-02
4000010 | 2018-08-02 | RSUD TARAKAN | C-Arm | Apelem/Evo R+ | 55118030653 | Service | image no connect | Ready to use |  | Bayu
7000036 | 2018-07-31 | RS Bhakti Asih Brebes | Ventilator | Draeger/Savina 300 | ASHD0011 | Service | tidal volume tidak tercapai | Ready to use |  | Ariawan
7000034 | 2018-07-30 | RSUD Dr. Soeselo Slawi | Patient Monitor | Draeger/Vista 120 | V3SHJ0001 | Service | perlu penggantian bart motor | Ready to use |  | Ariawan
7000035 | 2018-07-30 | RSUD Kardinah Tegal | ESWL | Medispec/EM1000 | 6111205 | Survey | alat mati total | Ready to use |  | Ariawan
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08031 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08426 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08427 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08428 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08430 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08431 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08432 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08433 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08434 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08435 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08436 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08437 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000033 | 2018-07-28 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | MH08443 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-28
7000032 | 2018-07-27 | RSUD DR. Soedirman - Kebumen | Mesin Anasthesia | Draeger/Primus | ASKJ0050 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-27
11000030 | 2018-07-27 | RSUD Undata Palu | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0075 | Service | Tidal volume tidak sesuai setting | Ready to use |  | Nasrul
11000031 | 2018-07-27 | RSUD Undata Palu | Mesin Anasthesia | Draeger/Fabius Plus | ASFN0081 | Service | O2 sensor failed | Ready to use |  | Nasrul
2000009 | 2018-07-26 | RSUD Kota Makassar | X-ray Digital | Apelem/Camargue DR | CPD 22055E15 | Service | Hasil gambar banyak bercak | Ready to use |  | Eddy
3000012 | 2018-07-24 | RSUD Malingping | X-ray Digital | Apelem/Camargue DR | IMA703916 | Service | Adjust kolimasi | Ready to use |  | Imam
11000029 | 2018-07-24 | RSU Inco Soruwako | Paramount Bed | Pediatric Bed/Paramount | 9003072-00002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000029 | 2018-07-24 | RSU Inco Soruwako | Paramount Bed | Pediatric Bed/Paramount | 9003072-00003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000029 | 2018-07-24 | RSU Inco Soruwako | Paramount Bed | Pediatric Bed/Paramount | 9003072-00004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000029 | 2018-07-24 | RSU Inco Soruwako | Paramount Bed | Pediatric Bed/Paramount | 9003072-00005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000028 | 2018-07-23 | RSUD Kota Makassar | Stretcher Trolly | Paramount Bed/KK 7285 | 1968630-00019 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000028 | 2018-07-23 | RSUD Kota Makassar | Stretcher Trolly | Paramount Bed/KK 7285 | 1968630-00023 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000028 | 2018-07-23 | RSUD Kota Makassar | Stretcher Trolly | Paramount Bed/KK 7285 | 1968630-00024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000028 | 2018-07-23 | RSUD Kota Makassar | Stretcher Trolly | Paramount Bed/KK 7285 | 1968630-00037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000028 | 2018-07-23 | RSUD Kota Makassar | Stretcher Trolly | Paramount Bed/KK 7285 | 1968640-00003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
7000031 | 2018-07-23 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius CE | ARWC0079 | Survey | presure regulator rusak | Ready to use |  | Ariawan
12000039 | 2018-07-23 | RSUP Dr. Kariadi | Infant Warmer | Draeger/Babytherm 8004 | ARUC0013 | Survey | Device Error | Ready to use |  | Amna
11000026 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Emergency Troly | Paramount | 500000051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000026 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Emergency Troly | Paramount | 5001400047 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000026 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Emergency Troly | Paramount | 5001400048 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 19683000022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 19684000007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 19686400009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000026 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Emergency Troly | Paramount | 50051400008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000026 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Emergency Troly | Paramount | 50051400043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000026 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Emergency Troly | Paramount | 50051400045 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000026 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Emergency Troly | Paramount | 50051400053 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000026 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Emergency Troly | Paramount | 50051400055 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000026 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Emergency Troly | Paramount | 05000514-00003 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000026 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Emergency Troly | Paramount | 05000514-00005 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 136830-00021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 196830-00027 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 196830-00033 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 196830-00038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 196830-00040 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 1968630-00020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 1968630-00028 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 1968630-00034 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 1968630-00037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 1968630-00039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 1968640-00004 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000027 | 2018-07-22 | RSUD H. ANDI SULTHAN DAENG RADJA KAB. BULUKUMBA | Transfer strecher | Paramound Bed/kk7285 | 1968640-00008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
4000009 | 2018-07-21 | RS. Dr. ERNALDI BAHAR - PALEMBANG | X-Ray Analog | Apelem/Camargue Analog | CPD30672C18 | Survey | Setting colimator dan uji paparan | Ready to use |  | Bayu
12000038 | 2018-07-20 | RS Mardi Rahayu | Mesin Anasthesia | Draeger/Fabius Tiro | ASKL0012 | Service | Interkoneksi anasthesi dan polaris | Ready to use |  | Amna
11000025 | 2018-07-20 | RSUD Bahteramas | Baby Incubator | Draeger/Caleo | ASKL0010 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-07-20
11000025 | 2018-07-20 | RSUD Bahteramas | Baby Incubator | Draeger/Caleo | ASKL0021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-07-20
11000025 | 2018-07-20 | RSUD Bahteramas | Baby Incubator | Draeger/Caleo | ASLE0001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-07-20
11000025 | 2018-07-20 | RSUD Bahteramas | Baby Incubator | Draeger/Caleo | ASLE0002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-07-20
11000042 | 2018-07-20 | RSUD Bahteramas | Infant Warmer | Draeger/Babytherm 8010 | ASLE0024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-07-20
11000032 | 2018-07-20 | RSUD Bahteramas | Incubator Transport | Draeger/Isolette TI500 | MN12469 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-07-20
11000032 | 2018-07-20 | RSUD Bahteramas | Incubator Transport | Draeger/Isolette TI500 | MN12470 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
7000027 | 2018-07-20 | RSUI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | ARLF0023 | Service | leak pada saat digunakan pasien | Ready to use |  | Ariawan
7000028 | 2018-07-20 | RSUI Harapan Anda | Ventilator | Draeger/Evita 4 Edition | ASCC0210 | Maintenance | Device Failure | Ready to use |  | Ariawan
4000008 | 2018-07-20 | RSUP Dr. Mohammad Hoesin (KSO) | Digital Radiografi Fluoroskopi | Apelem/Platinum | L156F0142 | Maintenance | Cleaning CPU | Ready to use |  | Bayu
12000037 | 2018-07-19 | RSI Sultan Agung | Mesin Anasthesia | Draeger/Primus | ASKF0024 | Maintenance | General Check | Ready to use |  | Amna
11000053 | 2018-07-19 | RSUD Bahteramas | Phototherapy | Draeger/PT 4000 | ASKE 0254 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul | 2019-07-19
2000008 | 2018-07-19 | RSUD Dr. Fauziah | ESWL | Medispec/EM1000 | 716249 | Service | Repair thermal | Ready to use |  | Eddy
3000011 | 2018-07-18 | RSUD Dr. Soehadi Prijonegoro | ESWL | Medispec/EM1000 | 1016250 | Service | Unit mati | Ready to use |  | Imam
7000029 | 2018-07-18 | RSUD Dr. M. Ashari | Mesin Anasthesia | Draeger/Primus | ASLD0025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-18
2000007 | 2018-07-18 | RSU Haji Medan | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/033/C0 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2020-07-18
7000026 | 2018-07-17 | RSIA Bunda Arif | Jaundice Meter | Draeger/JM 105 | B3501432 | Service | kalibrasi unit | Ready to use |  | Ariawan
2000010 | 2018-07-17 | RSUD Tenriawaru Bone | X-ray Digital | Apelem/Camargue DR | CPD26084L16 | Service | Hasil foto rongent objek terpotong | Ready to use |  | Eddy
4000006 | 2018-07-17 | RSUP Dr. Mohammad Hoesin (KSO) | Mobile X-ray | Apelem/Rafale EV30 | 009/15/01095 | Maintenance | Stting koliimator | Ready to use |  | Bayu
12000036 | 2018-07-17 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHA0114 | Maintenance | Device Check | Ready to use |  | Amna
12000036 | 2018-07-17 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Fabius Tiro | ASHC0076 | Maintenance | Device Check | Ready to use |  | Amna
12000036 | 2018-07-17 | Rumah Sakit St. Elisabeth | Mesin Anasthesia | Draeger/Primus | ASKE0213 | Maintenance | General Check | Ready to use |  | Amna
11000017 | 2018-07-16 | RSUD I Lagaligo | Emergency Troly | paramount/KY 40Q12 | 30015 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
12000035 | 2018-07-16 | RSUD Dr. R. Soedjati Soemodiarjo | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0103 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2019-07-16
12000033 | 2018-07-16 | RSUD Kraton Pekalongan | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASHM0009 | Service | Layar pada LCD error | Ready to use |  | Amna
7000025 | 2018-07-16 | RSUI Harapan Anda | Mesin Anasthesia | Draeger/Fabius Plus | URZD0137 | Service | pergantian Service set | Ready to use |  | Ariawan
5000002 | 2018-07-13 | RS Premier Jatinegara (KSO) | ESST | Medispec/ED1000 | 0311124 | Service | Shockwave menembak dengan pelan | Ready to use |  | Bilal
8000073 | 2018-07-13 | RSUD Kraton Pekalongan | Lampu Operasi | Draeger/Polaris | ASKD6461 | Service | Kamera lampu mati | Ready to use |  | Luhur
8000073 | 2018-07-13 | RSUD Kraton Pekalongan | Lampu Operasi | Draeger/Polaris | ASKD9499 | Maintenance | LCD dan kabel kamera error | Ready to use |  | Luhur
4000005 | 2018-07-13 | RSUD Malingping | X-ray Digital | Apelem/Camargue DR | CPD30501816 | Survey | Pelaksanaan uji kesesuaian dan uji paparan | Ready to use |  | Bayu
2000006 | 2018-07-13 | RSUD Ulin | Printer Dry | Carestream DV 5960 | 59541226 | Service | Printer berfungsi dengan baik | Ready to use |  | Eddy
11000023 | 2018-07-12 | RSUD Kota Makassar | Pedriatric Bed | Paramount | 3100043 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
7000024 | 2018-07-12 | RSUD Dr. M. Ashari | Mesin Anasthesia | Draeger/Primus | ASDL0025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-12
7000023 | 2018-07-11 | RSUD Brebes | Resuscitaire | Draeger/Resuscitaire | AF11911 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-11
11000055 | 2018-07-11 | RSUP Dr. Wahidin Sudirohusodo | USG/Ecno | E84OTE/7340 | 4325 | Service | Cursor tidak dapat bergerak | Ready to use |  | Nasrul
11000020 | 2018-07-11 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 110443 | Service | EMH Tidak mmaksimal | Ready to use |  | Nasrul
11000054 | 2018-07-11 | RSUP Dr. Wahidin Sudirohusodo | Baby Incubator | Medispec/CS 2 | 8111362 | Maintenance | Aplicator ESMR tidak maksimal | Ready to use |  | Nasrul
12000034 | 2018-07-10 | RSUD BENDAN | Baby Incubator | Draeger/Isolette C2000 | CU27888 | Maintenance | Device Check | Ready to use |  | Amna
12000034 | 2018-07-10 | RSUD BENDAN | Baby Incubator | Draeger/Isolette C2000 | CU29513 | Maintenance | Device Check | Ready to use |  | Amna
12000034 | 2018-07-10 | RSUD BENDAN | Baby Incubator | Draeger/Isolette C2000 | CU29513 | Maintenance | Device Check | Ready to use |  | Amna
12000034 | 2018-07-10 | RSUD BENDAN | Baby Incubator | Draeger/Isolette C2000 | CU29514 | Maintenance | Device Check | Ready to use |  | Amna
12000034 | 2018-07-10 | RSUD BENDAN | Baby Incubator | Draeger/Isolette C2000 | CU29515 | Maintenance | Device Check | Ready to use |  | Amna
12000034 | 2018-07-10 | RSUD BENDAN | Resuscitaire | Draeger/Resuscitaire | DR11735 | Maintenance | Device Check | Ready to use |  | Amna
7000022 | 2018-07-10 | RSUD Brebes | Resuscitaire | Draeger/Resuscitaire | AR11911 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan | 2019-07-10
7000021 | 2018-07-10 | RSUD Brebes | Ventilator | Draeger/Evita V300 | ASHA-0095 | Service | leak pada saat breathing circuit | Ready to use |  | Ariawan
8000072 | 2018-07-10 | RSUD Dr. Moewardi | HCU Neonatus | Draeger/Caleo | ASKL0011 | Service | O2 cell failed | Ready to use | HCrlu penggran module control electronic mainboard, valve block brake arm, gas connector | Luhur
6000073 | 2018-07-10 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Service | APL Unconnectional | Ready to use |  | Afrizal
8000071 | 2018-07-09 | RSUD Waras Wiris Kab. Boyolali | Patient Monitor | Draeger/Infinity Vista XL | 6007841767 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Luhur | 2019-07-09
8000071 | 2018-07-09 | RSUD Waras Wiris Kab. Boyolali | Patient Monitor | Draeger/Infinity Vista XL | 6007858072 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Luhur | 2019-07-09
12000032 | 2018-07-09 | RSUD Suradadi Kab. Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0182 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna | 2019-07-09
11000037 | 2018-07-09 | RSUD Embung Fatimah Batam | Pedriatric Bed | Paramount Bed | 900110-00020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000037 | 2018-07-09 | RSUD Embung Fatimah Batam | Pedriatric Bed | Paramount Bed | 900110-00021 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000037 | 2018-07-09 | RSUD Embung Fatimah Batam | Pedriatric Bed | Paramount Bed | 900110-00024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000037 | 2018-07-09 | RSUD Embung Fatimah Batam | Pedriatric Bed | Paramount Bed | 9005110-00001 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
7000020 | 2018-07-09 | RSUD Banyumas | Lampu Operasi | Draeger polaris 700 | 100000000122490 | Service | Lampu mati | Ready to use |  | Ariawan
7000020 | 2018-07-09 | RSUD Banyumas | Lampu Operasi | Draeger polaris 700 | 100000000123586 | Service | Lampu mati | Ready to use |  | Ariawan
11000037 | 2018-07-09 | RSUD Embung Fatimah Batam | Pedriatric Bed | Paramount Bed | 9005110-00014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000037 | 2018-07-09 | RSUD Embung Fatimah Batam | Pedriatric Bed | Paramount Bed | 9005110-00022 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000037 | 2018-07-09 | RSUD Embung Fatimah Batam | Pedriatric Bed | Paramount Bed | 9005110-00023 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000037 | 2018-07-09 | RSUD Embung Fatimah Batam | Pedriatric Bed | Paramount Bed | 9005110-00025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000037 | 2018-07-09 | RSUD Embung Fatimah Batam | Pedriatric Bed | Paramount Bed | 9005110-00026 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000037 | 2018-07-09 | RSUD Embung Fatimah Batam | Pedriatric Bed | Paramount Bed | 9005110-00031 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000035 | 2018-07-09 | RSUD Andi Makassau | Paramount Bed | Paramount Bed/RD 260 | 9005110-0037 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000035 | 2018-07-09 | RSUD Andi Makassau | Paramount Bed | Paramount Bed/RD 260 | 9005110-0038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000034 | 2018-07-09 | RSUD Andi Makassau | Emergency Stretcher | Paramount Bed/PK 80318 | 9006007-00012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000034 | 2018-07-09 | RSUD Andi Makassau | Emergency Stretcher | Paramount Bed/PK 80318 | 9006007-00013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000034 | 2018-07-09 | RSUD Andi Makassau | Emergency Stretcher | Paramount Bed/PK 80318 | 9006007-00016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
6000072 | 2018-07-09 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0039 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-07-09
6000072 | 2018-07-09 | RSJ Prof Dr Soerojo Magelang | Mesin Anasthesia | Draeger/Primus | ASLD0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-07-09
3000009 | 2018-07-06 | RS. Dr. ERNALDI BAHAR - PALEMBANG | X-Ray Analog | Apelem/Camargue Analog | CPD30672C18 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2020-07-06
12000031 | 2018-07-06 | RSUD dr. R. Soetijono Blora | Mesin Anasthesia | Draeger/Fabius Plus | ASJD0084 | Service | Piston tidak bergerak | Ready to use |  | Amna
2000005 | 2018-07-05 | RSUD Kab. Temanggung | ESWL | Medispec/EM1000 | 1115056 | Service | Center target kurang tepat | Ready to use |  | Eddy
11000016 | 2018-07-04 | RSUD Salewangan | Baby Incubator | Draeger/Isolette 8000 | BRO7492 | Service | Alarm humidity on display | Ready to use |  | Nasrul
2000004 | 2018-07-03 | RSUD Kab. Temanggung | X-Ray Analog | Apelem/Camargue Analog | CPD30673C18 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Eddy | 2019-07-03
4000004 | 2018-07-02 | RS. Dr. ERNALDI BAHAR - PALEMBANG | X-Ray Analog | Apelem/Camargue Analog | CPD30672C18 | Pra Instalasi | Pemasangan panel listrik | Ready to use |  | Bayu
12000030 | 2018-07-02 | RSI Sultan Agung | Baby Incubator | Draeger/Isolette C2000 | CR35553 | Service | Controller Failure 4 | Ready to use |  | Amna
11000015 | 2018-07-02 | RSUP Tadjudin Chalid | Ventilator | Draeger/Savina 300 | ASJL0030 | Service | Battery internal alarm low | Ready to use |  | Nasrul
8000070 | 2018-07-02 | RSUP Dr. Kariadi | Baby Incubator | Draeger/Isolette 8000 | MR07738 | Service | Arm lengan patah | Ready to use |  | Luhur
12000029 | 2018-06-28 | RSUD dr. R. Soetrasno | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
7000018B | 2018-06-26 | RS. MITRA SIAGA - TEGAL | Lampu Operasi | Draeger/Polaris 600 | ASKM0062 | Service | Pemasangan ring cover | Ready to use |  | Ariawan
7000018B | 2018-06-26 | RS. MITRA SIAGA - TEGAL | Lampu Operasi | Draeger/Polaris 600 | ASKM068 | Service | Pemasangan ring cover | Ready to use |  | Ariawan
7000018A | 2018-06-25 | RSUD Hj. Anna Lasmanah | Resuscitaire | Draeger/Resuscitaire | CI11218 | Maintenance | Penganntian corner blok kit | Ready to use |  | Ariawan
12000028 | 2018-06-25 | RSUP Dr. Kariadi | Ventilator | Draeger/Babylog VN500 | ASJC0044 | Maintenance | Device Check | Ready to use |  | Amna
12000027 | 2018-06-25 | RSUP Dr. Kariadi | Baby Incubator | Draeger/Isolette 8000 | DY06740 | Maintenance | Device Check | Ready to use |  | Amna
12000028 | 2018-06-25 | RSUP Dr. Kariadi | Ventilator | Draeger/Evita V300 | ASJM0044 | Maintenance | General Check | Ready to use |  | Amna
8000069 | 2018-06-23 | RSUD Brebes | Phototherapy | Draeger/PT 4000 | ASKB0013 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Luhur
8000069 | 2018-06-23 | RSUD Brebes | Phototherapy | Draeger/PT 4000 | ASKE0089 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Luhur
12000026 | 2018-06-22 | RSUI Harapan Anda | Ventilator | Draeger/Evita 4 edition | ASCC0210 | Service | Device Failure/General Check | Ready to use |  | Amna
12000025 | 2018-06-22 | RSUI Harapan Anda | Ventilator | Draeger/Savina 300 | ASHA0095 | Service | Alat tidak bisa di standby | Ready to use |  | Amna
8000067 | 2018-06-22 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Primus | ASKF0069 | Maintenance | Saat manual muncul getaran | Ready to use |  | Luhur
12000024 | 2018-06-15 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0118 | Maintenance | General Check | Ready to use |  | Amna
12000024 | 2018-06-15 | RSUD RA Kartini | Mesin Anasthesia | Draeger/Primus | ASHM0012 | Service | General Check | Ready to use |  | Amna
7000008 | 2018-06-11 | RSUD Kardinah Tegal | Mesin Anasthesia | Draeger/Fabius Plus | ASHL0027 | Service | Pemasangan o2 Cell | Ready to use |  | Ariawan
6000050 | 2018-06-11 | RSUD Tidar Magelang | Medical Compressor | Draeger/Medical Compressor | ASKD0007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan
6000050 | 2018-06-11 | RSUD Tidar Magelang | Medical Compressor | Draeger/Medical Compressor | ASKD0011 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan
12000023 | 2018-06-10 | RSUD Jepara | Ventilator | Draeger/Evita V300 | ASHD0187 | Survey | Device Check | Ready to use |  | Amna
12000023 | 2018-06-10 | RSUD Jepara | Ventilator | Draeger/Evita V300 | SW0251 | Service | Device Check | Ready to use |  | Amna
8000065 | 2018-06-08 | RSUD Kota Salatiga | Pendant | Draeger/Movita | ASJL0292 | Service | Arm pendant tidak bisa berhenti | Ready to use |  | Luhur
6000068 | 2018-06-08 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASKM0402 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000068 | 2018-06-08 | RSUD Prambanan | Patient Monitor | Draeger/Vista 120 | VTSKL0002 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
7000017 | 2018-06-08 | RSUD Prof. Dr. Margono Soekarjo | Mesin Anasthesia | Draeger/Fabius CE | ARWC0079 | Service | Mendengung di bagian flowneter | Ready to use |  | Ariawan
7000016 | 2018-06-07 | RS Panti Rapih | Mesin Anasthesia | Draeger/Fabius Plus | ASKL0009 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan
12000022 | 2018-06-07 | RSUD BENDAN | Resuscitaire | Draeger/Resuscitaire | DR11735 | Survey | Training User | Ready to use |  | Amna
12000022 | 2018-06-07 | RSUD BENDAN | Resuscitaire | Draeger/Resuscitaire | MR11828 | Service | Training User | Ready to use |  | Amna
8000064 | 2018-06-07 | RSUD Kraton Pekalongan | Mesin Anasthesia | Draeger/Primus | ASHM0011 | Survey | Self test result | Ready to use |  | Luhur
12000021 | 2018-06-06 | RSUD RA Kartini | Baby Incubator | Draeger/Isolette C2000 | NH35833 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
6000067 | 2018-06-06 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASHL0055 | Survey | O2 calibrating | Ready to use |  | Afrizal
6000067 | 2018-06-06 | RSUD Prambanan | Mesin Anasthesia | Draeger/Primus | ASHL0055 | Service | O2 calibrating | Ready to use |  | Afrizal
12000020 | 2018-06-05 | RSUD BENDAN | Resuscitaire | Draeger/Resuscitaire | DR11735 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000020 | 2018-06-05 | RSUD BENDAN | Resuscitaire | Draeger/Resuscitaire | MR11828 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
11000012 | 2018-06-05 | RSUD Tenriawaru Bone | X-ray Digital | Apelem/Camargue DR | CPD26084L16 | Maintenance | generator tidak ready | Ready to use |  | Nasrul
12000019 | 2018-06-02 | Hospital Siloam Semarang | Ventilator Transport | Draeger | ASKL0038 | Maintenance | Device Check | Ready to use |  | Amna
12000019 | 2018-06-02 | Hospital Siloam Semarang | Ventilator Transport | Draeger/Oxylog 2000 Plus | ASKN0288 | Maintenance | Device Check | Ready to use |  | Amna
12000019 | 2018-06-02 | Hospital Siloam Semarang | Incubator Transport | Draeger/Isolette TI500 | CR12070 | Maintenance | Device Check | Ready to use |  | Amna
12000019 | 2018-06-02 | Hospital Siloam Semarang | Baby Incubator | Draeger/Isolette C2000 | CR33565 | Service | Device Check | Ready to use |  | Amna
7000015 | 2018-05-31 | RSUD Cilacap | Resuscitaire | Draeger/Resuscitaire | CY11244 | Maintenance | Pemasangan kit corner blok K | Ready to use |  | Ariawan
12000018 | 2018-05-31 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 Edition | ASEL0142 | Maintenance | General Check | Ready to use |  | Amna
12000018 | 2018-05-31 | RSUD Dr. Loekmono Hadi | Ventilator | Draeger/Evita 4 Edition | ASEL0143 | Survey | General Check | Ready to use |  | Amna
12000017 | 2018-05-31 | RSUD Dr. Loekmono Hadi | Patient Monitor | Draeger/Vista 120 | V3SEF0005 | Survey | General Check | Ready to use |  | Amna
12000017 | 2018-05-31 | RSUD Dr. Loekmono Hadi | Patient Monitor | Draeger/Vista 120 | V3SEF0009 | Service | General Check | Ready to use |  | Amna
11000010 | 2018-05-31 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | AY06913 | Service | Error suply | Ready to use |  | Nasrul
11000010 | 2018-05-31 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | AY06914 | Service | Error suply | Ready to use |  | Nasrul
11000010 | 2018-05-31 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | AY06915 | Service | Error suply | Ready to use |  | Nasrul
11000010 | 2018-05-31 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | FR07841 | Service | Error suply | Ready to use |  | Nasrul
11000010 | 2018-05-31 | RSUD I Lagaligo | Baby Incubator | Draeger/Isolette 8000 | FR07842 | Service | Error suply | Ready to use |  | Nasrul
11000011 | 2018-05-31 | RSUD I Lagaligo | Patient Monitor | Draeger/Vista 120 | V3SHN0074 | Service | Spo2 adult tidak terbaca | Ready to use |  | Nasrul
11000011 | 2018-05-31 | RSUD I Lagaligo | Patient Monitor | Draeger/Vista 120 | V3SHN0098 | Service | Spo2 adult tidak terbaca | Ready to use |  | Nasrul
2000002 | 2018-05-31 | RSUP Dr. H. Adam Malik Medan | Bone Mineral Densitometer | Medilink/Medix DR | J17016M280 | Service | Ukes BPFK | Ready to use |  | Eddy
8000063 | 2018-05-30 | RS Nasional Diponegoro | Resuscitaire | Draeger/Resuscitaire | MF02734 | Service | C-Pap error | Ready to use |  | Luhur
8000063 | 2018-05-30 | RS Nasional Diponegoro | Resuscitaire | Draeger/Resuscitaire | MF10231 | Service | C-Pap error | Ready to use |  | Luhur
6000066 | 2018-05-30 | RS Panti Rapih | Ventilator | Draeger/Savina 300 | ARSC0003 | Maintenance | pemasangan pressure regulator | Ready to use |  | Afrizal
12000016 | 2018-05-28 | RSUD Dr. R. Soedjati Soemodiarjo | Incubator Transport | Draeger/Isolette TI500 | DY11073 | Maintenance | General Check | Ready to use |  | Amna
12000016 | 2018-05-28 | RSUD Dr. R. Soedjati Soemodiarjo | Incubator Transport | Draeger/Isolette TI500 | KF10739 | Service | General Check | Ready to use |  | Amna
6000064 | 2018-05-28 | RSUD Tidar Magelang | Incubator Transport | Draeger/Isolette TI500 | CY11377 | Survey | suhu tidak panas | Ready to use |  | Afrizal
6000063 | 2018-05-26 | RSUD Dr. Tjitrowardojo Purworejo | Pendant | Draeger/Movita | ASHE0161 | Service | pendant tidak berfungsi dengan baik | Ready to use | Disarankan penggran module control electronic mainboard, valve block brake arm, gas connector | Afrizal
8000062 | 2018-05-25 | RST Bhakti Wira Tamtama | Baby Incubator | Draeger/Isolette C2000 | CY34226 | Service | Suara van agak bising | Ready to use |  | Luhur
8000062 | 2018-05-25 | RST Bhakti Wira Tamtama | Baby Incubator | Draeger/Isolette C2000 | CY34227 | Service | Suara van agak bising | Ready to use |  | Luhur
12000015 | 2018-05-25 | RSUD Kayen Pati | Infant Warmer | Draeger/Babytherm 3004 | ASEJ0008 | Service | Error 25/Running Test | Ready to use |  | Amna
6000062 | 2018-05-25 | RSUD Tidar Magelang | Baby Incubator | Draeger/Caleo | ASJN0063 | Service | suhu 37 tidak tercapai | Ready to use |  | Afrizal
11000009 | 2018-05-24 | RS Siloam Hospital Makassar | Ventilator | Draeger/Evita V300 | ASDL0234 | Service | Tidal volume tidak sesuai setting | Ready to use |  | Nasrul
8000060 | 2018-05-24 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDL0002 | Service | O2 sensor failed | Ready to use |  | Luhur
8000061 | 2018-05-24 | RS Telogorejo | Mesin Anasthesia | Draeger/Fabius GS Premium | ASDLC3000 | Maintenance | O2 sensor failed | Ready to use |  | Luhur
12000014 | 2018-05-24 | RSUD dr. R. Soetijono Blora | Mesin Anasthesia | Draeger/Fabius Plus | ASJD00085 | Maintenance | General Check | Ready to use |  | Amna
12000014 | 2018-05-24 | RSUD dr. R. Soetijono Blora | Mesin Anasthesia | Draeger/Fabius Plus | ASJD0084 | Maintenance | General Check | Ready to use |  | Amna
2000001 | 2018-05-24 | RSUD Raden Mattaher | Digital Radiografi Fluoroskopi | Apelem/Platinum | G617F0228 | Survey | Error table platinum | Ready to use |  | Eddy
6000061 | 2018-05-24 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/fabius XL | ASDL0007 | Survey | cek unit | Ready to use |  | Afrizal
6000061 | 2018-05-24 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger Scio Four | ASKD0152 | Service | scio four ditarik kantor | Ready to use |  | Afrizal
12000013 | 2018-05-23 | RSUD dr. H. Soewondo Kendal | Patient Monitor | Draeger/Vista 120 | V25EB0109 | Maintenance | Kalibrasi/running Test | Ready to use |  | Amna
7000013 | 2018-05-23 | RSUD Kardinah Tegal | Ventilator | Draeger/Savina 300 | ASCF0129 | Service | Perlu penggantian Service set dan o2 cell | Ready to use |  | Ariawan
7000012 | 2018-05-23 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASEF0030 | Service | Presure sensor, expitatory, safet value failed | Ready to use |  | Ariawan
7000012 | 2018-05-23 | RSUD Kardinah Tegal | Ventilator | Draeger/Babylog VN500 | ASEF0108 | Survey | Presure sensor, expitatory, safet value failed | Ready to use |  | Ariawan
8000059 | 2018-05-23 | RSUD Kota Salatiga | Baby Incubator | Draeger/Isolette C2000 | CY34224 | Survey | General Check | Ready to use |  | Luhur
6000058 | 2018-05-23 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKE0069 | Survey | meng"set: ventilator ke nicu | Ready to use |  | Afrizal
6000058 | 2018-05-23 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKE0070 | Survey | meng"set: ventilator ke nicu | Ready to use |  | Afrizal
6000060 | 2018-05-23 | RSUD Panembahan Senopati Bantul | Mamografi | Metaltronica/ Helianthus DBT | HELIDBT/027/C0 | Service | UPS bermasalah di bawa ke kantor | Ready to use |  | Afrizal
11000008 | 2018-05-23 | RSUD Undata Palu | Mesin Anasthesia | Draeger/Fabius Plus | ASEH0075 | Maintenance | O2 sensor failed | Ready to use |  | Nasrul
12000012 | 2018-05-22 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASJF0033 | Maintenance | General Check | Ready to use |  | Amna
12000012 | 2018-05-22 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0198 | Maintenance | General Check | Ready to use |  | Amna
12000012 | 2018-05-22 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | ASKE0199 | Maintenance | General Check | Ready to use |  | Amna
12000012 | 2018-05-22 | RSUD Dr. Loekmono Hadi | Mesin Anasthesia | Draeger/Primus | AX11003 | Service | General Check | Ready to use |  | Amna
8000056 | 2018-05-22 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Caleo | ASJE0010 | Service | O2 cell failed | Ready to use |  | Luhur
7000011 | 2018-05-22 | RSUD DR. R. Goeteng Taroenadibrata | Baby Incubator | Draeger/Isolette 8000 | CU03921 | Service | Alarm stuck key | Ready to use |  | Ariawan
8000057 | 2018-05-22 | RSUD Kota Salatiga | Resuscitaire | Draeger/Resuscitaire | RM02182 | Service | General Check | Ready to use |  | Luhur
6000057 | 2018-05-21 | RS Panti Rapih | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0193 | Service | System leak test failed | Ready to use |  | Afrizal
7000010 | 2018-05-21 | RSUD Banyumas | Incubator Transport | Draeger/Isolette TI500 | AF11203 | Survey | battery rusak | Ready to use |  | Ariawan
4000002 | 2018-05-21 | RSUD Kab Tangerang | ESWL | Medispec/EM1000 | 710599 | Survey | Tabung pengisi air tidak ada | Ready to use |  | Bayu
12000011 | 2018-05-21 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0093 | Service | Training User | Ready to use |  | Amna
8000055 | 2018-05-21 | RSUD Pandan Arang | Incubator Transport | Draeger/Isolette TI500 | BH12345 | Service | Battery tidak mau ngecas | Ready to use |  | Luhur
6000056 | 2018-05-21 | RSUP Dr. Sardjito | Baby Incubator | Draeger/Isolette C2000 | HM31340 | Service | Humidity error | Ready to use |  | Afrizal
4000001 | 2018-05-19 | RSUD Kab Tangerang | ESWL | Medispec/EM1000 | 710599 | Service | Re-install | Ready to use |  | Bayu
1100006 | 2018-05-18 | RSUD Andi Djemma Masamba | Ventilator | Draeger/Carina | ASJK0068 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
6000055 | 2018-05-18 | RSUD Bagas Waras Klaten | Patient Monitor | Draeger/Vista 120 | V3SHK0141 | Service | SpO2 tidak terbaca | Ready to use |  | Afrizal
8000053 | 2018-05-18 | RSUP Dr. Kariadi | Ventilator Transport | Draeger/Oxylog | RSUC0012 | Survey | Service check running | Ready to use |  | Luhur
1200010 | 2018-05-17 | RS Nasional Diponegoro | Incubator Transport | Draeger/Isolette TI500 | AF02870 | Maintenance | Training User | Ready to use |  | Amna
12000009 | 2018-05-17 | RS Nasional Diponegoro | Baby Incubator | Draeger/Isolette C2000 | AF32530 | Maintenance | General Check | Ready to use |  | Amna
1200009 | 2018-05-17 | RS Nasional Diponegoro | Baby Incubator | Draeger/Isolette C2000 | AF32536 | Maintenance | General Check | Ready to use |  | Amna
12000009 | 2018-05-17 | RS Nasional Diponegoro | Baby Incubator | Draeger/Isolette C2000 | AF32538 | Maintenance | General Check | Ready to use |  | Amna
12000009 | 2018-05-17 | RS Nasional Diponegoro | Baby Incubator | Draeger/Isolette C2000 | AF32539 | Service | General Check | Ready to use |  | Amna
60000054 | 2018-05-17 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0079 | Service | Monitor touch screen tidak berfungsi | Ready to use |  | Afrizal
8000052 | 2018-05-16 | RSUD Dr. Soeratno Gemolong | Patient Monitor | Draeger/Vista 120 | VTSKK0049 | Service | SpO2 tidak terbaca | Ready to use |  | Luhur
12000007 | 2018-05-16 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKK0017 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000007 | 2018-05-16 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKK0464 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000007 | 2018-05-16 | RSUD Sunan Kalijaga | Mesin Anasthesia | Draeger/Fabius Plus | ASKK0842 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000008 | 2018-05-16 | RSUD Sunan Kalijaga | Baby Incubator | Draeger/Isolette C2000 | CF32801 | Service | General Check | Ready to use |  | Amna
11000004 | 2018-05-16 | RSUP Dr. Wahidin Sudirohusodo | Mesin Anasthesia | Draeger/Fabius Plus | ASEL0068 | Service | Kebocoran pada cosy | Ready to use |  | Nasrul
6000053 | 2018-05-15 | RSUP Dr. Sardjito | Mesin Anasthesia | Draeger/Perseus A500 | ASKF0069 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal | 2019-05-15
11000007 | 2018-05-15 | RSUD Andi Djemma Masamba | Ventilator | Draeger/Carina | ASJK0068 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
11000003 | 2018-05-14 | RSUP Tadjudin Chalid | Infant Warmer | Draeger/Babytherm 8004 | ASHD0038 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
6000052 | 2018-05-14 | RSUD Dr. Tjitrowardojo Purworejo | Baby Incubator | Draeger/Isolette 8000 | CR08033 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
8000051 | 2018-05-14 | RSUD BENDAN | Infant Warmer | Drger/Bsbytherm 8004 | ASCM0145 | Service | Heater kurang panas | Ready to use |  | Luhur
6000052 | 2018-05-14 | RSUD Dr. Tjitrowardojo Purworejo | Baby Incubator | Draeger/Isolette 8000 | CR8029 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
11000014 | 2018-05-14 | RSUD Andi Djemma Masamba | Ventilator | Draeger/Carina | ASJK0069 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Nasrul
12000006 | 2018-05-14 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0093 | Service | Kondensasi breating circuit tinggi | Ready to use |  | Amna
12000006 | 2018-05-14 | RSUD K.R.M.T Wongsonegoro | Ventilator | Draeger/Babylog VN500 | ASKL0094 | Service | Kondensasi breating circuit tinggi | Ready to use |  | Amna
12000005 | 2018-05-09 | RSI Sunan Kudus | Ventilator | Draeger/Savina 300 | ARUE0022 | Service | TV Low | Ready to use |  | Amna
12000005 | 2018-05-09 | RSI Sunan Kudus | Ventilator | Draeger/Savina 300 | ARUM0308 | Survey | TV Low | Ready to use |  | Amna
7000005 | 2018-05-09 | RSUD PREMBUN KAB. KEBUMEN | Ventilator | Draeger/Babylog VN500 | ASJC0044 | Survey | pengambilan demo unit | Ready to use |  | Ariawan
7000006 | 2018-05-09 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | CR08031 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan
7000005 | 2018-05-09 | RSUD PREMBUN KAB. KEBUMEN | Baby Incubator | Draeger/Isolette 8000 | NM06493 | Survey | pengambilan demo unit | Ready to use |  | Ariawan
6000049 | 2018-05-09 | RSUP dr. Soeradji Tirtonegoro | Incubator Transport | Draeger/Isolette TI500 | KF10706 | Service | alat bunyi dengung | Ready to use |  | Afrizal
8000050 | 2018-05-08 | RS Dr. Oen Solo Baru | Mesin Anasthesia | Draeger/Fabius Plus | ASHA0145 | Service | O2 cell failed | Ready to use |  | Luhur
6000048 | 2018-05-08 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASJH0126 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000048 | 2018-05-08 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASJH0127 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000048 | 2018-05-08 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASJH0128 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000047 | 2018-05-08 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKE0068 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000047 | 2018-05-08 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKE0070 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
6000048 | 2018-05-08 | RSUD Panembahan Senopati Bantul | Ventilator | Draeger/Evita V300 | ASKL0094 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Afrizal
12000002 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKF0007 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000002 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKF0008 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000004 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKF0016 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000003 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKF0030 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000003 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKF0031 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000004 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKF0042 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000001 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKF0044 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000002 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKF0049 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000002 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKF0051 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000004 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKL0006 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
7000004 | 2018-05-08 | RSUD DR. Soedirman - Kebumen | Incubator Transport | Draeger/Isolette TI500 | BH12346 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan
7000004 | 2018-05-08 | RSUD DR. Soedirman - Kebumen | Baby Incubator | Draeger/Isolette C2000 | NH35832 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan
12000001 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKL0012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000001 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKL0014 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000001 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKL0020 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000004 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKL0023 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000003 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKL0024 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
12000003 | 2018-05-08 | RSUD Dr. R. Soedjati Soemodiarjo | Patient Monitor | Draeger/Vista 120 | VTSKL0025 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
8000049 | 2018-05-07 | RSUD Pandan Arang | Baby Incubator | Draeger/Isolette C2000 | CF03034 | Service | Corner block patah | Ready to use |  | Luhur
8000048 | 2018-05-04 | RSUD Dr. Moewardi | Baby Incubator | Draeger/Caleo | ASJE0010 | Service | Clean mode muncul alarm | Ready to use |  | Luhur
11000001 | 2018-04-26 | RSUP Dr. Wahidin Sudirohusodo | C-Arm | Apelem/Evo R+ | 54114050674 | Survey | Mati total | Ready to use |  | Nasrul
7000001 | 2018-04-20 | RSUD Prof. Dr. Margono Soekarjo | Ventilator | Draeger/Savina 300 | ARWB 0347 | Service | Backup pressure sensor | Ready to use |  | Ariawan
11000022 | 2018-04-10 | RSUP Dr. Wahidin Sudirohusodo | ESWL | Medispec/EM1000 | 110443 | Service | EMH Tidak maksimal | Ready to use |  | Nasrul
3000007 | 2018-03-31 | RSUD Malingping | X-ray Digital | Apelem/Camargue DR | IMA703916 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Imam | 2020-03-31
3000006 | 2018-03-22 | RSUD Kab Tangerang | C-Arm | Apelem/Evo R+ | 551161245 | Service | Software DSA 500 tidak bekerja | Ready to use |  | Imam
7000085 | 2018-03-08 | RSUI Harapan Anda | Incubator Transport | Draeger/Isolette TI500 | HH12697 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan
12000084 | 2018-01-22 | RSUD RAA Soewondo Pati | Mesin Anasthesia | Draeger/Primus | ASJK0014 | Service | Mixer fail | Ready to use |  | Amna
12000057 | 2017-10-04 | RS Bhina Bhakti Husada | Central Monitor | Draeger/Vista 120 CMS | VLSKM 0058 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Amna
6000051 | 2017-06-11 | RSUD Tidar Magelang | Medical Compressor | Draeger/Medical Compressor | ASKD0012 | Instalasi | Alat berfungsi dengan baik | Ready to use |  | Ariawan
6000051 | 2017-06-11 | RSUD Tidar Magelang | Medical Compressor | Draeger/Medical Compressor | ASKF0059 |  | Alat berfungsi dengan baik | Ready to use |  | Ariawan
20000331 | 2002-02-20 | RSUD Morowali | CPAP | Draeger/Oxymixer | QM20230004 | Service | Flow meter bermasalah, pembersihan unit | Pending | Unit perlu diganti | Sujana
Formulas:
MASA BERAKHIR GARANSI = EDATE([TANGGAL],24)